## Summary
NASA Summary of VIIRS F2 Ambient Testing Calibration
Contents: results from F2 Ambient testing including RC-1, FP-10, FP-11, FP-12, and FP-14.
All quantities listed by electronics side, HAM side, and subsample where appropriate. Detector dependent, detector median, and detector average quantities listed.
Version 1: Intital results from all tests compiled. | 2014-04-01
Version 2: Added in results from post TVAC polarization testing | 2015-01-01

## TEB rad coeff
Band | HAM | Subsample | Gain State | Band Median | [7 empty] | Band Average
|  |  |  | Eside A |  |  |  | Eside B |  |  |  | Eside A |  |  |  | Eside B
|  |  |  | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1)
I4 | A | 1 | H | -0.0024258 | 0.00087187 | -3.3339e-09 | 1146.9636 | -0.0037754 | 0.00088049 | -6.737e-09 | 1135.7279 | -0.0024112 | 0.00087165 | -3.2895e-09 | 1147.2498 | -0.0037578 | 0.00088113 | -6.8301e-09 | 1134.9106
I4 | A | 2 | H | -0.0024219 | 0.00087208 | -3.2759e-09 | 1146.6845 | -0.0037649 | 0.00088057 | -6.7511e-09 | 1135.6301 | -0.0024075 | 0.0008717 | -3.3069e-09 | 1147.1872 | -0.0037527 | 0.00088113 | -6.8328e-09 | 1134.9069
I4 | B | 1 | H | -0.0024251 | 0.00087218 | -3.3758e-09 | 1146.5515 | -0.0037723 | 0.00088079 | -6.6503e-09 | 1135.3418 | -0.0024346 | 0.00087192 | -3.3399e-09 | 1146.8899 | -0.0036864 | 0.00088135 | -6.7149e-09 | 1134.617
I4 | B | 2 | H | -0.0024565 | 0.00087223 | -3.3508e-09 | 1146.4824 | -0.0037766 | 0.00088077 | -6.6399e-09 | 1135.3681 | -0.0024317 | 0.00087196 | -3.3497e-09 | 1146.843 | -0.0037195 | 0.0008814 | -6.7312e-09 | 1134.5559
I5 | A | 1 | H | -0.026157 | 0.0054008 | -2.2856e-08 | 185.1567 | -0.014407 | 0.0055308 | -4.016e-09 | 180.8051 | -0.025959 | 0.0054131 | -1.8927e-08 | 184.7363 | -0.013274 | 0.0055436 | 4.6305e-10 | 180.3869
I5 | A | 2 | H | -0.026878 | 0.005405 | -2.3498e-08 | 185.0139 | -0.013818 | 0.0055303 | -3.0133e-09 | 180.822 | -0.026379 | 0.0054141 | -1.9685e-08 | 184.7018 | -0.012999 | 0.0055425 | 9.8424e-10 | 180.4241
I5 | B | 1 | H | -0.026568 | 0.0054051 | -2.7429e-08 | 185.0103 | -0.012408 | 0.005533 | -1.8403e-09 | 180.7324 | -0.026219 | 0.0054188 | -2.0921e-08 | 184.5417 | -0.012348 | 0.0055456 | 2.6706e-09 | 180.3227
I5 | B | 2 | H | -0.025872 | 0.0054031 | -2.6488e-08 | 185.0799 | -0.012277 | 0.005533 | -2.7716e-09 | 180.7349 | -0.026288 | 0.0054185 | -2.0803e-08 | 184.5523 | -0.012102 | 0.0055446 | 3.1204e-09 | 180.3541
M12 | A | 1 | H | -0.0023728 | 0.00086665 | -5.035e-09 | 1153.8711 | -0.003616 | 0.00087697 | -8.745e-09 | 1140.2854 | -0.0023775 | 0.00086651 | -5.0164e-09 | 1154.0485 | -0.003643 | 0.00087769 | -8.7729e-09 | 1139.3583
M12 | B | 1 | H | -0.002404 | 0.00086693 | -5.0733e-09 | 1153.5016 | -0.0035175 | 0.00087721 | -8.6327e-09 | 1139.9742 | -0.0024021 | 0.00086678 | -5.0631e-09 | 1153.7003 | -0.0035853 | 0.00087794 | -8.6672e-09 | 1139.0265
M13 | A | 1 | H | -0.0051737 | 0.0016477 | -1.1777e-08 | 606.9119 | -0.0075895 | 0.0016677 | -1.9047e-08 | 599.6257 | -0.0052116 | 0.0016466 | -1.1623e-08 | 607.3091 | -0.0074838 | 0.0016672 | -1.9432e-08 | 599.8034
M13 | B | 1 | H | -0.0052056 | 0.0016476 | -1.1797e-08 | 606.946 | -0.0075032 | 0.0016679 | -1.8877e-08 | 599.553 | -0.0052591 | 0.0016466 | -1.1726e-08 | 607.3031 | -0.007442 | 0.0016674 | -1.9225e-08 | 599.7366
M14 | A | 1 | H | -0.036996 | 0.0052447 | 1.5049e-08 | 190.6673 | -0.020869 | 0.0054361 | 3.377e-08 | 183.9549 | -0.037162 | 0.0052673 | 1.5618e-08 | 189.852 | -0.021657 | 0.0054567 | 3.5174e-08 | 183.2606
M14 | B | 1 | H | -0.03703 | 0.0052323 | 1.5043e-08 | 191.1213 | -0.02025 | 0.0054233 | 3.4387e-08 | 184.3894 | -0.037085 | 0.0052543 | 1.522e-08 | 190.3215 | -0.020913 | 0.0054427 | 3.5614e-08 | 183.7326
M15 | A | 1 | H | -0.025863 | 0.0051753 | -1.1531e-08 | 193.227 | -0.022982 | 0.0053063 | -1.2839e-08 | 188.4536 | -0.02603 | 0.0051958 | -1.1914e-08 | 192.4613 | -0.024526 | 0.0053237 | -1.6081e-08 | 187.8378
M15 | B | 1 | H | -0.025981 | 0.005177 | -1.1863e-08 | 193.1611 | -0.022542 | 0.005307 | -1.2695e-08 | 188.4293 | -0.026074 | 0.0051968 | -1.2374e-08 | 192.4258 | -0.023946 | 0.0053243 | -1.5445e-08 | 187.8165
M16A | A | 1 | H | -0.026655 | 0.0049379 | -1.2992e-08 | 202.5134 | -0.015761 | 0.0050331 | 3.8612e-09 | 198.6832 | -0.026857 | 0.0049431 | -1.3278e-08 | 202.3029 | -0.015342 | 0.0050445 | 4.2418e-09 | 198.2353
M16A | B | 1 | H | -0.026718 | 0.0049406 | -1.3255e-08 | 202.4052 | -0.015328 | 0.0050344 | 4.9252e-09 | 198.6322 | -0.02687 | 0.0049477 | -1.3359e-08 | 202.1141 | -0.01453 | 0.005048 | 6.0205e-09 | 198.0989
M16B | A | 1 | H | -0.023508 | 0.0049448 | -3.1612e-09 | 202.2314 | -0.0085939 | 0.0050531 | 1.6946e-08 | 197.9002 | -0.022905 | 0.0049505 | -3.996e-09 | 201.9981 | -0.0084978 | 0.0050573 | 1.6804e-08 | 197.7347
M16B | B | 1 | H | -0.023148 | 0.0049516 | -5.2902e-09 | 201.9545 | -0.0078542 | 0.0050605 | 1.7895e-08 | 197.6088 | -0.023346 | 0.0049568 | -5.226e-09 | 201.7422 | -0.0079192 | 0.0050626 | 1.7173e-08 | 197.5278
M13 | A | 1 | L | 0.26753 | 0.12831 | -4.9455e-07 | 7.7936 | 0.23217 | 0.12966 | -5.2227e-07 | 7.7125 | 0.24053 | 0.12828 | -4.9856e-07 | 7.7952 | 0.19741 | 0.12964 | -5.2754e-07 | 7.7139
M13 | B | 1 | L | 0.26858 | 0.12833 | -4.9244e-07 | 7.7924 | 0.23314 | 0.12972 | -5.2333e-07 | 7.7092 | 0.24264 | 0.12831 | -4.9663e-07 | 7.7936 | 0.19839 | 0.1297 | -5.2643e-07 | 7.7101
Band | HAM | Detector | Subsample | Gain State | Detector Dependent | [8 empty] | Notes:
[5 empty] | Eside A |  |  |  | Eside B |  |  |  |  | 1. This table provides VIIRS F2 TEB radiometric coefficients.
[5 empty] | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) |  | 2. The results were derived from RC-1 using the LABB source (TMC BB for M13 low gain).
I4 | A | 1 | 1 | H | -0.0024471 | 0.00087666 | -3.8478e-09 | 1140.695 | -0.0041499 | 0.00088989 | -8.3181e-09 | 1123.7312 |  | 3. Quadratic fits of the LABB path radiance difference vs dn were used to determine the radiometric coefficients.
I4 | A | 2 | 1 | H | -0.002514 | 0.00088294 | -3.5605e-09 | 1132.5835 | -0.0041495 | 0.00089437 | -7.9268e-09 | 1118.1014 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I4 | A | 3 | 1 | H | -0.0024907 | 0.00087403 | -3.4094e-09 | 1144.1238 | -0.0042394 | 0.00088755 | -8.0413e-09 | 1126.6918 |  | 5. The test was set to operational day mode.
I4 | A | 4 | 1 | H | -0.0025551 | 0.00088225 | -3.9158e-09 | 1133.4683 | -0.0041911 | 0.00089408 | -8.3429e-09 | 1118.4619
I4 | A | 5 | 1 | H | -0.0023738 | 0.00087271 | -2.5612e-09 | 1145.8514 | -0.003915 | 0.00088576 | -6.9669e-09 | 1128.9708
I4 | A | 6 | 1 | H | -0.0025465 | 0.00088617 | -3.1698e-09 | 1128.4505 | -0.0041362 | 0.00089656 | -7.3624e-09 | 1115.3719
I4 | A | 7 | 1 | H | -0.002331 | 0.00087195 | -2.4655e-09 | 1146.8521 | -0.0038999 | 0.00088348 | -6.6186e-09 | 1131.8898
I4 | A | 8 | 1 | H | -0.0025557 | 0.00087977 | -3.7671e-09 | 1136.6545 | -0.0041103 | 0.00089015 | -7.8064e-09 | 1123.4007
I4 | A | 9 | 1 | H | -0.0024265 | 0.0008672 | -3.2098e-09 | 1153.1385 | -0.0039393 | 0.00087841 | -7.0199e-09 | 1138.4224
I4 | A | 10 | 1 | H | -0.0025252 | 0.00087608 | -3.5652e-09 | 1141.4423 | -0.0039691 | 0.00088623 | -7.4574e-09 | 1128.3805
I4 | A | 11 | 1 | H | -0.0024408 | 0.0008693 | -3.1892e-09 | 1150.353 | -0.0038664 | 0.00088056 | -7.053e-09 | 1135.6405
I4 | A | 12 | 1 | H | -0.0025353 | 0.00087319 | -3.117e-09 | 1145.2208 | -0.0039452 | 0.00088201 | -7.0047e-09 | 1133.7761
I4 | A | 13 | 1 | H | -0.0024252 | 0.00086179 | -3.5929e-09 | 1160.3759 | -0.0039035 | 0.0008721 | -7.3452e-09 | 1146.6629
I4 | A | 14 | 1 | H | -0.002576 | 0.00087161 | -3.5744e-09 | 1147.3081 | -0.0039726 | 0.00088033 | -7.1555e-09 | 1135.9355
I4 | A | 15 | 1 | H | -0.002441 | 0.00086087 | -3.436e-09 | 1161.6092 | -0.0036693 | 0.00087068 | -6.9026e-09 | 1148.5226
I4 | A | 16 | 1 | H | -0.0025202 | 0.00087667 | -3.6161e-09 | 1140.6755 | -0.0038895 | 0.00088505 | -7.0955e-09 | 1129.8766
I4 | A | 17 | 1 | H | -0.0022687 | 0.00086889 | -2.7733e-09 | 1150.9001 | -0.0035821 | 0.00087878 | -6.4438e-09 | 1137.943
I4 | A | 18 | 1 | H | -0.00258 | 0.00087515 | -3.215e-09 | 1142.6663 | -0.0038362 | 0.00088293 | -6.6047e-09 | 1132.5962
I4 | A | 19 | 1 | H | -0.0022547 | 0.00086558 | -2.9457e-09 | 1155.2997 | -0.0035474 | 0.00087489 | -6.2828e-09 | 1143.004
I4 | A | 20 | 1 | H | -0.0023495 | 0.00087863 | -3.0722e-09 | 1138.1388 | -0.0036293 | 0.00088674 | -6.3983e-09 | 1127.7218
I4 | A | 21 | 1 | H | -0.0023393 | 0.0008627 | -3.3399e-09 | 1159.1482 | -0.0035206 | 0.00087181 | -6.5331e-09 | 1147.0441
I4 | A | 22 | 1 | H | -0.0024063 | 0.00087082 | -3.4781e-09 | 1148.3491 | -0.0036257 | 0.00087864 | -6.6787e-09 | 1138.1174
I4 | A | 23 | 1 | H | -0.0022315 | 0.00086654 | -2.9527e-09 | 1154.0102 | -0.003532 | 0.00087596 | -6.2515e-09 | 1141.5985
I4 | A | 24 | 1 | H | -0.0024829 | 0.00087316 | -3.3278e-09 | 1145.2676 | -0.0037145 | 0.00088042 | -6.3402e-09 | 1135.8152
I4 | A | 25 | 1 | H | -0.0022974 | 0.0008677 | -3.7643e-09 | 1152.4761 | -0.0034773 | 0.00087683 | -6.7954e-09 | 1140.4655
I4 | A | 26 | 1 | H | -0.0024231 | 0.00087178 | -3.49e-09 | 1147.0751 | -0.0035764 | 0.00087895 | -6.3832e-09 | 1137.72
I4 | A | 27 | 1 | H | -0.0022383 | 0.00086382 | -3.039e-09 | 1157.6448 | -0.0034413 | 0.00087258 | -6.0342e-09 | 1146.0264
I4 | A | 28 | 1 | H | -0.0025151 | 0.00086923 | -3.2534e-09 | 1150.4387 | -0.0035091 | 0.00087592 | -6.0372e-09 | 1141.6629
I4 | A | 29 | 1 | H | -0.0023172 | 0.00085895 | -2.9059e-09 | 1164.2128 | -0.003381 | 0.00086698 | -5.6213e-09 | 1153.4231
I4 | A | 30 | 1 | H | -0.0021957 | 0.00087501 | -3.3634e-09 | 1142.8415 | -0.0033519 | 0.00088156 | -6.1298e-09 | 1134.3547
I4 | A | 31 | 1 | H | -0.0021633 | 0.0008661 | -2.6795e-09 | 1154.6055 | -0.0031999 | 0.00087428 | -5.4116e-09 | 1143.7944
I4 | A | 32 | 1 | H | -0.00239 | 0.00087554 | -3.6647e-09 | 1142.1556 | -0.0033778 | 0.00088154 | -6.2002e-09 | 1134.3771
I4 | A | 1 | 2 | H | -0.0025355 | 0.0008768 | -3.8864e-09 | 1140.5085 | -0.0042571 | 0.00089 | -8.369e-09 | 1123.6016
I4 | A | 2 | 2 | H | -0.0024532 | 0.00088279 | -3.4962e-09 | 1132.776 | -0.0041137 | 0.00089433 | -7.891e-09 | 1118.1508
I4 | A | 3 | 2 | H | -0.0026205 | 0.00087425 | -3.4751e-09 | 1143.8334 | -0.0042152 | 0.00088739 | -7.9705e-09 | 1126.9044
I4 | A | 4 | 2 | H | -0.0024856 | 0.00088236 | -3.9947e-09 | 1133.3303 | -0.0040422 | 0.00089386 | -8.2671e-09 | 1118.7474
I4 | A | 5 | 2 | H | -0.0024779 | 0.000873 | -2.6454e-09 | 1145.4724 | -0.0039369 | 0.0008858 | -7.0097e-09 | 1128.9254
I4 | A | 6 | 2 | H | -0.0024162 | 0.00088607 | -3.1215e-09 | 1128.5795 | -0.0039639 | 0.00089629 | -7.2942e-09 | 1115.7137
I4 | A | 7 | 2 | H | -0.0024267 | 0.00087215 | -2.5419e-09 | 1146.5935 | -0.0040831 | 0.00088373 | -6.678e-09 | 1131.5691
I4 | A | 8 | 2 | H | -0.0024005 | 0.00087946 | -3.6607e-09 | 1137.0675 | -0.0039758 | 0.00089007 | -7.7926e-09 | 1123.5044
I4 | A | 9 | 2 | H | -0.002496 | 0.00086714 | -3.1506e-09 | 1153.2207 | -0.0039363 | 0.00087841 | -7.0193e-09 | 1138.4189
I4 | A | 10 | 2 | H | -0.0025353 | 0.00087607 | -3.5332e-09 | 1141.4604 | -0.0037966 | 0.00088583 | -7.2897e-09 | 1128.8805
I4 | A | 11 | 2 | H | -0.0025125 | 0.00086935 | -3.1951e-09 | 1150.2889 | -0.0040283 | 0.0008807 | -7.0889e-09 | 1135.4585
I4 | A | 12 | 2 | H | -0.0025025 | 0.00087341 | -3.2321e-09 | 1144.9374 | -0.0038515 | 0.00088196 | -6.9934e-09 | 1133.8352
I4 | A | 13 | 2 | H | -0.0025751 | 0.00086189 | -3.5929e-09 | 1160.239 | -0.0039579 | 0.000872 | -7.2573e-09 | 1146.7851
I4 | A | 14 | 2 | H | -0.0025419 | 0.00087201 | -3.7648e-09 | 1146.7755 | -0.0038265 | 0.00088042 | -7.2488e-09 | 1135.8211
I4 | A | 15 | 2 | H | -0.0024278 | 0.00086063 | -3.2972e-09 | 1161.9335 | -0.0038293 | 0.00087083 | -6.9411e-09 | 1148.3361
I4 | A | 16 | 2 | H | -0.0024171 | 0.00087669 | -3.6678e-09 | 1140.657 | -0.0038639 | 0.00088527 | -7.2467e-09 | 1129.5957
I4 | A | 17 | 2 | H | -0.0024663 | 0.00086951 | -3.0591e-09 | 1150.0716 | -0.0037288 | 0.00087893 | -6.4629e-09 | 1137.7499
I4 | A | 18 | 2 | H | -0.0024565 | 0.00087508 | -3.2088e-09 | 1142.7515 | -0.0036645 | 0.00088281 | -6.6157e-09 | 1132.7459
I4 | A | 19 | 2 | H | -0.0023966 | 0.00086568 | -2.9673e-09 | 1155.1653 | -0.0036388 | 0.00087506 | -6.344e-09 | 1142.7806
I4 | A | 20 | 2 | H | -0.0021447 | 0.00087851 | -3.0636e-09 | 1138.2921 | -0.0036633 | 0.00088706 | -6.5506e-09 | 1127.3228
I4 | A | 21 | 2 | H | -0.0023878 | 0.00086257 | -3.2546e-09 | 1159.3279 | -0.0037332 | 0.00087198 | -6.5366e-09 | 1146.8099
I4 | A | 22 | 2 | H | -0.0022581 | 0.00087077 | -3.4965e-09 | 1148.4046 | -0.0035287 | 0.00087852 | -6.6396e-09 | 1138.2724
I4 | A | 23 | 2 | H | -0.0023315 | 0.00086687 | -3.0819e-09 | 1153.5699 | -0.0034727 | 0.00087566 | -6.1147e-09 | 1141.9979
I4 | A | 24 | 2 | H | -0.0023866 | 0.00087321 | -3.3598e-09 | 1145.1958 | -0.0036613 | 0.00088044 | -6.3711e-09 | 1135.8019
I4 | A | 25 | 2 | H | -0.0023982 | 0.00086775 | -3.7533e-09 | 1152.403 | -0.0035371 | 0.00087691 | -6.8242e-09 | 1140.3628
I4 | A | 26 | 2 | H | -0.0022804 | 0.00087164 | -3.4469e-09 | 1147.2583 | -0.0034561 | 0.00087866 | -6.2761e-09 | 1138.0907
I4 | A | 27 | 2 | H | -0.0024001 | 0.00086397 | -3.0782e-09 | 1157.4502 | -0.0036241 | 0.00087283 | -6.0935e-09 | 1145.7046
I4 | A | 28 | 2 | H | -0.0022714 | 0.00086897 | -3.1733e-09 | 1150.7913 | -0.0033232 | 0.00087567 | -5.942e-09 | 1141.9838
I4 | A | 29 | 2 | H | -0.0024957 | 0.00085909 | -2.9064e-09 | 1164.0236 | -0.0033775 | 0.00086692 | -5.5656e-09 | 1153.5134
I4 | A | 30 | 2 | H | -0.0021718 | 0.00087523 | -3.4581e-09 | 1142.5631 | -0.0032765 | 0.00088149 | -6.1276e-09 | 1134.443
I4 | A | 31 | 2 | H | -0.0022442 | 0.00086613 | -2.6746e-09 | 1154.561 | -0.0033268 | 0.00087452 | -5.4952e-09 | 1143.4849
I4 | A | 32 | 2 | H | -0.0021265 | 0.00087527 | -3.5829e-09 | 1142.5035 | -0.003395 | 0.0008818 | -6.3337e-09 | 1134.0434
I4 | B | 1 | 1 | H | -0.0024467 | 0.00087696 | -3.9614e-09 | 1140.3092 | -0.0039951 | 0.00089008 | -8.2205e-09 | 1123.4927
I4 | B | 2 | 1 | H | -0.0026217 | 0.00088324 | -3.5949e-09 | 1132.1917 | -0.0040747 | 0.00089463 | -7.8166e-09 | 1117.7821
I4 | B | 3 | 1 | H | -0.0025397 | 0.00087454 | -3.5309e-09 | 1143.4559 | -0.003944 | 0.00088732 | -7.684e-09 | 1126.9937
I4 | B | 4 | 1 | H | -0.0025841 | 0.00088239 | -3.9625e-09 | 1133.2873 | -0.0039949 | 0.00089384 | -8.0699e-09 | 1118.7659
I4 | B | 5 | 1 | H | -0.0022747 | 0.00087286 | -2.6438e-09 | 1145.6635 | -0.0038142 | 0.00088585 | -6.8157e-09 | 1128.862
I4 | B | 6 | 1 | H | -0.0025172 | 0.0008861 | -3.0795e-09 | 1128.5374 | -0.004037 | 0.00089661 | -7.2036e-09 | 1115.3166
I4 | B | 7 | 1 | H | -0.0023356 | 0.00087197 | -2.3957e-09 | 1146.8237 | -0.0038531 | 0.00088379 | -6.5076e-09 | 1131.4911
I4 | B | 8 | 1 | H | -0.0026685 | 0.00088005 | -3.8393e-09 | 1136.2959 | -0.0040271 | 0.00089032 | -7.6726e-09 | 1123.1882
I4 | B | 9 | 1 | H | -0.0024376 | 0.00086728 | -3.1342e-09 | 1153.0325 | -0.0038416 | 0.00087869 | -6.9262e-09 | 1138.0612
I4 | B | 10 | 1 | H | -0.0025917 | 0.0008764 | -3.566e-09 | 1141.026 | -0.0039534 | 0.00088656 | -7.3362e-09 | 1127.9573
I4 | B | 11 | 1 | H | -0.0023726 | 0.00086943 | -3.2147e-09 | 1150.1737 | -0.003927 | 0.00088084 | -6.9168e-09 | 1135.2836
I4 | B | 12 | 1 | H | -0.0026372 | 0.00087367 | -3.2366e-09 | 1144.6032 | -0.0038984 | 0.00088226 | -6.8442e-09 | 1133.4464
I4 | B | 13 | 1 | H | -0.0022533 | 0.00086161 | -3.5043e-09 | 1160.6218 | -0.0037303 | 0.00087226 | -7.2449e-09 | 1146.4482
I4 | B | 14 | 1 | H | -0.0026197 | 0.00087221 | -3.7742e-09 | 1146.512 | -0.0038944 | 0.00088075 | -7.1549e-09 | 1135.3999
I4 | B | 15 | 1 | H | -0.002403 | 0.00086108 | -3.3967e-09 | 1161.3361 | -0.0038538 | 0.00087136 | -6.9038e-09 | 1147.6258
I4 | B | 16 | 1 | H | -0.0026307 | 0.00087701 | -3.6691e-09 | 1140.2329 | -0.0038639 | 0.0008854 | -7.075e-09 | 1129.433
I4 | B | 17 | 1 | H | -0.0023481 | 0.00086971 | -3.0426e-09 | 1149.8088 | -0.0035462 | 0.00087903 | -6.2915e-09 | 1137.6122
I4 | B | 18 | 1 | H | -0.0025428 | 0.00087531 | -3.2263e-09 | 1142.4466 | -0.0038369 | 0.00088343 | -6.6336e-09 | 1131.9547
I4 | B | 19 | 1 | H | -0.0024082 | 0.00086621 | -3.1008e-09 | 1154.453 | -0.0035431 | 0.00087549 | -6.325e-09 | 1142.2151
I4 | B | 20 | 1 | H | -0.0023748 | 0.00087888 | -3.1798e-09 | 1137.807 | -0.0036985 | 0.00088706 | -6.3381e-09 | 1127.3224
I4 | B | 21 | 1 | H | -0.0022744 | 0.00086284 | -3.4053e-09 | 1158.9608 | -0.0035621 | 0.00087217 | -6.4687e-09 | 1146.5679
I4 | B | 22 | 1 | H | -0.0023547 | 0.00087082 | -3.4805e-09 | 1148.3388 | -0.0035541 | 0.00087872 | -6.535e-09 | 1138.0195
I4 | B | 23 | 1 | H | -0.0022721 | 0.0008671 | -3.1045e-09 | 1153.2709 | -0.0033766 | 0.00087581 | -5.9601e-09 | 1141.7953
I4 | B | 24 | 1 | H | -0.0024317 | 0.00087344 | -3.3548e-09 | 1144.8954 | -0.0034954 | 0.00088062 | -6.2455e-09 | 1135.5698
I4 | B | 25 | 1 | H | -0.0024186 | 0.00086815 | -3.8824e-09 | 1151.8804 | -0.0033703 | 0.00087703 | -6.667e-09 | 1140.2086
I4 | B | 26 | 1 | H | -0.0025742 | 0.00087215 | -3.5576e-09 | 1146.591 | -0.0035882 | 0.00087924 | -6.288e-09 | 1137.34
I4 | B | 27 | 1 | H | -0.0022077 | 0.00086378 | -2.9897e-09 | 1157.7081 | -0.0034955 | 0.00087306 | -6.0704e-09 | 1145.3995
I4 | B | 28 | 1 | H | -0.0025705 | 0.00086944 | -3.2532e-09 | 1150.1681 | -0.0032747 | 0.00087575 | -5.7634e-09 | 1141.8788
I4 | B | 29 | 1 | H | -0.0022328 | 0.00085914 | -2.8942e-09 | 1163.9583 | -0.0033977 | 0.00086758 | -5.6388e-09 | 1152.6348
I4 | B | 30 | 1 | H | -0.0022684 | 0.00087529 | -3.4051e-09 | 1142.4806 | -0.0032097 | 0.0008817 | -6.0267e-09 | 1134.1707
I4 | B | 31 | 1 | H | -0.0022104 | 0.00086642 | -2.7364e-09 | 1154.1717 | -0.0029575 | 0.00087412 | -5.1652e-09 | 1144.0014
I4 | B | 32 | 1 | H | -0.0024833 | 0.00087606 | -3.7593e-09 | 1141.4763 | -0.0033565 | 0.00088198 | -6.0671e-09 | 1133.8135
I4 | B | 1 | 2 | H | -0.0025996 | 0.00087706 | -3.9396e-09 | 1140.167 | -0.0042369 | 0.00089025 | -8.2622e-09 | 1123.2755
I4 | B | 2 | 2 | H | -0.0024696 | 0.00088327 | -3.6196e-09 | 1132.1515 | -0.0040094 | 0.00089457 | -7.8276e-09 | 1117.8586
I4 | B | 3 | 2 | H | -0.0026847 | 0.00087469 | -3.5476e-09 | 1143.2656 | -0.0040808 | 0.00088739 | -7.6631e-09 | 1126.9047
I4 | B | 4 | 2 | H | -0.0025438 | 0.00088271 | -4.1384e-09 | 1132.88 | -0.0039867 | 0.000894 | -8.1365e-09 | 1118.5679
I4 | B | 5 | 2 | H | -0.0024622 | 0.00087298 | -2.6271e-09 | 1145.5008 | -0.0039193 | 0.00088596 | -6.8701e-09 | 1128.7254
I4 | B | 6 | 2 | H | -0.0023899 | 0.00088601 | -3.0641e-09 | 1128.6525 | -0.0039351 | 0.00089651 | -7.1872e-09 | 1115.4377
I4 | B | 7 | 2 | H | -0.002417 | 0.00087231 | -2.5512e-09 | 1146.3769 | -0.0040563 | 0.00088399 | -6.57e-09 | 1131.2364
I4 | B | 8 | 2 | H | -0.0024794 | 0.00087983 | -3.7494e-09 | 1136.5847 | -0.0040277 | 0.0008903 | -7.651e-09 | 1123.2211
I4 | B | 9 | 2 | H | -0.0025901 | 0.00086748 | -3.1929e-09 | 1152.7658 | -0.0039824 | 0.00087889 | -7.0155e-09 | 1137.7925
I4 | B | 10 | 2 | H | -0.0024088 | 0.00087629 | -3.5356e-09 | 1141.1721 | -0.0038477 | 0.00088637 | -7.2837e-09 | 1128.1922
I4 | B | 11 | 2 | H | -0.0025977 | 0.00086967 | -3.2758e-09 | 1149.8572 | -0.0039288 | 0.00088081 | -6.8951e-09 | 1135.3187
I4 | B | 12 | 2 | H | -0.0025069 | 0.00087356 | -3.2012e-09 | 1144.7399 | -0.0038925 | 0.00088239 | -6.9282e-09 | 1133.2805
I4 | B | 13 | 2 | H | -0.0024956 | 0.000862 | -3.6346e-09 | 1160.0944 | -0.0038311 | 0.00087219 | -7.2166e-09 | 1146.5373
I4 | B | 14 | 2 | H | -0.0025089 | 0.00087196 | -3.6574e-09 | 1146.8369 | -0.0038025 | 0.00088073 | -7.1526e-09 | 1135.4176
I4 | B | 15 | 2 | H | -0.0025471 | 0.00086113 | -3.3813e-09 | 1161.2615 | -0.0038858 | 0.00087139 | -6.9017e-09 | 1147.5972
I4 | B | 16 | 2 | H | -0.0024508 | 0.00087687 | -3.6831e-09 | 1140.4254 | -0.0038137 | 0.00088555 | -7.151e-09 | 1129.2363
I4 | B | 17 | 2 | H | -0.00247 | 0.00086964 | -2.9488e-09 | 1149.8959 | -0.0036866 | 0.00087925 | -6.3595e-09 | 1137.3323
I4 | B | 18 | 2 | H | -0.0024718 | 0.00087543 | -3.3051e-09 | 1142.2946 | -0.00363 | 0.00088306 | -6.5022e-09 | 1132.4226
I4 | B | 19 | 2 | H | -0.0024791 | 0.00086616 | -3.065e-09 | 1154.5182 | -0.0037506 | 0.00087562 | -6.3358e-09 | 1142.0453
I4 | B | 20 | 2 | H | -0.0022493 | 0.00087875 | -3.1311e-09 | 1137.9749 | -0.0036617 | 0.00088702 | -6.3265e-09 | 1127.3735
I4 | B | 21 | 2 | H | -0.0023913 | 0.00086277 | -3.3203e-09 | 1159.0569 | -0.0035363 | 0.00087208 | -6.4368e-09 | 1146.6792
I4 | B | 22 | 2 | H | -0.0022941 | 0.00087098 | -3.5554e-09 | 1148.1277 | -0.0034698 | 0.00087872 | -6.5431e-09 | 1138.0141
I4 | B | 23 | 2 | H | -0.0024337 | 0.00086709 | -3.0582e-09 | 1153.2854 | -0.0035284 | 0.00087594 | -5.9806e-09 | 1141.6318
I4 | B | 24 | 2 | H | -0.0024065 | 0.00087357 | -3.4456e-09 | 1144.7222 | -0.0034333 | 0.00088065 | -6.2624e-09 | 1135.526
I4 | B | 25 | 2 | H | -0.0023979 | 0.00086796 | -3.7734e-09 | 1152.1316 | -0.0034867 | 0.0008772 | -6.7097e-09 | 1139.9887
I4 | B | 26 | 2 | H | -0.0023822 | 0.00087215 | -3.5913e-09 | 1146.5879 | -0.0034957 | 0.00087905 | -6.2357e-09 | 1137.5855
I4 | B | 27 | 2 | H | -0.0024861 | 0.00086429 | -3.1383e-09 | 1157.0251 | -0.0035161 | 0.00087294 | -5.993e-09 | 1145.5565
I4 | B | 28 | 2 | H | -0.0022455 | 0.00086909 | -3.1601e-09 | 1150.6318 | -0.0033602 | 0.00087598 | -5.8551e-09 | 1141.5753
I4 | B | 29 | 2 | H | -0.002334 | 0.00085928 | -2.9434e-09 | 1163.7626 | -0.00343 | 0.00086759 | -5.6193e-09 | 1152.6196
I4 | B | 30 | 2 | H | -0.0022039 | 0.00087543 | -3.4984e-09 | 1142.3001 | -0.0032728 | 0.00088188 | -6.0894e-09 | 1133.9469
I4 | B | 31 | 2 | H | -0.0021516 | 0.00086632 | -2.7152e-09 | 1154.3042 | -0.0031108 | 0.00087442 | -5.2805e-09 | 1143.61
I4 | B | 32 | 2 | H | -0.0022651 | 0.00087593 | -3.7422e-09 | 1141.6403 | -0.0034171 | 0.00088216 | -6.1568e-09 | 1133.5813
I5 | A | 1 | 1 | H | -0.033773 | 0.0052249 | -2.6122e-08 | 191.3926 | -0.018398 | 0.0053242 | -1.0007e-08 | 187.8222
I5 | A | 2 | 1 | H | -0.0277 | 0.005413 | 1.2147e-08 | 184.7394 | -0.019256 | 0.0055616 | 2.1797e-08 | 179.8038
I5 | A | 3 | 1 | H | -0.029691 | 0.0051837 | -2.5772e-08 | 192.9138 | -0.015333 | 0.005314 | -8.5951e-09 | 188.1822
I5 | A | 4 | 1 | H | -0.029903 | 0.0052384 | 7.0469e-09 | 190.897 | -0.014676 | 0.0053702 | 2.5239e-08 | 186.2117
I5 | A | 5 | 1 | H | -0.02941 | 0.0050603 | -2.2475e-08 | 197.6165 | -0.015789 | 0.0051907 | -5.6741e-09 | 192.6539
I5 | A | 6 | 1 | H | -0.029634 | 0.0051843 | 2.407e-09 | 192.8913 | -0.01642 | 0.0053028 | 1.7036e-08 | 188.5784
I5 | A | 7 | 1 | H | -0.025888 | 0.0050174 | -1.8717e-08 | 199.3073 | -0.014463 | 0.005135 | -2.8318e-09 | 194.7405
I5 | A | 8 | 1 | H | -0.028036 | 0.0053305 | -2.3979e-08 | 187.5989 | -0.013682 | 0.0054345 | -3.7985e-09 | 184.0099
I5 | A | 9 | 1 | H | -0.022506 | 0.0050684 | -1.5972e-08 | 197.3005 | -0.015186 | 0.0052089 | -1.2119e-08 | 191.9792
I5 | A | 10 | 1 | H | -0.03311 | 0.0054212 | -3.5056e-08 | 184.4616 | -0.015408 | 0.0055392 | -1.3871e-08 | 180.5315
I5 | A | 11 | 1 | H | -0.025382 | 0.0049939 | -3.144e-08 | 200.2448 | -0.018405 | 0.0051277 | -1.9501e-08 | 195.0209
I5 | A | 12 | 1 | H | -0.027923 | 0.0053297 | -3.9767e-09 | 187.6288 | -0.016709 | 0.0054589 | 1.2059e-08 | 183.1887
I5 | A | 13 | 1 | H | -0.025985 | 0.0050724 | -1.8222e-08 | 197.1436 | -0.012505 | 0.0051828 | -3.9043e-09 | 192.9452
I5 | A | 14 | 1 | H | -0.01716 | 0.0055678 | -1.3643e-08 | 179.6029 | -0.01814 | 0.0057233 | -1.7119e-08 | 174.7247
I5 | A | 15 | 1 | H | -0.026867 | 0.0051063 | -2.4927e-08 | 195.8358 | -0.015736 | 0.0052535 | -7.8474e-09 | 190.3506
I5 | A | 16 | 1 | H | -0.025355 | 0.0055341 | -2.9829e-08 | 180.6983 | -0.014351 | 0.0056747 | -9.3855e-09 | 176.22
I5 | A | 17 | 1 | H | -0.027409 | 0.0053523 | -2.2056e-08 | 186.8365 | -0.015709 | 0.005501 | -1.5067e-08 | 181.7865
I5 | A | 18 | 1 | H | -0.028709 | 0.0055616 | -3.1596e-08 | 179.8032 | -0.011422 | 0.0056717 | -4.1278e-09 | 176.314
I5 | A | 19 | 1 | H | -0.029757 | 0.0053684 | -3.0282e-08 | 186.274 | -0.0093306 | 0.0054716 | 1.9947e-09 | 182.7605
I5 | A | 20 | 1 | H | -0.02788 | 0.0056336 | -3.1624e-08 | 177.5074 | -0.012422 | 0.0057636 | -2.3569e-09 | 173.5022
I5 | A | 21 | 1 | H | -0.022645 | 0.0052942 | -2.2453e-08 | 188.8846 | -0.014594 | 0.0054447 | -9.4896e-09 | 183.6643
I5 | A | 22 | 1 | H | -0.020734 | 0.0058134 | -2.0079e-08 | 172.0166 | -0.0077168 | 0.0059694 | -5.2016e-09 | 167.5219
I5 | A | 23 | 1 | H | -0.023969 | 0.0053886 | -2.3236e-08 | 185.5759 | -0.0092616 | 0.0055224 | -9.4414e-10 | 181.0795
I5 | A | 24 | 1 | H | -0.025048 | 0.0056265 | -7.9471e-09 | 177.7315 | -0.0066928 | 0.0057466 | 3.8073e-08 | 174.0145
I5 | A | 25 | 1 | H | -0.028999 | 0.0055313 | -3.2062e-08 | 180.7888 | -0.0061831 | 0.0056404 | 6.005e-09 | 177.2923
I5 | A | 26 | 1 | H | -0.026329 | 0.005784 | -3.092e-08 | 172.8894 | -0.013017 | 0.0059112 | -7.886e-09 | 169.1697
I5 | A | 27 | 1 | H | -0.023319 | 0.0055017 | -1.9472e-08 | 181.7625 | -0.011516 | 0.0056581 | -2.0016e-09 | 176.7366
I5 | A | 28 | 1 | H | -0.02326 | 0.005803 | -2.7763e-08 | 172.3232 | -0.0065078 | 0.0059293 | 9.4759e-09 | 168.6531
I5 | A | 29 | 1 | H | -0.023353 | 0.0056199 | -3.0224e-08 | 177.938 | -0.015816 | 0.0057723 | -1.3883e-08 | 173.2418
I5 | A | 30 | 1 | H | -0.020955 | 0.0058777 | -2.7787e-08 | 170.1333 | -0.011109 | 0.0060154 | -1.2861e-09 | 166.2395
I5 | A | 31 | 1 | H | -0.021436 | 0.0054408 | 3.8185e-08 | 183.7976 | -0.0086112 | 0.0055555 | 6.5259e-08 | 180.0024
I5 | A | 32 | 1 | H | -0.018557 | 0.0058764 | -1.781e-08 | 170.1713 | -0.010394 | 0.0060212 | -5.2214e-09 | 166.0801
I5 | A | 1 | 2 | H | -0.033068 | 0.005231 | -2.5904e-08 | 191.1698 | -0.016519 | 0.0053255 | -6.9779e-09 | 187.7744
I5 | A | 2 | 2 | H | -0.027662 | 0.0054071 | 1.2614e-08 | 184.9408 | -0.018364 | 0.0055524 | 2.3147e-08 | 180.1022
I5 | A | 3 | 2 | H | -0.030251 | 0.0051918 | -2.7106e-08 | 192.6112 | -0.01718 | 0.0053266 | -1.1992e-08 | 187.7374
I5 | A | 4 | 2 | H | -0.030218 | 0.0052343 | 5.777e-09 | 191.0475 | -0.016663 | 0.005368 | 2.3148e-08 | 186.289
I5 | A | 5 | 2 | H | -0.031249 | 0.0050704 | -2.4514e-08 | 197.2224 | -0.014478 | 0.0051941 | -4.0188e-09 | 192.5252
I5 | A | 6 | 2 | H | -0.030826 | 0.005181 | 1.2902e-09 | 193.0144 | -0.015902 | 0.0052934 | 1.9559e-08 | 188.916
I5 | A | 7 | 2 | H | -0.026629 | 0.0050243 | -1.9137e-08 | 199.0308 | -0.012646 | 0.005136 | -1.8511e-10 | 194.7031
I5 | A | 8 | 2 | H | -0.027593 | 0.005323 | -2.3303e-08 | 187.8651 | -0.012763 | 0.0054244 | -1.1948e-09 | 184.3507
I5 | A | 9 | 2 | H | -0.023002 | 0.0050744 | -1.5661e-08 | 197.0694 | -0.015541 | 0.0052144 | -1.131e-08 | 191.7759
I5 | A | 10 | 2 | H | -0.032362 | 0.0054125 | -3.3383e-08 | 184.759 | -0.015342 | 0.0055331 | -1.412e-08 | 180.7301
I5 | A | 11 | 2 | H | -0.025544 | 0.0049977 | -3.0149e-08 | 200.0921 | -0.018374 | 0.0051335 | -2.0242e-08 | 194.8005
I5 | A | 12 | 2 | H | -0.027874 | 0.0053253 | -5.5835e-09 | 187.7826 | -0.018152 | 0.0054557 | 1.0571e-08 | 183.2939
I5 | A | 13 | 2 | H | -0.024861 | 0.0050754 | -1.7014e-08 | 197.0303 | -0.012556 | 0.0051902 | -5.4021e-09 | 192.67
I5 | A | 14 | 2 | H | -0.018334 | 0.0055648 | -1.5232e-08 | 179.7 | -0.017988 | 0.0057167 | -1.7237e-08 | 174.9248
I5 | A | 15 | 2 | H | -0.025555 | 0.0051085 | -2.3156e-08 | 195.7537 | -0.015649 | 0.005258 | -7.3997e-09 | 190.1877
I5 | A | 16 | 2 | H | -0.023661 | 0.0055232 | -2.7014e-08 | 181.0544 | -0.014028 | 0.005668 | -9.2235e-09 | 176.4304
I5 | A | 17 | 2 | H | -0.027885 | 0.0053598 | -2.3318e-08 | 186.5727 | -0.01539 | 0.0055068 | -1.4989e-08 | 181.5942
I5 | A | 18 | 2 | H | -0.029655 | 0.0055573 | -3.2868e-08 | 179.9442 | -0.010122 | 0.0056626 | -2.8839e-09 | 176.5968
I5 | A | 19 | 2 | H | -0.029943 | 0.0053778 | -3.3389e-08 | 185.9488 | -0.010649 | 0.0054828 | -1.3063e-09 | 182.3898
I5 | A | 20 | 2 | H | -0.029575 | 0.0056319 | -3.4879e-08 | 177.5588 | -0.010659 | 0.0057535 | -1.0272e-09 | 173.8087
I5 | A | 21 | 2 | H | -0.021267 | 0.0052958 | -1.9449e-08 | 188.8291 | -0.015752 | 0.0054558 | -1.2037e-08 | 183.2899
I5 | A | 22 | 2 | H | -0.022289 | 0.0058088 | -2.204e-08 | 172.1525 | -0.0066005 | 0.0059547 | -5.5568e-10 | 167.9353
I5 | A | 23 | 2 | H | -0.027186 | 0.0054029 | -2.6725e-08 | 185.087 | -0.0086456 | 0.0055275 | -2.5761e-10 | 180.9139
I5 | A | 24 | 2 | H | -0.026142 | 0.0056217 | -9.0731e-09 | 177.8813 | -0.008298 | 0.0057439 | 3.5959e-08 | 174.0967
I5 | A | 25 | 2 | H | -0.027984 | 0.005535 | -3.0197e-08 | 180.6674 | -0.0051051 | 0.005645 | 7.5746e-09 | 177.1467
I5 | A | 26 | 2 | H | -0.027126 | 0.0057802 | -3.3577e-08 | 173.0033 | -0.013826 | 0.0059055 | -8.9927e-09 | 169.3326
I5 | A | 27 | 2 | H | -0.025463 | 0.0055135 | -2.1692e-08 | 181.3717 | -0.0096379 | 0.0056617 | 3.101e-10 | 176.6254
I5 | A | 28 | 2 | H | -0.021212 | 0.0057908 | -2.5524e-08 | 172.6881 | -0.0057981 | 0.0059182 | 1.2214e-08 | 168.9714
I5 | A | 29 | 2 | H | -0.024183 | 0.0056304 | -3.248e-08 | 177.6088 | -0.013809 | 0.0057731 | -9.7889e-09 | 173.2184
I5 | A | 30 | 2 | H | -0.021592 | 0.0058704 | -2.7302e-08 | 170.3467 | -0.010499 | 0.0060076 | -3.1709e-09 | 166.456
I5 | A | 31 | 2 | H | -0.022738 | 0.0054547 | 3.3751e-08 | 183.3295 | -0.0093853 | 0.0055626 | 6.6469e-08 | 179.7721
I5 | A | 32 | 2 | H | -0.021209 | 0.0058767 | -2.3679e-08 | 170.1646 | -0.0096375 | 0.0060086 | -3.1427e-09 | 166.4276
I5 | B | 1 | 1 | H | -0.034795 | 0.0052353 | -3.347e-08 | 191.0119 | -0.020219 | 0.0053327 | -1.1765e-08 | 187.5235
I5 | B | 2 | 1 | H | -0.02384 | 0.0054078 | 1.7954e-08 | 184.9197 | -0.014737 | 0.0055538 | 2.8722e-08 | 180.0561
I5 | B | 3 | 1 | H | -0.028132 | 0.0051846 | -2.3936e-08 | 192.8781 | -0.01709 | 0.0053224 | -9.8214e-09 | 187.8854
I5 | B | 4 | 1 | H | -0.027051 | 0.0052332 | 1.2541e-08 | 191.0869 | -0.015558 | 0.005373 | 2.7781e-08 | 186.1163
I5 | B | 5 | 1 | H | -0.03141 | 0.0050702 | -2.7872e-08 | 197.2298 | -0.015021 | 0.0051949 | -7.0333e-09 | 192.4976
I5 | B | 6 | 1 | H | -0.027507 | 0.0051819 | 4.894e-09 | 192.9789 | -0.013711 | 0.0052938 | 2.5438e-08 | 188.8991
I5 | B | 7 | 1 | H | -0.022232 | 0.0050096 | -1.3094e-08 | 199.615 | -0.016631 | 0.0051453 | -4.392e-09 | 194.3533
I5 | B | 8 | 1 | H | -0.026916 | 0.0053334 | -2.4396e-08 | 187.4974 | -0.014389 | 0.0054445 | -7.9682e-09 | 183.6715
I5 | B | 9 | 1 | H | -0.022982 | 0.0050666 | -1.3517e-08 | 197.3697 | -0.010591 | 0.0052015 | -4.4241e-09 | 192.2527
I5 | B | 10 | 1 | H | -0.031692 | 0.005424 | -3.5447e-08 | 184.3668 | -0.011905 | 0.0055359 | -9.7792e-09 | 180.6391
I5 | B | 11 | 1 | H | -0.028645 | 0.0050016 | -3.285e-08 | 199.9357 | -0.01545 | 0.0051286 | -1.8503e-08 | 194.9863
I5 | B | 12 | 1 | H | -0.02457 | 0.0053262 | -1.0577e-09 | 187.7513 | -0.01395 | 0.0054575 | 1.679e-08 | 183.2342
I5 | B | 13 | 1 | H | -0.027661 | 0.0050838 | -2.5741e-08 | 196.7015 | -0.019512 | 0.0052035 | -1.315e-08 | 192.1781
I5 | B | 14 | 1 | H | -0.016901 | 0.0055743 | -1.7091e-08 | 179.396 | -0.014852 | 0.0057195 | -1.0349e-08 | 174.8404
I5 | B | 15 | 1 | H | -0.030191 | 0.0051155 | -2.7194e-08 | 195.4853 | -0.016233 | 0.0052626 | -1.2752e-08 | 190.0193
I5 | B | 16 | 1 | H | -0.023171 | 0.0055331 | -2.7826e-08 | 180.7297 | -0.015942 | 0.005678 | -5.8286e-09 | 176.1171
I5 | B | 17 | 1 | H | -0.026161 | 0.0053557 | -2.3366e-08 | 186.716 | -0.0054522 | 0.0054736 | 9.4867e-09 | 182.6949
I5 | B | 18 | 1 | H | -0.025665 | 0.0055587 | -2.7257e-08 | 179.8987 | -0.011173 | 0.0056742 | -9.7862e-10 | 176.237
I5 | B | 19 | 1 | H | -0.027704 | 0.0053685 | -2.8859e-08 | 186.2721 | -0.0093912 | 0.0054782 | 1.0315e-09 | 182.5411
I5 | B | 20 | 1 | H | -0.024096 | 0.0056273 | -2.7602e-08 | 177.7064 | -0.0097119 | 0.0057656 | -1.1003e-09 | 173.4411
I5 | B | 21 | 1 | H | -0.026921 | 0.0053101 | -2.9085e-08 | 188.3196 | -0.0073616 | 0.0054272 | 4.5093e-09 | 184.2561
I5 | B | 22 | 1 | H | -0.023394 | 0.0058323 | -3.1809e-08 | 171.4583 | -0.0057217 | 0.0059611 | 7.5493e-09 | 167.7538
I5 | B | 23 | 1 | H | -0.026219 | 0.0054025 | -2.7885e-08 | 185.101 | -0.012528 | 0.0055302 | -1.3032e-09 | 180.8258
I5 | B | 24 | 1 | H | -0.029778 | 0.0056377 | -1.3526e-08 | 177.3768 | -0.0079174 | 0.0057552 | 3.603e-08 | 173.7559
I5 | B | 25 | 1 | H | -0.029372 | 0.0055374 | -3.5219e-08 | 180.5898 | -0.0062595 | 0.0056504 | 4.789e-10 | 176.9797
I5 | B | 26 | 1 | H | -0.024692 | 0.0057865 | -3.196e-08 | 172.8169 | -0.012287 | 0.0059153 | -5.2721e-09 | 169.0519
I5 | B | 27 | 1 | H | -0.023244 | 0.0055148 | -2.778e-08 | 181.3294 | -0.012142 | 0.0056662 | -3.5171e-09 | 176.4861
I5 | B | 28 | 1 | H | -0.021959 | 0.0057978 | -2.035e-08 | 172.4782 | -0.0040883 | 0.0059285 | 1.0088e-08 | 168.6771
I5 | B | 29 | 1 | H | -0.029278 | 0.0056343 | -3.562e-08 | 177.4849 | -0.010521 | 0.0057593 | -2.3773e-09 | 173.632
I5 | B | 30 | 1 | H | -0.027351 | 0.0058979 | -3.564e-08 | 169.5515 | -0.014114 | 0.0060382 | -1.365e-08 | 165.6135
I5 | B | 31 | 1 | H | -0.021542 | 0.0054548 | 3.0088e-08 | 183.3241 | -0.011422 | 0.0055692 | 6.0976e-08 | 179.5591
I5 | B | 32 | 1 | H | -0.023945 | 0.005905 | -3.549e-08 | 169.3473 | -0.0092394 | 0.0060198 | 5.4294e-10 | 166.1187
I5 | B | 1 | 2 | H | -0.033017 | 0.0052388 | -3.2652e-08 | 190.8836 | -0.018208 | 0.005334 | -8.4059e-09 | 187.4767
I5 | B | 2 | 2 | H | -0.023259 | 0.0054001 | 1.8417e-08 | 185.183 | -0.015748 | 0.0055498 | 2.7433e-08 | 180.1876
I5 | B | 3 | 2 | H | -0.025053 | 0.0051817 | -1.8754e-08 | 192.988 | -0.015291 | 0.0053242 | -7.4669e-09 | 187.8219
I5 | B | 4 | 2 | H | -0.027067 | 0.0052287 | 1.1163e-08 | 191.2511 | -0.015075 | 0.0053675 | 2.6785e-08 | 186.3048
I5 | B | 5 | 2 | H | -0.030983 | 0.0050759 | -2.7702e-08 | 197.0087 | -0.016982 | 0.0052077 | -1.0977e-08 | 192.0233
I5 | B | 6 | 2 | H | -0.029126 | 0.0051788 | 3.9321e-09 | 193.0956 | -0.013606 | 0.0052877 | 2.5059e-08 | 189.1172
I5 | B | 7 | 2 | H | -0.023899 | 0.0050215 | -1.6596e-08 | 199.1425 | -0.014785 | 0.0051485 | -3.6841e-09 | 194.2314
I5 | B | 8 | 2 | H | -0.029073 | 0.0053333 | -2.8018e-08 | 187.4997 | -0.013544 | 0.0054385 | -8.9911e-09 | 183.873
I5 | B | 9 | 2 | H | -0.022495 | 0.0050728 | -1.4364e-08 | 197.1311 | -0.012979 | 0.0052134 | -7.4542e-09 | 191.8148
I5 | B | 10 | 2 | H | -0.030705 | 0.0054139 | -3.2755e-08 | 184.7094 | -0.013059 | 0.0055343 | -1.236e-08 | 180.6922
I5 | B | 11 | 2 | H | -0.027191 | 0.005006 | -3.3096e-08 | 199.7588 | -0.014972 | 0.0051324 | -1.7316e-08 | 194.8405
I5 | B | 12 | 2 | H | -0.025478 | 0.0053255 | -4.4824e-09 | 187.7748 | -0.014452 | 0.005453 | 1.6688e-08 | 183.3842
I5 | B | 13 | 2 | H | -0.029257 | 0.0050929 | -2.7596e-08 | 196.3509 | -0.017074 | 0.0052047 | -1.1797e-08 | 192.1353
I5 | B | 14 | 2 | H | -0.020045 | 0.0055712 | -1.7391e-08 | 179.4957 | -0.011987 | 0.0057071 | -7.0537e-09 | 175.2206
I5 | B | 15 | 2 | H | -0.026786 | 0.0051099 | -2.0939e-08 | 195.6996 | -0.015727 | 0.0052646 | -1.0121e-08 | 189.9491
I5 | B | 16 | 2 | H | -0.02365 | 0.005528 | -2.8201e-08 | 180.8984 | -0.01667 | 0.0056723 | -5.9318e-09 | 176.2957
I5 | B | 17 | 2 | H | -0.026444 | 0.0053638 | -2.5417e-08 | 186.4341 | -0.0047736 | 0.0054776 | 1.0682e-08 | 182.5606
I5 | B | 18 | 2 | H | -0.026378 | 0.005553 | -2.687e-08 | 180.0817 | -0.012074 | 0.0056715 | -3.474e-09 | 176.3209
I5 | B | 19 | 2 | H | -0.026961 | 0.0053729 | -2.8325e-08 | 186.119 | -0.0098886 | 0.0054835 | 2.9127e-09 | 182.3662
I5 | B | 20 | 2 | H | -0.024308 | 0.0056184 | -2.5463e-08 | 177.9851 | -0.0096019 | 0.0057577 | -4.5093e-10 | 173.6805
I5 | B | 21 | 2 | H | -0.025658 | 0.0053099 | -2.4395e-08 | 188.3281 | -0.0080826 | 0.0054335 | 5.0103e-09 | 184.0431
I5 | B | 22 | 2 | H | -0.024757 | 0.00583 | -3.5301e-08 | 171.5276 | -0.0064355 | 0.0059538 | 6.6459e-09 | 167.9598
I5 | B | 23 | 2 | H | -0.025403 | 0.0054061 | -2.6106e-08 | 184.9769 | -0.011313 | 0.0055317 | 2.3498e-09 | 180.7777
I5 | B | 24 | 2 | H | -0.029524 | 0.0056305 | -1.3532e-08 | 177.6056 | -0.0085637 | 0.0057482 | 3.5952e-08 | 173.9675
I5 | B | 25 | 2 | H | -0.028591 | 0.0055413 | -3.3392e-08 | 180.4618 | -0.0080283 | 0.005662 | -2.0692e-09 | 176.617
I5 | B | 26 | 2 | H | -0.026085 | 0.0057837 | -3.4728e-08 | 172.8984 | -0.012352 | 0.005908 | -5.0679e-09 | 169.2608
I5 | B | 27 | 2 | H | -0.025223 | 0.0055257 | -2.9912e-08 | 180.9739 | -0.010682 | 0.0056686 | -8.1478e-10 | 176.4109
I5 | B | 28 | 2 | H | -0.024556 | 0.0057955 | -2.2816e-08 | 172.5466 | -0.004214 | 0.0059204 | 1.0675e-08 | 168.9064
I5 | B | 29 | 2 | H | -0.029371 | 0.0056425 | -3.6319e-08 | 177.2271 | -0.011274 | 0.0057714 | -5.3462e-09 | 173.2691
I5 | B | 30 | 2 | H | -0.024908 | 0.0058811 | -2.9248e-08 | 170.0367 | -0.012203 | 0.006023 | -8.7566e-09 | 166.0297
I5 | B | 31 | 2 | H | -0.02292 | 0.005466 | 2.8451e-08 | 182.9486 | -0.0094768 | 0.0055704 | 6.516e-08 | 179.5214
I5 | B | 32 | 2 | H | -0.023032 | 0.0058932 | -3.3297e-08 | 169.6883 | -0.0081534 | 0.0060078 | 2.0399e-09 | 166.45
M12 | A | 1 | 1 | H | -0.0025407 | 0.00087616 | -5.0631e-09 | 1141.3472 | -0.0038719 | 0.00088847 | -9.468e-09 | 1125.5251
M12 | A | 2 | 1 | H | -0.002514 | 0.00085925 | -4.9863e-09 | 1163.8089 | -0.0037835 | 0.00087248 | -9.1927e-09 | 1146.1554
M12 | A | 3 | 1 | H | -0.0025628 | 0.00086257 | -5.0006e-09 | 1159.3274 | -0.0041017 | 0.0008759 | -9.4551e-09 | 1141.6778
M12 | A | 4 | 1 | H | -0.0025054 | 0.00086267 | -5.051e-09 | 1159.1874 | -0.0039328 | 0.0008752 | -9.2521e-09 | 1142.5953
M12 | A | 5 | 1 | H | -0.002441 | 0.00086141 | -4.7554e-09 | 1160.8916 | -0.0038835 | 0.00087432 | -8.9849e-09 | 1143.7407
M12 | A | 6 | 1 | H | -0.0023607 | 0.00086689 | -5.0403e-09 | 1153.5442 | -0.0038002 | 0.00087916 | -9.2008e-09 | 1137.4486
M12 | A | 7 | 1 | H | -0.0024183 | 0.00086396 | -4.911e-09 | 1157.4558 | -0.003912 | 0.00087603 | -9.0964e-09 | 1141.5071
M12 | A | 8 | 1 | H | -0.0023778 | 0.00086678 | -5.0536e-09 | 1153.6984 | -0.0035553 | 0.00087734 | -8.8014e-09 | 1139.8097
M12 | A | 9 | 1 | H | -0.0024232 | 0.00086464 | -5.0586e-09 | 1156.5491 | -0.0036386 | 0.00087553 | -8.6885e-09 | 1142.1644
M12 | A | 10 | 1 | H | -0.0022616 | 0.00086781 | -5.001e-09 | 1152.3263 | -0.003531 | 0.00087832 | -8.6658e-09 | 1138.5388
M12 | A | 11 | 1 | H | -0.0023679 | 0.0008648 | -4.9749e-09 | 1156.3308 | -0.0035934 | 0.00087558 | -8.5648e-09 | 1142.0938
M12 | A | 12 | 1 | H | -0.0022261 | 0.00086798 | -5.0671e-09 | 1152.1033 | -0.0035067 | 0.00087825 | -8.5941e-09 | 1138.6295
M12 | A | 13 | 1 | H | -0.0023345 | 0.00086652 | -4.988e-09 | 1154.0439 | -0.0034024 | 0.00087661 | -8.2575e-09 | 1140.7614
M12 | A | 14 | 1 | H | -0.0022351 | 0.00087123 | -5.1718e-09 | 1147.8009 | -0.0033059 | 0.00088032 | -8.1788e-09 | 1135.9466
M12 | A | 15 | 1 | H | -0.0022705 | 0.00086824 | -5.0298e-09 | 1151.7534 | -0.0033818 | 0.00087778 | -8.1116e-09 | 1139.2321
M12 | A | 16 | 1 | H | -0.0022004 | 0.00087332 | -5.1092e-09 | 1145.0535 | -0.0030875 | 0.00088167 | -7.8547e-09 | 1134.2051
M12 | B | 1 | 1 | H | -0.0025258 | 0.00087634 | -5.0525e-09 | 1141.1057 | -0.0038645 | 0.00088888 | -9.4182e-09 | 1125.0163
M12 | B | 2 | 1 | H | -0.0025658 | 0.00085948 | -5.013e-09 | 1163.4984 | -0.0037353 | 0.00087271 | -9.0941e-09 | 1145.8596
M12 | B | 3 | 1 | H | -0.0025262 | 0.00086269 | -5.0087e-09 | 1159.1684 | -0.0040361 | 0.00087608 | -9.3121e-09 | 1141.446
M12 | B | 4 | 1 | H | -0.0024959 | 0.00086287 | -5.0761e-09 | 1158.9224 | -0.0039253 | 0.00087555 | -9.1725e-09 | 1142.1432
M12 | B | 5 | 1 | H | -0.0024763 | 0.0008616 | -4.7911e-09 | 1160.6265 | -0.0037857 | 0.00087442 | -8.8181e-09 | 1143.6204
M12 | B | 6 | 1 | H | -0.0024122 | 0.00086729 | -5.1274e-09 | 1153.0224 | -0.0037459 | 0.00087942 | -9.094e-09 | 1137.1091
M12 | B | 7 | 1 | H | -0.0023958 | 0.00086415 | -4.9505e-09 | 1157.2071 | -0.0038867 | 0.00087642 | -9.0271e-09 | 1141.0088
M12 | B | 8 | 1 | H | -0.0024434 | 0.00086717 | -5.1364e-09 | 1153.1787 | -0.0034915 | 0.00087759 | -8.6775e-09 | 1139.4835
M12 | B | 9 | 1 | H | -0.0024327 | 0.00086499 | -5.142e-09 | 1156.08 | -0.003534 | 0.00087576 | -8.5879e-09 | 1141.8711
M12 | B | 10 | 1 | H | -0.0023283 | 0.00086817 | -5.0705e-09 | 1151.8451 | -0.0034466 | 0.00087855 | -8.55e-09 | 1138.2396
M12 | B | 11 | 1 | H | -0.0023934 | 0.00086507 | -5.0459e-09 | 1155.9804 | -0.0035011 | 0.00087571 | -8.4112e-09 | 1141.9355
M12 | B | 12 | 1 | H | -0.002234 | 0.00086818 | -5.0914e-09 | 1151.8357 | -0.0034238 | 0.00087848 | -8.477e-09 | 1138.3303
M12 | B | 13 | 1 | H | -0.0023283 | 0.00086668 | -5.0211e-09 | 1153.8248 | -0.0033639 | 0.00087684 | -8.1634e-09 | 1140.4654
M12 | B | 14 | 1 | H | -0.0022537 | 0.00087144 | -5.1897e-09 | 1147.5237 | -0.0032335 | 0.00088048 | -8.0258e-09 | 1135.7404
M12 | B | 15 | 1 | H | -0.0023302 | 0.00086859 | -5.0988e-09 | 1151.2926 | -0.0033341 | 0.00087815 | -8.0362e-09 | 1138.758
M12 | B | 16 | 1 | H | -0.0022916 | 0.00087371 | -5.1954e-09 | 1144.5442 | -0.0030571 | 0.00088207 | -7.8104e-09 | 1133.7023
M13 | A | 1 | 1 | H | -0.0051721 | 0.0016076 | -1.1087e-08 | 622.028 | -0.0082583 | 0.0016315 | -2.0817e-08 | 612.9367
M13 | A | 2 | 1 | H | -0.0050698 | 0.0016644 | -1.1782e-08 | 600.8216 | -0.0075973 | 0.0016888 | -2.1369e-08 | 592.1261
M13 | A | 3 | 1 | H | -0.0051225 | 0.0015977 | -1.0716e-08 | 625.9021 | -0.007926 | 0.0016214 | -1.9697e-08 | 616.7616
M13 | A | 4 | 1 | H | -0.0054072 | 0.001666 | -1.1803e-08 | 600.2415 | -0.0078955 | 0.0016905 | -2.1286e-08 | 591.5476
M13 | A | 5 | 1 | H | -0.0053463 | 0.0015963 | -1.0888e-08 | 626.4449 | -0.007892 | 0.001619 | -1.8981e-08 | 617.6693
M13 | A | 6 | 1 | H | -0.0053796 | 0.0016795 | -1.1943e-08 | 595.413 | -0.0076748 | 0.0017022 | -2.089e-08 | 587.4853
M13 | A | 7 | 1 | H | -0.0052801 | 0.0016093 | -1.0856e-08 | 621.398 | -0.0078726 | 0.0016309 | -1.892e-08 | 613.1457
M13 | A | 8 | 1 | H | -0.0054055 | 0.001681 | -1.2114e-08 | 594.8979 | -0.0075816 | 0.001702 | -2.051e-08 | 587.558
M13 | A | 9 | 1 | H | -0.0053845 | 0.0016149 | -1.1454e-08 | 619.2449 | -0.0079304 | 0.0016358 | -1.9112e-08 | 611.3262
M13 | A | 10 | 1 | H | -0.0053308 | 0.0016783 | -1.2179e-08 | 595.8389 | -0.0073977 | 0.0016979 | -1.9928e-08 | 588.9457
M13 | A | 11 | 1 | H | -0.0051258 | 0.0016164 | -1.125e-08 | 618.6542 | -0.0073785 | 0.0016352 | -1.8074e-08 | 611.5322
M13 | A | 12 | 1 | H | -0.0050576 | 0.0016832 | -1.1805e-08 | 594.1029 | -0.0069387 | 0.0017016 | -1.8693e-08 | 587.6732
M13 | A | 13 | 1 | H | -0.0050667 | 0.0016163 | -1.1195e-08 | 618.6829 | -0.0070946 | 0.0016338 | -1.7474e-08 | 612.0578
M13 | A | 14 | 1 | H | -0.0050863 | 0.0016906 | -1.2236e-08 | 591.5149 | -0.0067465 | 0.0017082 | -1.8786e-08 | 585.4004
M13 | A | 15 | 1 | H | -0.0049753 | 0.001631 | -1.1773e-08 | 613.1268 | -0.0068001 | 0.0016466 | -1.7392e-08 | 607.3177
M13 | A | 16 | 1 | H | -0.0051753 | 0.0017133 | -1.2892e-08 | 583.6826 | -0.0067564 | 0.0017299 | -1.8979e-08 | 578.0579
M13 | B | 1 | 1 | H | -0.0052384 | 0.0016075 | -1.1189e-08 | 622.073 | -0.0081908 | 0.0016317 | -2.0716e-08 | 612.8612
M13 | B | 2 | 1 | H | -0.0051655 | 0.0016644 | -1.1797e-08 | 600.811 | -0.0075958 | 0.0016891 | -2.1213e-08 | 592.0223
M13 | B | 3 | 1 | H | -0.0051738 | 0.0015976 | -1.0813e-08 | 625.9214 | -0.0079059 | 0.0016215 | -1.9491e-08 | 616.6973
M13 | B | 4 | 1 | H | -0.0054711 | 0.0016659 | -1.1798e-08 | 600.2795 | -0.0077724 | 0.0016904 | -2.0969e-08 | 591.5589
M13 | B | 5 | 1 | H | -0.0054368 | 0.0015964 | -1.1069e-08 | 626.4184 | -0.0078953 | 0.0016191 | -1.8778e-08 | 617.6167
M13 | B | 6 | 1 | H | -0.005395 | 0.0016795 | -1.2007e-08 | 595.4082 | -0.0076216 | 0.0017022 | -2.0594e-08 | 587.4586
M13 | B | 7 | 1 | H | -0.0052047 | 0.001609 | -1.0964e-08 | 621.4848 | -0.0079849 | 0.0016313 | -1.882e-08 | 613.0139
M13 | B | 8 | 1 | H | -0.0054389 | 0.001681 | -1.2175e-08 | 594.8758 | -0.0074106 | 0.001702 | -2.0206e-08 | 587.5538
M13 | B | 9 | 1 | H | -0.0054434 | 0.0016149 | -1.1613e-08 | 619.2158 | -0.0078701 | 0.001636 | -1.8933e-08 | 611.2402
M13 | B | 10 | 1 | H | -0.0054359 | 0.0016787 | -1.2404e-08 | 595.7058 | -0.007265 | 0.001698 | -1.9644e-08 | 588.9173
M13 | B | 11 | 1 | H | -0.0051066 | 0.0016164 | -1.1343e-08 | 618.6667 | -0.0072023 | 0.0016352 | -1.7762e-08 | 611.5544
M13 | B | 12 | 1 | H | -0.0050924 | 0.0016834 | -1.1969e-08 | 594.0318 | -0.0069559 | 0.001702 | -1.8535e-08 | 587.5577
M13 | B | 13 | 1 | H | -0.0050994 | 0.0016162 | -1.1284e-08 | 618.7239 | -0.0070367 | 0.0016339 | -1.7211e-08 | 612.0201
M13 | B | 14 | 1 | H | -0.0051967 | 0.0016908 | -1.2437e-08 | 591.4272 | -0.0068051 | 0.0017088 | -1.877e-08 | 585.1943
M13 | B | 15 | 1 | H | -0.0050406 | 0.0016308 | -1.1796e-08 | 613.2076 | -0.0068145 | 0.0016467 | -1.7138e-08 | 607.2776
M13 | B | 16 | 1 | H | -0.0052065 | 0.0017133 | -1.2964e-08 | 583.6699 | -0.0067451 | 0.0017303 | -1.8819e-08 | 577.9364
M14 | A | 1 | 1 | H | -0.037999 | 0.0052725 | 1.8468e-08 | 189.6618 | -0.024375 | 0.0054725 | 3.2804e-08 | 182.731
M14 | A | 2 | 1 | H | -0.03797 | 0.0050982 | 1.2382e-08 | 196.1469 | -0.026676 | 0.0052794 | 2.2147e-08 | 189.4166
M14 | A | 3 | 1 | H | -0.037999 | 0.0051967 | 1.4917e-08 | 192.4313 | -0.024239 | 0.0053933 | 3.0251e-08 | 185.4147
M14 | A | 4 | 1 | H | -0.036415 | 0.0054463 | 9.8043e-09 | 183.6106 | -0.026213 | 0.0056635 | 2.0591e-08 | 176.57
M14 | A | 5 | 1 | H | -0.037768 | 0.0051894 | 1.5123e-08 | 192.7015 | -0.023565 | 0.0053787 | 3.0585e-08 | 185.9193
M14 | A | 6 | 1 | H | -0.038429 | 0.0052018 | 1.5166e-08 | 192.2411 | -0.022963 | 0.0053904 | 3.2831e-08 | 185.5143
M14 | A | 7 | 1 | H | -0.037295 | 0.0052974 | 2.123e-08 | 188.7727 | -0.020301 | 0.005497 | 4.3817e-08 | 181.9161
M14 | A | 8 | 1 | H | -0.037042 | 0.0052165 | 1.5946e-08 | 191.6992 | -0.022792 | 0.0054006 | 3.375e-08 | 185.1639
M14 | A | 9 | 1 | H | -0.036992 | 0.005284 | 1.7733e-08 | 189.2503 | -0.020481 | 0.0054716 | 4.0906e-08 | 182.7616
M14 | A | 10 | 1 | H | -0.037 | 0.0051938 | 1.4662e-08 | 192.537 | -0.021256 | 0.0053711 | 3.3789e-08 | 186.1799
M14 | A | 11 | 1 | H | -0.036933 | 0.0054315 | 2.1252e-08 | 184.1107 | -0.018127 | 0.0056306 | 4.9586e-08 | 177.6012
M14 | A | 12 | 1 | H | -0.036665 | 0.0052929 | 1.7147e-08 | 188.9309 | -0.019639 | 0.0054749 | 4.1483e-08 | 182.6502
M14 | A | 13 | 1 | H | -0.036038 | 0.0052169 | 1.4715e-08 | 191.6835 | -0.019224 | 0.0053927 | 3.5435e-08 | 185.4345
M14 | A | 14 | 1 | H | -0.036873 | 0.0051953 | 1.1974e-08 | 192.4826 | -0.020117 | 0.0053587 | 3.2501e-08 | 186.6126
M14 | A | 15 | 1 | H | -0.036648 | 0.0053086 | 1.4975e-08 | 188.3745 | -0.018283 | 0.0054834 | 4.0928e-08 | 182.3672
M14 | A | 16 | 1 | H | -0.036528 | 0.0054343 | 1.4391e-08 | 184.0157 | -0.018259 | 0.0056487 | 4.1384e-08 | 177.0307
M14 | B | 1 | 1 | H | -0.037863 | 0.0052602 | 1.7858e-08 | 190.1083 | -0.023971 | 0.0054596 | 3.2915e-08 | 183.1643
M14 | B | 2 | 1 | H | -0.03774 | 0.0050849 | 1.2135e-08 | 196.6594 | -0.025929 | 0.0052653 | 2.2647e-08 | 189.9228
M14 | B | 3 | 1 | H | -0.037504 | 0.0051818 | 1.5488e-08 | 192.9837 | -0.02362 | 0.00538 | 3.0396e-08 | 185.8735
M14 | B | 4 | 1 | H | -0.037315 | 0.0054355 | 8.2832e-09 | 183.9743 | -0.025138 | 0.0056486 | 2.1019e-08 | 177.0365
M14 | B | 5 | 1 | H | -0.037525 | 0.005176 | 1.4986e-08 | 193.2012 | -0.022516 | 0.0053646 | 3.0927e-08 | 186.4072
M14 | B | 6 | 1 | H | -0.038036 | 0.0051884 | 1.5101e-08 | 192.737 | -0.022201 | 0.0053762 | 3.3586e-08 | 186.0044
M14 | B | 7 | 1 | H | -0.036822 | 0.0052843 | 2.0835e-08 | 189.2407 | -0.019209 | 0.0054823 | 4.4684e-08 | 182.4064
M14 | B | 8 | 1 | H | -0.037007 | 0.0052044 | 1.5241e-08 | 192.1451 | -0.022559 | 0.0053879 | 3.3959e-08 | 185.6008
M14 | B | 9 | 1 | H | -0.037297 | 0.0052716 | 1.7314e-08 | 189.6939 | -0.020247 | 0.0054587 | 4.0917e-08 | 183.1937
M14 | B | 10 | 1 | H | -0.037 | 0.0051815 | 1.4257e-08 | 192.9959 | -0.020252 | 0.0053565 | 3.4815e-08 | 186.6892
M14 | B | 11 | 1 | H | -0.036533 | 0.0054174 | 2.1318e-08 | 184.5911 | -0.01645 | 0.0056141 | 5.1026e-08 | 178.1217
M14 | B | 12 | 1 | H | -0.037053 | 0.0052811 | 1.6567e-08 | 189.3553 | -0.018265 | 0.0054604 | 4.213e-08 | 183.1351
M14 | B | 13 | 1 | H | -0.035727 | 0.005204 | 1.4091e-08 | 192.1603 | -0.018502 | 0.0053781 | 3.6193e-08 | 185.9401
M14 | B | 14 | 1 | H | -0.036753 | 0.0051816 | 1.1796e-08 | 192.9914 | -0.019717 | 0.0053454 | 3.2399e-08 | 187.0769
M14 | B | 15 | 1 | H | -0.036587 | 0.0052953 | 1.4496e-08 | 188.8477 | -0.0168 | 0.0054671 | 4.2198e-08 | 182.9108
M14 | B | 16 | 1 | H | -0.036597 | 0.0054204 | 1.3754e-08 | 184.4872 | -0.019233 | 0.0056383 | 4.001e-08 | 177.3597
M15 | A | 1 | 1 | H | -0.028175 | 0.0051723 | -1.6591e-08 | 193.337 | -0.030188 | 0.0053148 | -3.0211e-08 | 188.155
M15 | A | 2 | 1 | H | -0.027996 | 0.0052357 | -1.3877e-08 | 190.9971 | -0.028682 | 0.0053749 | -2.6174e-08 | 186.0487
M15 | A | 3 | 1 | H | -0.026254 | 0.0051364 | -1.2322e-08 | 194.6882 | -0.02908 | 0.005275 | -2.3363e-08 | 189.575
M15 | A | 4 | 1 | H | -0.025562 | 0.0052452 | -1.0804e-08 | 190.649 | -0.027539 | 0.0053874 | -2.1691e-08 | 185.6173
M15 | A | 5 | 1 | H | -0.026065 | 0.0051281 | -9.6553e-09 | 195.0056 | -0.025133 | 0.0052631 | -1.6568e-08 | 190.003
M15 | A | 6 | 1 | H | -0.026474 | 0.0051626 | -1.1318e-08 | 193.7019 | -0.025477 | 0.0052907 | -1.7244e-08 | 189.0096
M15 | A | 7 | 1 | H | -0.026614 | 0.005139 | -1.1221e-08 | 194.5905 | -0.024943 | 0.0052672 | -1.6523e-08 | 189.8557
M15 | A | 8 | 1 | H | -0.025661 | 0.0052169 | -1.0971e-08 | 191.6864 | -0.022631 | 0.0053398 | -1.2706e-08 | 187.2715
M15 | A | 9 | 1 | H | -0.026197 | 0.0051595 | -1.0854e-08 | 193.8163 | -0.023203 | 0.0052838 | -1.2024e-08 | 189.2571
M15 | A | 10 | 1 | H | -0.025296 | 0.0052219 | -1.0309e-08 | 191.5022 | -0.022761 | 0.005346 | -1.1839e-08 | 187.0555
M15 | A | 11 | 1 | H | -0.025112 | 0.0051608 | -1.1218e-08 | 193.7684 | -0.022245 | 0.0052806 | -1.2972e-08 | 189.3712
M15 | A | 12 | 1 | H | -0.025583 | 0.00525 | -1.2066e-08 | 190.4749 | -0.022692 | 0.0053703 | -1.1245e-08 | 186.2109
M15 | A | 13 | 1 | H | -0.026186 | 0.0051459 | -1.1745e-08 | 194.3288 | -0.022463 | 0.0052671 | -1.2335e-08 | 189.8561
M15 | A | 14 | 1 | H | -0.02491 | 0.0052291 | -1.3352e-08 | 191.2358 | -0.022079 | 0.0053449 | -1.1223e-08 | 187.093
M15 | A | 15 | 1 | H | -0.025234 | 0.0051782 | -1.2102e-08 | 193.117 | -0.020645 | 0.0052979 | -9.9522e-09 | 188.7531
M15 | A | 16 | 1 | H | -0.025158 | 0.0053519 | -1.222e-08 | 186.8483 | -0.022652 | 0.0054763 | -1.1228e-08 | 182.6063
M15 | B | 1 | 1 | H | -0.028024 | 0.0051735 | -1.7343e-08 | 193.2931 | -0.029438 | 0.0053143 | -2.8966e-08 | 188.1722
M15 | B | 2 | 1 | H | -0.028162 | 0.0052372 | -1.4907e-08 | 190.9407 | -0.027155 | 0.0053721 | -2.3384e-08 | 186.1471
M15 | B | 3 | 1 | H | -0.026065 | 0.0051358 | -1.2036e-08 | 194.7115 | -0.02847 | 0.0052749 | -2.2762e-08 | 189.5778
M15 | B | 4 | 1 | H | -0.025645 | 0.0052459 | -1.1189e-08 | 190.6241 | -0.026644 | 0.0053875 | -2.0836e-08 | 185.616
M15 | B | 5 | 1 | H | -0.026269 | 0.005129 | -1.0127e-08 | 194.9685 | -0.025072 | 0.0052644 | -1.6415e-08 | 189.9536
M15 | B | 6 | 1 | H | -0.0262 | 0.0051627 | -1.1149e-08 | 193.6979 | -0.025749 | 0.0052943 | -1.7946e-08 | 188.8815
M15 | B | 7 | 1 | H | -0.026609 | 0.0051392 | -1.1286e-08 | 194.5824 | -0.02446 | 0.0052672 | -1.5805e-08 | 189.8546
M15 | B | 8 | 1 | H | -0.02545 | 0.0052179 | -1.1678e-08 | 191.6468 | -0.021427 | 0.0053397 | -1.1628e-08 | 187.2776
M15 | B | 9 | 1 | H | -0.026181 | 0.0051603 | -1.1127e-08 | 193.789 | -0.022619 | 0.0052841 | -1.1136e-08 | 189.2473
M15 | B | 10 | 1 | H | -0.024983 | 0.0052213 | -9.9039e-09 | 191.5216 | -0.021555 | 0.0053448 | -1.0224e-08 | 187.0973
M15 | B | 11 | 1 | H | -0.024983 | 0.0051614 | -1.1691e-08 | 193.7472 | -0.02181 | 0.0052813 | -1.2611e-08 | 189.347
M15 | B | 12 | 1 | H | -0.02579 | 0.0052513 | -1.2891e-08 | 190.43 | -0.022134 | 0.0053708 | -1.0878e-08 | 186.1927
M15 | B | 13 | 1 | H | -0.026736 | 0.0051482 | -1.2905e-08 | 194.2409 | -0.021966 | 0.0052679 | -1.1654e-08 | 189.8295
M15 | B | 14 | 1 | H | -0.024911 | 0.0052303 | -1.3795e-08 | 191.1938 | -0.022466 | 0.0053491 | -1.278e-08 | 186.9468
M15 | B | 15 | 1 | H | -0.02528 | 0.0051806 | -1.2728e-08 | 193.0293 | -0.020051 | 0.0052998 | -9.3493e-09 | 188.6872
M15 | B | 16 | 1 | H | -0.025897 | 0.0053543 | -1.3233e-08 | 186.7652 | -0.022126 | 0.0054774 | -1.0746e-08 | 182.5673
M16A | A | 1 | 1 | H | -0.027833 | 0.0049566 | -1.0167e-08 | 201.7521 | -0.021098 | 0.0050739 | -5.8004e-09 | 197.0881
M16A | A | 2 | 1 | H | -0.028852 | 0.0048421 | -1.2047e-08 | 206.522 | -0.018355 | 0.00497 | -2.0068e-09 | 201.2075
M16A | A | 3 | 1 | H | -0.030552 | 0.0047977 | -1.4878e-08 | 208.4339 | -0.02096 | 0.0049041 | -5.4514e-09 | 203.9124
M16A | A | 4 | 1 | H | -0.027679 | 0.0048457 | -1.033e-08 | 206.3701 | -0.018546 | 0.0049599 | 8.0686e-10 | 201.6169
M16A | A | 5 | 1 | H | -0.027676 | 0.0048702 | -1.3102e-08 | 205.3304 | -0.016033 | 0.0049704 | -3.6935e-10 | 201.1929
M16A | A | 6 | 1 | H | -0.025711 | 0.004867 | -9.0596e-09 | 205.4651 | -0.015762 | 0.0049799 | 5.2464e-09 | 200.8074
M16A | A | 7 | 1 | H | -0.025205 | 0.0048854 | -8.7234e-09 | 204.6899 | -0.016983 | 0.0049936 | 1.3342e-09 | 200.2564
M16A | A | 8 | 1 | H | -0.02794 | 0.0049237 | -1.2881e-08 | 203.1012 | -0.017281 | 0.0050125 | 2.4761e-09 | 199.5002
M16A | A | 9 | 1 | H | -0.027157 | 0.0049256 | -1.4568e-08 | 203.0196 | -0.015761 | 0.0050279 | 2.1563e-09 | 198.889
M16A | A | 10 | 1 | H | -0.02541 | 0.0049762 | -1.284e-08 | 200.9578 | -0.014495 | 0.0050879 | 7.0269e-09 | 196.5463
M16A | A | 11 | 1 | H | -0.028327 | 0.0049503 | -1.5533e-08 | 202.0097 | -0.012893 | 0.0050383 | 6.9817e-09 | 198.4778
M16A | A | 12 | 1 | H | -0.025833 | 0.0050367 | -1.3314e-08 | 198.5411 | -0.010815 | 0.0051301 | 1.5127e-08 | 194.9271
M16A | A | 13 | 1 | H | -0.023974 | 0.0050081 | -1.1156e-08 | 199.6753 | -0.011629 | 0.0050997 | 1.3046e-08 | 196.0904
M16A | A | 14 | 1 | H | -0.026153 | 0.0049629 | -1.6783e-08 | 201.4954 | -0.01192 | 0.0050603 | 8.7508e-09 | 197.6174
M16A | A | 15 | 1 | H | -0.02538 | 0.0051131 | -1.5659e-08 | 195.5749 | -0.011425 | 0.0051938 | 1.0984e-08 | 192.5384
M16A | A | 16 | 1 | H | -0.026026 | 0.0051281 | -2.1415e-08 | 195.0053 | -0.011517 | 0.0052099 | 7.5611e-09 | 191.9409
M16A | B | 1 | 1 | H | -0.028045 | 0.004961 | -1.0151e-08 | 201.5721 | -0.01911 | 0.0050751 | -3.7019e-09 | 197.0396
M16A | B | 2 | 1 | H | -0.028986 | 0.0048462 | -1.2263e-08 | 206.3453 | -0.017869 | 0.0049741 | -1.2222e-09 | 201.0398
M16A | B | 3 | 1 | H | -0.02842 | 0.0047977 | -1.2973e-08 | 208.434 | -0.018005 | 0.0049024 | -9.8244e-10 | 203.9832
M16A | B | 4 | 1 | H | -0.027975 | 0.0048505 | -1.063e-08 | 206.1649 | -0.017021 | 0.0049638 | 1.8965e-09 | 201.4603
M16A | B | 5 | 1 | H | -0.028061 | 0.0048761 | -1.4214e-08 | 205.0805 | -0.016239 | 0.0049741 | 1.7804e-09 | 201.0408
M16A | B | 6 | 1 | H | -0.029044 | 0.0048799 | -1.3578e-08 | 204.9208 | -0.015263 | 0.0049859 | 4.7818e-09 | 200.5671
M16A | B | 7 | 1 | H | -0.026749 | 0.004895 | -1.186e-08 | 204.2899 | -0.016733 | 0.0049991 | 1.9324e-09 | 200.0372
M16A | B | 8 | 1 | H | -0.026687 | 0.0049264 | -1.2247e-08 | 202.9866 | -0.015974 | 0.0050145 | 5.0687e-09 | 199.4197
M16A | B | 9 | 1 | H | -0.026679 | 0.0049291 | -1.453e-08 | 202.8766 | -0.015392 | 0.005031 | 4.7545e-09 | 198.7665
M16A | B | 10 | 1 | H | -0.024884 | 0.0049804 | -1.2594e-08 | 200.7885 | -0.013832 | 0.0050914 | 9.062e-09 | 196.408
M16A | B | 11 | 1 | H | -0.027707 | 0.0049521 | -1.3537e-08 | 201.9359 | -0.011018 | 0.0050378 | 1.1486e-08 | 198.4981
M16A | B | 12 | 1 | H | -0.026613 | 0.0050434 | -1.4492e-08 | 198.2772 | -0.01137 | 0.0051377 | 1.4677e-08 | 194.6383
M16A | B | 13 | 1 | H | -0.024961 | 0.005014 | -1.1636e-08 | 199.4433 | -0.012063 | 0.0051077 | 1.1664e-08 | 195.7844
M16A | B | 14 | 1 | H | -0.025476 | 0.0049646 | -1.457e-08 | 201.4251 | -0.012176 | 0.0050649 | 1.0671e-08 | 197.4354
M16A | B | 15 | 1 | H | -0.025377 | 0.0051184 | -1.5586e-08 | 195.3721 | -0.0097171 | 0.005196 | 1.3807e-08 | 192.4543
M16A | B | 16 | 1 | H | -0.024249 | 0.0051283 | -1.888e-08 | 194.9966 | -0.010692 | 0.0052121 | 1.0654e-08 | 191.8619
M16B | A | 1 | 1 | H | -0.024876 | 0.0048391 | -1.6369e-09 | 206.6506 | -0.01301 | 0.0049617 | 8.5728e-09 | 201.5443
M16B | A | 2 | 1 | H | -0.024761 | 0.0048676 | -4.6029e-09 | 205.4405 | -0.012601 | 0.0049958 | 6.7997e-09 | 200.1691
M16B | A | 3 | 1 | H | -0.024179 | 0.004971 | -1.3673e-09 | 201.1657 | -0.010816 | 0.005085 | 1.5397e-08 | 196.6552
M16B | A | 4 | 1 | H | -0.023591 | 0.00482 | -9.2662e-10 | 207.4706 | -0.01106 | 0.0049464 | 1.2028e-08 | 202.1671
M16B | A | 5 | 1 | H | -0.022839 | 0.0048965 | -3.0931e-11 | 204.2294 | -0.010883 | 0.0050099 | 1.5019e-08 | 199.6041
M16B | A | 6 | 1 | H | -0.026077 | 0.0049186 | -5.8927e-09 | 203.3084 | -0.0098743 | 0.0050339 | 1.3801e-08 | 198.6529
M16B | A | 7 | 1 | H | -0.023541 | 0.0049182 | -2.078e-09 | 203.3271 | -0.0086117 | 0.0050188 | 1.8495e-08 | 199.2525
M16B | A | 8 | 1 | H | -0.023885 | 0.0049166 | -4.2445e-09 | 203.3931 | -0.0076244 | 0.0050316 | 1.9394e-08 | 198.745
M16B | A | 9 | 1 | H | -0.02138 | 0.0049933 | -6.236e-10 | 200.2697 | -0.0085761 | 0.0050869 | 1.9224e-08 | 196.5834
M16B | A | 10 | 1 | H | -0.022786 | 0.004906 | -5.4299e-09 | 203.8319 | -0.0098787 | 0.005005 | 1.4188e-08 | 199.7983
M16B | A | 11 | 1 | H | -0.020554 | 0.0049864 | -5.3388e-10 | 200.5457 | -0.0055519 | 0.0050781 | 2.3507e-08 | 196.9254
M16B | A | 12 | 1 | H | -0.019174 | 0.0049801 | -1.6927e-09 | 200.8004 | -0.0056583 | 0.0050907 | 2.2326e-08 | 196.4367
M16B | A | 13 | 1 | H | -0.023476 | 0.0049869 | -9.3078e-09 | 200.5265 | -0.0044993 | 0.0050722 | 2.3884e-08 | 197.1532
M16B | A | 14 | 1 | H | -0.023934 | 0.0050111 | -9.7282e-09 | 199.5584 | -0.0034526 | 0.0050945 | 2.5152e-08 | 196.2883
M16B | A | 15 | 1 | H | -0.020106 | 0.0051041 | -5.2938e-09 | 195.9223 | -0.0057866 | 0.0052204 | 1.873e-08 | 191.5578
M16B | A | 16 | 1 | H | -0.021323 | 0.0050934 | -1.0546e-08 | 196.332 | -0.0080811 | 0.0051856 | 1.2343e-08 | 192.84
M16B | B | 1 | 1 | H | -0.024731 | 0.0048431 | -1.5388e-09 | 206.4785 | -0.012012 | 0.004967 | 9.1584e-09 | 201.3294
M16B | B | 2 | 1 | H | -0.026231 | 0.004876 | -7.6155e-09 | 205.088 | -0.0076196 | 0.0049922 | 1.0558e-08 | 200.311
M16B | B | 3 | 1 | H | -0.023718 | 0.0049748 | -1.1269e-09 | 201.0128 | -0.008642 | 0.0050855 | 1.8656e-08 | 196.6389
M16B | B | 4 | 1 | H | -0.022813 | 0.0048236 | -1.025e-09 | 207.3147 | -0.011979 | 0.0049539 | 1.1386e-08 | 201.8598
M16B | B | 5 | 1 | H | -0.023483 | 0.0049025 | -1.3571e-09 | 203.9775 | -0.010109 | 0.0050135 | 1.6824e-08 | 199.4621
M16B | B | 6 | 1 | H | -0.024884 | 0.0049215 | -5.6858e-09 | 203.1884 | -0.0083947 | 0.0050359 | 1.5571e-08 | 198.5749
M16B | B | 7 | 1 | H | -0.025256 | 0.0049284 | -5.1526e-09 | 202.905 | -0.0080887 | 0.0050235 | 1.9597e-08 | 199.0641
M16B | B | 8 | 1 | H | -0.025229 | 0.0049229 | -4.7815e-09 | 203.1318 | -0.0070916 | 0.0050391 | 1.7785e-08 | 198.449
M16B | B | 9 | 1 | H | -0.022485 | 0.0050022 | -3.656e-09 | 199.9103 | -0.0086466 | 0.0050929 | 1.9346e-08 | 196.3525
M16B | B | 10 | 1 | H | -0.025617 | 0.0049186 | -1.0279e-08 | 203.3114 | -0.0093709 | 0.0050112 | 1.4067e-08 | 199.5512
M16B | B | 11 | 1 | H | -0.02236 | 0.0049966 | -4.2501e-09 | 200.1351 | -0.0055039 | 0.0050829 | 2.4189e-08 | 196.7384
M16B | B | 12 | 1 | H | -0.021651 | 0.0049938 | -7.5452e-09 | 200.2503 | -0.0057043 | 0.0050985 | 2.1235e-08 | 196.1377
M16B | B | 13 | 1 | H | -0.020719 | 0.0049849 | -5.4278e-09 | 200.6077 | -0.0062396 | 0.0050819 | 2.2589e-08 | 196.7758
M16B | B | 14 | 1 | H | -0.022531 | 0.005012 | -7.2201e-09 | 199.5193 | -0.0043084 | 0.0051047 | 2.2509e-08 | 195.8978
M16B | B | 15 | 1 | H | -0.022036 | 0.005115 | -9.2337e-09 | 195.504 | -0.0057643 | 0.005228 | 1.8005e-08 | 191.277
M16B | B | 16 | 1 | H | -0.019793 | 0.0050932 | -7.7209e-09 | 196.3384 | -0.0072339 | 0.0051905 | 1.3291e-08 | 192.6582
M13 | A | 1 | 1 | L | 0.064474 | 0.12519 | -5.3429e-07 | 7.9877 | 0.027241 | 0.12641 | -5.6362e-07 | 7.9107
M13 | A | 2 | 1 | L | 0.080792 | 0.12968 | -5.338e-07 | 7.7112 | 0.004363 | 0.1311 | -5.709e-07 | 7.6279
M13 | A | 3 | 1 | L | 0.099315 | 0.12462 | -4.9786e-07 | 8.0247 | 0.012676 | 0.12596 | -5.2104e-07 | 7.9392
M13 | A | 4 | 1 | L | 0.11174 | 0.12996 | -5.0894e-07 | 7.6946 | 0.049271 | 0.13138 | -5.3947e-07 | 7.6114
M13 | A | 5 | 1 | L | 0.22611 | 0.12449 | -4.7075e-07 | 8.0325 | 0.19773 | 0.12594 | -5.1267e-07 | 7.9404
M13 | A | 6 | 1 | L | 0.29432 | 0.13091 | -5.0031e-07 | 7.639 | 0.23137 | 0.13236 | -5.4698e-07 | 7.5549
M13 | A | 7 | 1 | L | 0.28284 | 0.12531 | -4.9125e-07 | 7.98 | 0.25089 | 0.12666 | -5.2351e-07 | 7.8952
M13 | A | 8 | 1 | L | 0.3804 | 0.13073 | -5.355e-07 | 7.6491 | 0.42201 | 0.132 | -5.3569e-07 | 7.5759
M13 | A | 9 | 1 | L | 0.26641 | 0.12597 | -5.073e-07 | 7.9381 | 0.24108 | 0.12724 | -5.1945e-07 | 7.8594
M13 | A | 10 | 1 | L | 0.29869 | 0.13068 | -4.8651e-07 | 7.6525 | 0.23297 | 0.13196 | -4.8958e-07 | 7.5783
M13 | A | 11 | 1 | L | 0.3325 | 0.12606 | -5.041e-07 | 7.9325 | 0.35926 | 0.12736 | -5.3028e-07 | 7.8516
M13 | A | 12 | 1 | L | 0.36095 | 0.13122 | -4.824e-07 | 7.6209 | 0.31334 | 0.13264 | -5.2099e-07 | 7.5394
M13 | A | 13 | 1 | L | 0.26865 | 0.12579 | -4.7306e-07 | 7.9497 | 0.25047 | 0.12713 | -5.1043e-07 | 7.8659
M13 | A | 14 | 1 | L | 0.22573 | 0.13175 | -4.8983e-07 | 7.5898 | 0.1824 | 0.13311 | -5.0868e-07 | 7.5126
M13 | A | 15 | 1 | L | 0.18131 | 0.12694 | -4.713e-07 | 7.8779 | 0.1086 | 0.12822 | -4.9355e-07 | 7.799
M13 | A | 16 | 1 | L | 0.37427 | 0.13323 | -4.897e-07 | 7.5056 | 0.27488 | 0.13473 | -5.538e-07 | 7.4223
M13 | B | 1 | 1 | L | 0.064589 | 0.12524 | -5.3586e-07 | 7.9844 | 0.030299 | 0.12649 | -5.5908e-07 | 7.906
M13 | B | 2 | 1 | L | 0.082987 | 0.12971 | -5.3701e-07 | 7.7093 | 0.0095948 | 0.13115 | -5.6687e-07 | 7.625
M13 | B | 3 | 1 | L | 0.10465 | 0.12463 | -4.9425e-07 | 8.0238 | 0.0035186 | 0.12602 | -5.1944e-07 | 7.9351
M13 | B | 4 | 1 | L | 0.1167 | 0.12998 | -5.008e-07 | 7.6935 | 0.053444 | 0.13143 | -5.3224e-07 | 7.6085
M13 | B | 5 | 1 | L | 0.22748 | 0.12454 | -4.7017e-07 | 8.0297 | 0.19516 | 0.12603 | -5.1665e-07 | 7.9347
M13 | B | 6 | 1 | L | 0.29351 | 0.13095 | -4.9998e-07 | 7.6368 | 0.23084 | 0.13243 | -5.466e-07 | 7.5513
M13 | B | 7 | 1 | L | 0.28631 | 0.12534 | -4.8524e-07 | 7.9783 | 0.25694 | 0.12673 | -5.2436e-07 | 7.8905
M13 | B | 8 | 1 | L | 0.38444 | 0.13077 | -5.3077e-07 | 7.6473 | 0.42133 | 0.13208 | -5.4033e-07 | 7.5712
M13 | B | 9 | 1 | L | 0.26686 | 0.126 | -5.0663e-07 | 7.9365 | 0.24592 | 0.12728 | -5.1111e-07 | 7.8569
M13 | B | 10 | 1 | L | 0.30356 | 0.13069 | -4.835e-07 | 7.6516 | 0.23543 | 0.13201 | -4.8852e-07 | 7.5751
M13 | B | 11 | 1 | L | 0.33404 | 0.12607 | -5.0184e-07 | 7.9322 | 0.36344 | 0.1274 | -5.2757e-07 | 7.8493
M13 | B | 12 | 1 | L | 0.3656 | 0.13125 | -4.7998e-07 | 7.6191 | 0.31726 | 0.13271 | -5.223e-07 | 7.5352
M13 | B | 13 | 1 | L | 0.27031 | 0.12581 | -4.7268e-07 | 7.9485 | 0.24961 | 0.1272 | -5.1522e-07 | 7.8616
M13 | B | 14 | 1 | L | 0.22592 | 0.1318 | -4.891e-07 | 7.587 | 0.18483 | 0.13319 | -5.0835e-07 | 7.5083
M13 | B | 15 | 1 | L | 0.17866 | 0.12695 | -4.6765e-07 | 7.8772 | 0.10185 | 0.12828 | -4.9436e-07 | 7.7952
M13 | B | 16 | 1 | L | 0.37666 | 0.13325 | -4.9063e-07 | 7.5046 | 0.27479 | 0.13477 | -5.4988e-07 | 7.4199

## RSB rad coeff
Band | HAM | Subsample | Gain State | Band Median | [7 empty] | Band Average
|  |  |  | Eside A |  |  |  | Eside B |  |  |  | Eside A |  |  |  | Eside B
|  |  |  | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1)
I1 | A | 1 | H | 1.3587 | 0.2332 | -1.446e-06 | 4.2882 | 1.7212 | 0.23865 | -2.6946e-06 | 4.1902 | 1.3775 | 0.23316 | -1.4221e-06 | 4.2889 | 1.728 | 0.23865 | -2.677e-06 | 4.1902
I1 | A | 2 | H | 1.4177 | 0.2332 | -1.4376e-06 | 4.2882 | 1.7784 | 0.2387 | -2.686e-06 | 4.1894 | 1.4161 | 0.23317 | -1.4217e-06 | 4.2887 | 1.7739 | 0.23866 | -2.6776e-06 | 4.19
I1 | B | 1 | H | 1.3642 | 0.23215 | -1.4257e-06 | 4.3076 | 1.7119 | 0.23775 | -2.6588e-06 | 4.2061 | 1.3811 | 0.23219 | -1.407e-06 | 4.3069 | 1.7237 | 0.23771 | -2.6573e-06 | 4.2068
I1 | B | 2 | H | 1.4323 | 0.23225 | -1.4155e-06 | 4.3057 | 1.7841 | 0.23775 | -2.658e-06 | 4.2061 | 1.4223 | 0.23218 | -1.4061e-06 | 4.3069 | 1.7761 | 0.23772 | -2.6547e-06 | 4.2067
I2 | A | 1 | H | -0.2749 | 0.11545 | -4.705e-07 | 8.6618 | -0.9427 | 0.11925 | -1.4232e-06 | 8.3857 | -0.26038 | 0.11541 | -4.6644e-07 | 8.6648 | -0.94338 | 0.11957 | -1.418e-06 | 8.3636
I2 | A | 2 | H | -0.2512 | 0.11545 | -4.689e-07 | 8.6618 | -0.93185 | 0.11925 | -1.4161e-06 | 8.3857 | -0.25569 | 0.11542 | -4.6613e-07 | 8.6643 | -0.93374 | 0.11957 | -1.415e-06 | 8.3632
I2 | B | 1 | H | -0.2726 | 0.11555 | -4.734e-07 | 8.6543 | -0.9463 | 0.11935 | -1.427e-06 | 8.3787 | -0.26159 | 0.11551 | -4.6817e-07 | 8.6573 | -0.94441 | 0.11963 | -1.4197e-06 | 8.359
I2 | B | 2 | H | -0.25315 | 0.11555 | -4.7175e-07 | 8.6543 | -0.9338 | 0.11935 | -1.4139e-06 | 8.3787 | -0.25646 | 0.11552 | -4.6716e-07 | 8.6566 | -0.93466 | 0.11965 | -1.4177e-06 | 8.3575
I3 | A | 1 | H | 0.05595 | 0.0187 | -4.34e-08 | 53.4759 | 0.09305 | 0.0184 | 8.105e-08 | 54.3478 | 0.047234 | 0.019238 | -4.3881e-08 | 51.9818 | 0.087378 | 0.01915 | 1.0064e-07 | 52.2193
I3 | B | 2 | H | 0.0491 | 0.0187 | -4.355e-08 | 53.4759 | 0.07885 | 0.0184 | 8.14e-08 | 54.3478 | 0.055519 | 0.019238 | -4.5813e-08 | 51.9818 | 0.086006 | 0.019134 | 1.0871e-07 | 52.262
I3 | A | 1 | H | 0.0559 | 0.01875 | -4.39e-08 | 53.3333 | 0.0942 | 0.01845 | 8.145e-08 | 54.2005 | 0.046772 | 0.019272 | -4.4156e-08 | 51.8891 | 0.091422 | 0.019156 | 1.1427e-07 | 52.2023
I3 | B | 2 | H | 0.0495 | 0.01875 | -4.365e-08 | 53.3333 | 0.0777 | 0.01845 | 8.075e-08 | 54.2005 | 0.053997 | 0.019269 | -4.5431e-08 | 51.8975 | 0.084031 | 0.019169 | 1.0784e-07 | 52.1682
M1 | A | 1 | H | -0.95285 | 0.0547 | -2.0105e-07 | 18.2815 | -1.0819 | 0.05785 | -5.5395e-07 | 17.2861 | -0.95866 | 0.0549 | -2.059e-07 | 18.2149 | -1.0853 | 0.058 | -5.6676e-07 | 17.2414
M1 | B | 1 | H | -0.96475 | 0.0559 | -2.0505e-07 | 17.8891 | -1.0876 | 0.0591 | -5.739e-07 | 16.9205 | -0.9627 | 0.056112 | -2.1279e-07 | 17.8213 | -1.0791 | 0.059294 | -5.8222e-07 | 16.8652
M2 | A | 1 | H | 2.6297 | 0.04345 | 7.9525e-07 | 23.015 | 3.1368 | 0.04445 | 8.6545e-07 | 22.4972 | 2.6343 | 0.043569 | 7.9551e-07 | 22.9522 | 3.1347 | 0.044556 | 8.6679e-07 | 22.4435
M2 | B | 1 | H | 2.6399 | 0.04385 | 8.153e-07 | 22.805 | 3.1307 | 0.0449 | 8.815e-07 | 22.2717 | 2.6385 | 0.04395 | 8.1459e-07 | 22.7531 | 3.1277 | 0.044956 | 8.8031e-07 | 22.2438
M3 | A | 1 | H | 0.0492 | 0.0386 | -3.1165e-07 | 25.9067 | 0.71675 | 0.03965 | -4.825e-07 | 25.2207 | 0.041425 | 0.038725 | -3.1143e-07 | 25.8231 | 0.71604 | 0.039756 | -4.8015e-07 | 25.1533
M3 | B | 1 | H | 0.0446 | 0.03875 | -3.161e-07 | 25.8065 | 0.71495 | 0.03975 | -4.8735e-07 | 25.1572 | 0.0416 | 0.038862 | -3.1397e-07 | 25.7317 | 0.71824 | 0.039906 | -4.8359e-07 | 25.0587
M4 | A | 1 | H | 0.32375 | 0.02825 | 2.498e-07 | 35.3982 | 1.1909 | 0.0275 | 9.6955e-07 | 36.3636 | 0.32221 | 0.028338 | 2.4867e-07 | 35.2889 | 1.1899 | 0.027594 | 9.7208e-07 | 36.2401
M4 | B | 1 | H | 0.32215 | 0.0282 | 2.4565e-07 | 35.461 | 1.1885 | 0.0275 | 9.656e-07 | 36.3636 | 0.32228 | 0.028288 | 2.4799e-07 | 35.3513 | 1.1869 | 0.027588 | 9.6827e-07 | 36.2483
M5 | A | 1 | H | -0.0207 | 0.0209 | -1.755e-07 | 47.8469 | 0.0672 | 0.0216 | -2.203e-07 | 46.2963 | -0.022725 | 0.020975 | -1.7418e-07 | 47.6758 | 0.063212 | 0.021619 | -2.2121e-07 | 46.2561
M5 | B | 1 | H | -0.02385 | 0.0208 | -1.752e-07 | 48.0769 | 0.06045 | 0.0215 | -2.23e-07 | 46.5116 | -0.024625 | 0.020881 | -1.737e-07 | 47.8899 | 0.060556 | 0.021537 | -2.2111e-07 | 46.4306
M6 | A | 1 | H | 0.054 | 0.0131 | -2.715e-08 | 76.3359 | 0.0773 | 0.01345 | -4.655e-08 | 74.3494 | 0.0532 | 0.013112 | -2.5937e-08 | 76.2631 | 0.077725 | 0.013463 | -4.6363e-08 | 74.2804
M6 | B | 1 | H | 0.05285 | 0.01305 | -2.675e-08 | 76.6284 | 0.0762 | 0.0134 | -4.68e-08 | 74.6269 | 0.053375 | 0.013069 | -2.5806e-08 | 76.5184 | 0.077606 | 0.013406 | -4.6294e-08 | 74.5921
M7 | A | 1 | H | 0.034 | 0.0101 | -2.84e-08 | 99.0099 | 0.088 | 0.0102 | -1.865e-08 | 98.0392 | 0.034656 | 0.010081 | -2.8131e-08 | 99.194 | 0.088294 | 0.0102 | -1.86e-08 | 98.0392
M7 | B | 1 | H | 0.03475 | 0.0101 | -2.825e-08 | 99.0099 | 0.08525 | 0.0102 | -1.93e-08 | 98.0392 | 0.0353 | 0.010094 | -2.8075e-08 | 99.0712 | 0.08705 | 0.0102 | -1.9194e-08 | 98.0392
M8 | A | 1 | H | 0.62985 | 0.03 | 2.5465e-07 | 33.3333 | 1.0484 | 0.0292 | 6.399e-07 | 34.2466 | 0.62699 | 0.029956 | 2.5317e-07 | 33.382 | 1.0493 | 0.029119 | 6.3535e-07 | 34.3421
M8 | B | 1 | H | 0.62875 | 0.0301 | 2.567e-07 | 33.2226 | 1.0489 | 0.0293 | 6.437e-07 | 34.1297 | 0.62711 | 0.030056 | 2.5503e-07 | 33.271 | 1.0498 | 0.029219 | 6.3997e-07 | 34.2246
M9 | A | 1 | H | 0.43035 | 0.025 | 2.053e-07 | 40 | 0.58115 | 0.02445 | 3.1665e-07 | 40.8998 | 0.43437 | 0.024981 | 2.0786e-07 | 40.03 | 0.58828 | 0.024463 | 3.1637e-07 | 40.8789
M9 | B | 1 | H | 0.43075 | 0.025 | 2.0695e-07 | 40 | 0.58185 | 0.0245 | 3.175e-07 | 40.8163 | 0.43469 | 0.025056 | 2.0871e-07 | 39.9102 | 0.58861 | 0.024538 | 3.1767e-07 | 40.7539
M10 | A | 1 | H | 0.10725 | 0.0215 | -9.52e-08 | 46.5116 | 0.15275 | 0.02105 | 7.195e-08 | 47.5059 | 0.099894 | 0.021412 | -9.4481e-08 | 46.7017 | 0.14216 | 0.020988 | 7.3819e-08 | 47.6474
M10 | B | 1 | H | 0.1075 | 0.02155 | -9.56e-08 | 46.4037 | 0.1535 | 0.02115 | 7.205e-08 | 47.2813 | 0.10007 | 0.021463 | -9.4606e-08 | 46.5929 | 0.14226 | 0.021038 | 7.38e-08 | 47.5342
M11 | A | 1 | H | 0.0226 | 0.00875 | -3.64e-08 | 114.2857 | 0.035 | 0.00865 | -3.69e-08 | 115.6069 | 0.020625 | 0.00875 | -3.5625e-08 | 114.2857 | 0.0321 | 0.008675 | -3.5975e-08 | 115.2738
M11 | B | 1 | H | 0.02265 | 0.00875 | -3.66e-08 | 114.2857 | 0.0352 | 0.00865 | -3.7e-08 | 115.6069 | 0.020587 | 0.0087562 | -3.5794e-08 | 114.2041 | 0.032025 | 0.008675 | -3.6094e-08 | 115.2738
M1 | A | 1 | L | 16.2269 | 0.185 | 3.0219e-05 | 5.4054 | 20.4546 | 0.18005 | 4.1643e-05 | 5.554 | 16.0457 | 0.18403 | 3.0145e-05 | 5.4339 | 20.5007 | 0.18003 | 4.2097e-05 | 5.5548
M1 | B | 1 | L | 16.0906 | 0.18745 | 3.1691e-05 | 5.3348 | 21.2812 | 0.18185 | 4.5341e-05 | 5.499 | 16.0845 | 0.18797 | 3.1521e-05 | 5.32 | 21.1542 | 0.18221 | 4.5329e-05 | 5.4883
M2 | A | 1 | L | 3.1775 | 0.22655 | 1.375e-06 | 4.414 | 3.09 | 0.23215 | 2.429e-06 | 4.3076 | 3.1763 | 0.22705 | 1.3701e-06 | 4.4043 | 3.1203 | 0.23274 | 2.4003e-06 | 4.2966
M2 | B | 1 | L | 3.2619 | 0.2282 | 1.497e-06 | 4.3821 | 3.1146 | 0.2339 | 2.4445e-06 | 4.2753 | 3.2549 | 0.22889 | 1.4755e-06 | 4.369 | 3.1302 | 0.23485 | 2.4616e-06 | 4.258
M3 | A | 1 | L | -4.0466 | 0.23935 | -1.779e-06 | 4.178 | -3.4181 | 0.24275 | -1.2775e-06 | 4.1195 | -4.0396 | 0.24007 | -1.7888e-06 | 4.1654 | -3.4826 | 0.24358 | -1.2893e-06 | 4.1054
M3 | B | 1 | L | -4.0407 | 0.24025 | -1.791e-06 | 4.1623 | -3.3937 | 0.2437 | -1.2505e-06 | 4.1034 | -4.0345 | 0.24088 | -1.7948e-06 | 4.1514 | -3.4294 | 0.24437 | -1.2668e-06 | 4.0922
M4 | A | 1 | L | 1.9894 | 0.2221 | 6.07e-07 | 4.5025 | 3.3369 | 0.22505 | 3.26e-07 | 4.4435 | 2.0004 | 0.22299 | 6.0725e-07 | 4.4844 | 3.3244 | 0.22595 | 3.2563e-07 | 4.4258
M4 | B | 1 | L | 1.9893 | 0.22185 | 6.115e-07 | 4.5076 | 3.3477 | 0.22485 | 3.28e-07 | 4.4474 | 2.0009 | 0.22278 | 6.0612e-07 | 4.4888 | 3.3316 | 0.22576 | 3.3006e-07 | 4.4294
M5 | A | 1 | L | 3.6421 | 0.20225 | -1.525e-07 | 4.9444 | 4.9001 | 0.2057 | -7.275e-07 | 4.8614 | 3.584 | 0.20251 | -1.8625e-07 | 4.938 | 4.8883 | 0.20594 | -7.0775e-07 | 4.8557
M5 | B | 1 | L | 3.5324 | 0.2017 | -2.515e-07 | 4.9579 | 4.8928 | 0.2048 | -7.115e-07 | 4.8828 | 3.5231 | 0.20171 | -2.4775e-07 | 4.9576 | 4.8889 | 0.20502 | -7.0163e-07 | 4.8776
M7 | A | 1 | L | -0.6758 | 0.1153 | -8.355e-07 | 8.673 | -1.4554 | 0.11925 | -1.5375e-06 | 8.3857 | -0.68074 | 0.11531 | -8.1994e-07 | 8.6721 | -1.4648 | 0.1193 | -1.5235e-06 | 8.3822
M7 | B | 1 | L | -0.66765 | 0.1154 | -8.35e-07 | 8.6655 | -1.4508 | 0.1193 | -1.537e-06 | 8.3822 | -0.67844 | 0.11542 | -8.2062e-07 | 8.6641 | -1.4652 | 0.11938 | -1.5252e-06 | 8.3765
Band | HAM | Detector | Subsample | Gain State | Detector Dependent | [8 empty] | Notes:
[5 empty] | Eside A |  |  |  | Eside B |  |  |  |  | 1. This table provides VIIRS F2 RSB radiometric coefficients.
[5 empty] | c0 | c1 | c2 | gain (1/c1) | c0 | c1 | c2 | gain (1/c1) |  | 2. The results were derived from RC-1 using the SIS-100 source (ScMAfor M1, M2, and M3 low gain).
I1 | A | 1 | 1 | H | 1.4329 | 0.2291 | -1.3509e-06 | 4.3649 | 1.7998 | 0.2348 | -2.5564e-06 | 4.2589 |  | 3. Quadratic fits of the SIS-100 radiance vs dn were used to determine the radiometric coefficients.
I1 | A | 2 | 1 | H | 1.4966 | 0.2291 | -1.2822e-06 | 4.3649 | 1.7816 | 0.2348 | -2.5207e-06 | 4.2589 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | A | 3 | 1 | H | 1.4104 | 0.2295 | -1.3755e-06 | 4.3573 | 1.7888 | 0.2351 | -2.6298e-06 | 4.2535 |  | 5. The test was set to operational day mode.
I1 | A | 4 | 1 | H | 1.4097 | 0.2303 | -1.3598e-06 | 4.3422 | 1.7219 | 0.2359 | -2.6283e-06 | 4.2391
I1 | A | 5 | 1 | H | 1.4674 | 0.2302 | -1.3271e-06 | 4.344 | 1.7594 | 0.2358 | -2.564e-06 | 4.2409
I1 | A | 6 | 1 | H | 1.3487 | 0.2299 | -1.2474e-06 | 4.3497 | 1.7061 | 0.2352 | -2.4705e-06 | 4.2517
I1 | A | 7 | 1 | H | 1.4463 | 0.2318 | -1.25e-06 | 4.3141 | 1.7487 | 0.2373 | -2.4998e-06 | 4.2141
I1 | A | 8 | 1 | H | 1.365 | 0.2306 | -1.2649e-06 | 4.3365 | 1.6908 | 0.2359 | -2.4931e-06 | 4.2391
I1 | A | 9 | 1 | H | 1.3682 | 0.2299 | -1.2771e-06 | 4.3497 | 1.8874 | 0.2351 | -2.4193e-06 | 4.2535
I1 | A | 10 | 1 | H | 1.3839 | 0.2302 | -1.2711e-06 | 4.344 | 1.6507 | 0.2356 | -2.5173e-06 | 4.2445
I1 | A | 11 | 1 | H | 1.4186 | 0.2345 | -1.4335e-06 | 4.2644 | 1.7439 | 0.2401 | -2.7109e-06 | 4.1649
I1 | A | 12 | 1 | H | 1.3485 | 0.2345 | -1.4456e-06 | 4.2644 | 1.8162 | 0.2397 | -2.654e-06 | 4.1719
I1 | A | 13 | 1 | H | 1.3429 | 0.2331 | -1.4597e-06 | 4.29 | 1.7785 | 0.2385 | -2.6879e-06 | 4.1929
I1 | A | 14 | 1 | H | 1.7373 | 0.2333 | -1.5323e-06 | 4.2863 | 1.9605 | 0.2389 | -2.8397e-06 | 4.1859
I1 | A | 15 | 1 | H | 1.4227 | 0.235 | -1.4961e-06 | 4.2553 | 1.8149 | 0.2406 | -2.8078e-06 | 4.1563
I1 | A | 16 | 1 | H | 1.3539 | 0.2361 | -1.5451e-06 | 4.2355 | 1.7295 | 0.2416 | -2.8179e-06 | 4.1391
I1 | A | 17 | 1 | H | 1.3364 | 0.2342 | -1.4694e-06 | 4.2699 | 1.7571 | 0.2396 | -2.7039e-06 | 4.1736
I1 | A | 18 | 1 | H | 1.3596 | 0.2328 | -1.4464e-06 | 4.2955 | 1.6889 | 0.2381 | -2.7013e-06 | 4.1999
I1 | A | 19 | 1 | H | 1.3768 | 0.2316 | -1.3331e-06 | 4.3178 | 1.7094 | 0.2371 | -2.6152e-06 | 4.2176
I1 | A | 20 | 1 | H | 1.3123 | 0.2309 | -1.2688e-06 | 4.3309 | 1.6549 | 0.2363 | -2.53e-06 | 4.2319
I1 | A | 21 | 1 | H | 1.3504 | 0.2325 | -1.3589e-06 | 4.3011 | 1.6923 | 0.2379 | -2.597e-06 | 4.2034
I1 | A | 22 | 1 | H | 1.1671 | 0.2321 | -1.2114e-06 | 4.3085 | 1.5073 | 0.2373 | -2.4085e-06 | 4.2141
I1 | A | 23 | 1 | H | 1.2979 | 0.2345 | -1.4816e-06 | 4.2644 | 1.6539 | 0.2402 | -2.7765e-06 | 4.1632
I1 | A | 24 | 1 | H | 1.3381 | 0.2376 | -1.5777e-06 | 4.2088 | 1.6977 | 0.2431 | -2.8588e-06 | 4.1135
I1 | A | 25 | 1 | H | 1.3578 | 0.2384 | -1.5816e-06 | 4.1946 | 1.7398 | 0.244 | -2.8855e-06 | 4.0984
I1 | A | 26 | 1 | H | 1.399 | 0.2371 | -1.5902e-06 | 4.2176 | 1.8147 | 0.2426 | -2.8587e-06 | 4.122
I1 | A | 27 | 1 | H | 1.3553 | 0.2369 | -1.5766e-06 | 4.2212 | 1.6976 | 0.2426 | -2.8692e-06 | 4.122
I1 | A | 28 | 1 | H | 1.3121 | 0.2333 | -1.473e-06 | 4.2863 | 1.6603 | 0.2388 | -2.7445e-06 | 4.1876
I1 | A | 29 | 1 | H | 1.312 | 0.2339 | -1.5303e-06 | 4.2753 | 1.7206 | 0.2394 | -2.7612e-06 | 4.1771
I1 | A | 30 | 1 | H | 1.4038 | 0.2349 | -1.4979e-06 | 4.2571 | 1.5942 | 0.2405 | -2.764e-06 | 4.158
I1 | A | 31 | 1 | H | 1.3342 | 0.2383 | -1.6131e-06 | 4.1964 | 1.6585 | 0.2441 | -2.9403e-06 | 4.0967
I1 | A | 32 | 1 | H | 1.3153 | 0.235 | -1.5789e-06 | 4.2553 | 1.671 | 0.2404 | -2.8329e-06 | 4.1597
I1 | A | 1 | 2 | H | 1.4087 | 0.2291 | -1.3572e-06 | 4.3649 | 1.8028 | 0.2347 | -2.5464e-06 | 4.2608
I1 | A | 2 | 2 | H | 1.4576 | 0.2292 | -1.2912e-06 | 4.363 | 1.7744 | 0.2348 | -2.5131e-06 | 4.2589
I1 | A | 3 | 2 | H | 1.3843 | 0.2295 | -1.3735e-06 | 4.3573 | 1.7617 | 0.2352 | -2.6375e-06 | 4.2517
I1 | A | 4 | 2 | H | 1.393 | 0.2304 | -1.3563e-06 | 4.3403 | 1.7189 | 0.2359 | -2.6136e-06 | 4.2391
I1 | A | 5 | 2 | H | 1.4394 | 0.2302 | -1.3263e-06 | 4.344 | 1.7505 | 0.2359 | -2.5731e-06 | 4.2391
I1 | A | 6 | 2 | H | 1.3115 | 0.2299 | -1.2383e-06 | 4.3497 | 1.6957 | 0.2352 | -2.4651e-06 | 4.2517
I1 | A | 7 | 2 | H | 1.4332 | 0.2318 | -1.2464e-06 | 4.3141 | 1.7335 | 0.2373 | -2.497e-06 | 4.2141
I1 | A | 8 | 2 | H | 1.3497 | 0.2306 | -1.2526e-06 | 4.3365 | 1.682 | 0.2359 | -2.4946e-06 | 4.2391
I1 | A | 9 | 2 | H | 1.3637 | 0.23 | -1.2796e-06 | 4.3478 | 1.8568 | 0.2351 | -2.4287e-06 | 4.2535
I1 | A | 10 | 2 | H | 1.3722 | 0.2302 | -1.2663e-06 | 4.344 | 1.6383 | 0.2356 | -2.5233e-06 | 4.2445
I1 | A | 11 | 2 | H | 1.4213 | 0.2345 | -1.434e-06 | 4.2644 | 1.7355 | 0.2401 | -2.7319e-06 | 4.1649
I1 | A | 12 | 2 | H | 1.3408 | 0.2345 | -1.4412e-06 | 4.2644 | 1.7911 | 0.2398 | -2.676e-06 | 4.1701
I1 | A | 13 | 2 | H | 1.3431 | 0.2331 | -1.4591e-06 | 4.29 | 1.8139 | 0.2385 | -2.679e-06 | 4.1929
I1 | A | 14 | 2 | H | 1.6623 | 0.2333 | -1.5438e-06 | 4.2863 | 1.9002 | 0.2389 | -2.8271e-06 | 4.1859
I1 | A | 15 | 2 | H | 1.4353 | 0.235 | -1.4999e-06 | 4.2553 | 1.8216 | 0.2407 | -2.8216e-06 | 4.1545
I1 | A | 16 | 2 | H | 1.4046 | 0.2361 | -1.537e-06 | 4.2355 | 1.7445 | 0.2416 | -2.849e-06 | 4.1391
I1 | A | 17 | 2 | H | 1.4647 | 0.2342 | -1.4708e-06 | 4.2699 | 1.8157 | 0.2397 | -2.721e-06 | 4.1719
I1 | A | 18 | 2 | H | 1.4891 | 0.2327 | -1.4249e-06 | 4.2974 | 1.7804 | 0.2381 | -2.693e-06 | 4.1999
I1 | A | 19 | 2 | H | 1.4948 | 0.2316 | -1.362e-06 | 4.3178 | 1.78 | 0.2372 | -2.6158e-06 | 4.2159
I1 | A | 20 | 2 | H | 1.4139 | 0.2309 | -1.2446e-06 | 4.3309 | 1.7113 | 0.2363 | -2.5258e-06 | 4.2319
I1 | A | 21 | 2 | H | 1.465 | 0.2324 | -1.3259e-06 | 4.3029 | 1.8258 | 0.2378 | -2.5847e-06 | 4.2052
I1 | A | 22 | 2 | H | 1.149 | 0.232 | -1.1666e-06 | 4.3103 | 1.475 | 0.2373 | -2.3837e-06 | 4.2141
I1 | A | 23 | 2 | H | 1.4141 | 0.2345 | -1.4647e-06 | 4.2644 | 1.7831 | 0.2401 | -2.7556e-06 | 4.1649
I1 | A | 24 | 2 | H | 1.4538 | 0.2376 | -1.5742e-06 | 4.2088 | 1.7769 | 0.2432 | -2.8811e-06 | 4.1118
I1 | A | 25 | 2 | H | 1.4819 | 0.2383 | -1.5612e-06 | 4.1964 | 1.798 | 0.2441 | -2.8989e-06 | 4.0967
I1 | A | 26 | 2 | H | 1.4677 | 0.2372 | -1.6142e-06 | 4.2159 | 1.9325 | 0.2425 | -2.8356e-06 | 4.1237
I1 | A | 27 | 2 | H | 1.4643 | 0.2369 | -1.5998e-06 | 4.2212 | 1.8558 | 0.2425 | -2.8591e-06 | 4.1237
I1 | A | 28 | 2 | H | 1.3424 | 0.2335 | -1.5014e-06 | 4.2827 | 1.7373 | 0.2389 | -2.7598e-06 | 4.1859
I1 | A | 29 | 2 | H | 1.4277 | 0.2339 | -1.5163e-06 | 4.2753 | 1.9221 | 0.2393 | -2.7141e-06 | 4.1789
I1 | A | 30 | 2 | H | 1.3995 | 0.235 | -1.5022e-06 | 4.2553 | 1.7491 | 0.2405 | -2.788e-06 | 4.158
I1 | A | 31 | 2 | H | 1.4118 | 0.2384 | -1.6525e-06 | 4.1946 | 1.8527 | 0.244 | -2.9265e-06 | 4.0984
I1 | A | 32 | 2 | H | 1.4555 | 0.235 | -1.6118e-06 | 4.2553 | 1.7462 | 0.2405 | -2.8645e-06 | 4.158
I1 | B | 1 | 1 | H | 1.4352 | 0.2281 | -1.3422e-06 | 4.384 | 1.8116 | 0.2339 | -2.5389e-06 | 4.2753
I1 | B | 2 | 1 | H | 1.4936 | 0.2282 | -1.2647e-06 | 4.3821 | 1.7746 | 0.2339 | -2.5036e-06 | 4.2753
I1 | B | 3 | 1 | H | 1.4005 | 0.2286 | -1.3721e-06 | 4.3745 | 1.7992 | 0.2342 | -2.602e-06 | 4.2699
I1 | B | 4 | 1 | H | 1.4236 | 0.2294 | -1.3355e-06 | 4.3592 | 1.7025 | 0.235 | -2.5998e-06 | 4.2553
I1 | B | 5 | 1 | H | 1.4628 | 0.2293 | -1.3178e-06 | 4.3611 | 1.766 | 0.2349 | -2.541e-06 | 4.2571
I1 | B | 6 | 1 | H | 1.3393 | 0.229 | -1.2416e-06 | 4.3668 | 1.7064 | 0.2343 | -2.4533e-06 | 4.268
I1 | B | 7 | 1 | H | 1.446 | 0.2309 | -1.2348e-06 | 4.3309 | 1.7424 | 0.2364 | -2.4908e-06 | 4.2301
I1 | B | 8 | 1 | H | 1.3655 | 0.2297 | -1.266e-06 | 4.3535 | 1.6893 | 0.235 | -2.4719e-06 | 4.2553
I1 | B | 9 | 1 | H | 1.383 | 0.229 | -1.2522e-06 | 4.3668 | 1.8719 | 0.2342 | -2.4114e-06 | 4.2699
I1 | B | 10 | 1 | H | 1.3803 | 0.2293 | -1.2658e-06 | 4.3611 | 1.6339 | 0.2347 | -2.503e-06 | 4.2608
I1 | B | 11 | 1 | H | 1.4307 | 0.2334 | -1.4023e-06 | 4.2845 | 1.7256 | 0.2392 | -2.7013e-06 | 4.1806
I1 | B | 12 | 1 | H | 1.363 | 0.2335 | -1.4291e-06 | 4.2827 | 1.8032 | 0.2388 | -2.6526e-06 | 4.1876
I1 | B | 13 | 1 | H | 1.3589 | 0.232 | -1.4286e-06 | 4.3103 | 1.7718 | 0.2376 | -2.6614e-06 | 4.2088
I1 | B | 14 | 1 | H | 1.7451 | 0.2323 | -1.5173e-06 | 4.3048 | 1.9568 | 0.2379 | -2.813e-06 | 4.2034
I1 | B | 15 | 1 | H | 1.4273 | 0.234 | -1.4771e-06 | 4.2735 | 1.8108 | 0.2397 | -2.7852e-06 | 4.1719
I1 | B | 16 | 1 | H | 1.3809 | 0.2351 | -1.5176e-06 | 4.2535 | 1.7186 | 0.2406 | -2.8152e-06 | 4.1563
I1 | B | 17 | 1 | H | 1.3416 | 0.2332 | -1.4532e-06 | 4.2882 | 1.737 | 0.2387 | -2.7014e-06 | 4.1894
I1 | B | 18 | 1 | H | 1.3707 | 0.2318 | -1.4229e-06 | 4.3141 | 1.6943 | 0.2371 | -2.6562e-06 | 4.2176
I1 | B | 19 | 1 | H | 1.3617 | 0.2306 | -1.332e-06 | 4.3365 | 1.7104 | 0.2362 | -2.5875e-06 | 4.2337
I1 | B | 20 | 1 | H | 1.3029 | 0.23 | -1.266e-06 | 4.3478 | 1.6384 | 0.2354 | -2.5145e-06 | 4.2481
I1 | B | 21 | 1 | H | 1.3531 | 0.2315 | -1.3508e-06 | 4.3197 | 1.6954 | 0.2369 | -2.5758e-06 | 4.2212
I1 | B | 22 | 1 | H | 1.1766 | 0.2311 | -1.1859e-06 | 4.3271 | 1.4841 | 0.2363 | -2.3941e-06 | 4.2319
I1 | B | 23 | 1 | H | 1.2977 | 0.2335 | -1.4602e-06 | 4.2827 | 1.6481 | 0.2392 | -2.7603e-06 | 4.1806
I1 | B | 24 | 1 | H | 1.3425 | 0.2366 | -1.5678e-06 | 4.2265 | 1.6895 | 0.2421 | -2.8327e-06 | 4.1305
I1 | B | 25 | 1 | H | 1.3589 | 0.2374 | -1.5726e-06 | 4.2123 | 1.7118 | 0.2431 | -2.8797e-06 | 4.1135
I1 | B | 26 | 1 | H | 1.4142 | 0.2361 | -1.5581e-06 | 4.2355 | 1.8221 | 0.2415 | -2.826e-06 | 4.1408
I1 | B | 27 | 1 | H | 1.3257 | 0.236 | -1.5956e-06 | 4.2373 | 1.7121 | 0.2416 | -2.8334e-06 | 4.1391
I1 | B | 28 | 1 | H | 1.3109 | 0.2323 | -1.4567e-06 | 4.3048 | 1.6523 | 0.2379 | -2.7185e-06 | 4.2034
I1 | B | 29 | 1 | H | 1.3331 | 0.2329 | -1.5046e-06 | 4.2937 | 1.7222 | 0.2384 | -2.7356e-06 | 4.1946
I1 | B | 30 | 1 | H | 1.4392 | 0.2339 | -1.4626e-06 | 4.2753 | 1.6098 | 0.2395 | -2.7447e-06 | 4.1754
I1 | B | 31 | 1 | H | 1.3268 | 0.2373 | -1.6009e-06 | 4.2141 | 1.6881 | 0.243 | -2.899e-06 | 4.1152
I1 | B | 32 | 1 | H | 1.3024 | 0.234 | -1.5702e-06 | 4.2735 | 1.659 | 0.2395 | -2.8307e-06 | 4.1754
I1 | B | 1 | 2 | H | 1.407 | 0.2282 | -1.3413e-06 | 4.3821 | 1.8069 | 0.2339 | -2.5334e-06 | 4.2753
I1 | B | 2 | 2 | H | 1.46 | 0.2282 | -1.2679e-06 | 4.3821 | 1.7866 | 0.2339 | -2.4859e-06 | 4.2753
I1 | B | 3 | 2 | H | 1.3911 | 0.2286 | -1.3607e-06 | 4.3745 | 1.7825 | 0.2342 | -2.6011e-06 | 4.2699
I1 | B | 4 | 2 | H | 1.4106 | 0.2294 | -1.3286e-06 | 4.3592 | 1.706 | 0.235 | -2.5851e-06 | 4.2553
I1 | B | 5 | 2 | H | 1.4323 | 0.2293 | -1.3172e-06 | 4.3611 | 1.7766 | 0.2349 | -2.5327e-06 | 4.2571
I1 | B | 6 | 2 | H | 1.3173 | 0.229 | -1.2278e-06 | 4.3668 | 1.7132 | 0.2342 | -2.4379e-06 | 4.2699
I1 | B | 7 | 2 | H | 1.4323 | 0.2309 | -1.2368e-06 | 4.3309 | 1.7142 | 0.2365 | -2.4961e-06 | 4.2283
I1 | B | 8 | 2 | H | 1.3455 | 0.2297 | -1.2603e-06 | 4.3535 | 1.6752 | 0.235 | -2.4791e-06 | 4.2553
I1 | B | 9 | 2 | H | 1.3751 | 0.229 | -1.2498e-06 | 4.3668 | 1.8486 | 0.2342 | -2.4146e-06 | 4.2699
I1 | B | 10 | 2 | H | 1.3704 | 0.2293 | -1.2598e-06 | 4.3611 | 1.6262 | 0.2347 | -2.5096e-06 | 4.2608
I1 | B | 11 | 2 | H | 1.456 | 0.2334 | -1.3888e-06 | 4.2845 | 1.7557 | 0.2391 | -2.6873e-06 | 4.1824
I1 | B | 12 | 2 | H | 1.3663 | 0.2335 | -1.4184e-06 | 4.2827 | 1.7889 | 0.2389 | -2.6673e-06 | 4.1859
I1 | B | 13 | 2 | H | 1.3625 | 0.2321 | -1.4389e-06 | 4.3085 | 1.8146 | 0.2376 | -2.6488e-06 | 4.2088
I1 | B | 14 | 2 | H | 1.6366 | 0.2324 | -1.5498e-06 | 4.3029 | 1.9106 | 0.2379 | -2.7926e-06 | 4.2034
I1 | B | 15 | 2 | H | 1.4446 | 0.234 | -1.4865e-06 | 4.2735 | 1.8252 | 0.2397 | -2.7898e-06 | 4.1719
I1 | B | 16 | 2 | H | 1.427 | 0.235 | -1.5108e-06 | 4.2553 | 1.7588 | 0.2406 | -2.8247e-06 | 4.1563
I1 | B | 17 | 2 | H | 1.4809 | 0.2332 | -1.4566e-06 | 4.2882 | 1.816 | 0.2387 | -2.7064e-06 | 4.1894
I1 | B | 18 | 2 | H | 1.4829 | 0.2317 | -1.4126e-06 | 4.3159 | 1.8 | 0.2371 | -2.6474e-06 | 4.2176
I1 | B | 19 | 2 | H | 1.5056 | 0.2306 | -1.344e-06 | 4.3365 | 1.7806 | 0.2362 | -2.5941e-06 | 4.2337
I1 | B | 20 | 2 | H | 1.4113 | 0.2299 | -1.246e-06 | 4.3497 | 1.73 | 0.2354 | -2.495e-06 | 4.2481
I1 | B | 21 | 2 | H | 1.4687 | 0.2314 | -1.3104e-06 | 4.3215 | 1.8279 | 0.2369 | -2.5641e-06 | 4.2212
I1 | B | 22 | 2 | H | 1.1566 | 0.231 | -1.1466e-06 | 4.329 | 1.4796 | 0.2363 | -2.3567e-06 | 4.2319
I1 | B | 23 | 2 | H | 1.433 | 0.2335 | -1.4352e-06 | 4.2827 | 1.7521 | 0.2392 | -2.7562e-06 | 4.1806
I1 | B | 24 | 2 | H | 1.445 | 0.2367 | -1.5781e-06 | 4.2248 | 1.7858 | 0.2422 | -2.8443e-06 | 4.1288
I1 | B | 25 | 2 | H | 1.4744 | 0.2373 | -1.55e-06 | 4.2141 | 1.8016 | 0.2431 | -2.8735e-06 | 4.1135
I1 | B | 26 | 2 | H | 1.4703 | 0.2361 | -1.5892e-06 | 4.2355 | 1.9207 | 0.2415 | -2.8225e-06 | 4.1408
I1 | B | 27 | 2 | H | 1.4692 | 0.2359 | -1.5956e-06 | 4.2391 | 1.8516 | 0.2416 | -2.8404e-06 | 4.1391
I1 | B | 28 | 2 | H | 1.3662 | 0.2324 | -1.4732e-06 | 4.3029 | 1.7415 | 0.2379 | -2.7347e-06 | 4.2034
I1 | B | 29 | 2 | H | 1.4504 | 0.2328 | -1.4884e-06 | 4.2955 | 1.9161 | 0.2383 | -2.6965e-06 | 4.1964
I1 | B | 30 | 2 | H | 1.3956 | 0.234 | -1.4926e-06 | 4.2735 | 1.7427 | 0.2396 | -2.7749e-06 | 4.1736
I1 | B | 31 | 2 | H | 1.4196 | 0.2374 | -1.6355e-06 | 4.2123 | 1.8437 | 0.243 | -2.9116e-06 | 4.1152
I1 | B | 32 | 2 | H | 1.449 | 0.234 | -1.5978e-06 | 4.2735 | 1.7557 | 0.2396 | -2.8449e-06 | 4.1736
I2 | A | 1 | 1 | H | -0.1866 | 0.113 | -4.237e-07 | 8.8496 | -0.9144 | 0.1175 | -1.3797e-06 | 8.5106
I2 | A | 2 | 1 | H | -0.2244 | 0.1139 | -4.61e-07 | 8.7796 | -0.9172 | 0.1179 | -1.4303e-06 | 8.4818
I2 | A | 3 | 1 | H | -0.2081 | 0.1144 | -4.842e-07 | 8.7413 | -0.9166 | 0.1188 | -1.4617e-06 | 8.4175
I2 | A | 4 | 1 | H | -0.2081 | 0.1134 | -4.977e-07 | 8.8183 | -0.9183 | 0.1174 | -1.4326e-06 | 8.5179
I2 | A | 5 | 1 | H | -0.2276 | 0.1134 | -4.717e-07 | 8.8183 | -0.9598 | 0.1178 | -1.4042e-06 | 8.489
I2 | A | 6 | 1 | H | -0.228 | 0.1133 | -4.678e-07 | 8.8261 | -0.9232 | 0.1171 | -1.3982e-06 | 8.5397
I2 | A | 7 | 1 | H | -0.2362 | 0.1136 | -4.691e-07 | 8.8028 | -0.9658 | 0.118 | -1.4288e-06 | 8.4746
I2 | A | 8 | 1 | H | -0.2788 | 0.1114 | -4.441e-07 | 8.9767 | -0.9574 | 0.1152 | -1.3392e-06 | 8.6806
I2 | A | 9 | 1 | H | -0.2738 | 0.1114 | -4.309e-07 | 8.9767 | -0.942 | 0.1156 | -1.313e-06 | 8.6505
I2 | A | 10 | 1 | H | -0.281 | 0.1113 | -4.4e-07 | 8.9847 | -0.9396 | 0.1151 | -1.3381e-06 | 8.6881
I2 | A | 11 | 1 | H | -0.2522 | 0.1133 | -4.814e-07 | 8.8261 | -0.9604 | 0.1176 | -1.3979e-06 | 8.5034
I2 | A | 12 | 1 | H | -0.2874 | 0.1132 | -4.896e-07 | 8.8339 | -0.9403 | 0.1171 | -1.3946e-06 | 8.5397
I2 | A | 13 | 1 | H | -0.1517 | 0.1141 | -5e-07 | 8.7642 | -0.8666 | 0.1185 | -1.4477e-06 | 8.4388
I2 | A | 14 | 1 | H | -0.1656 | 0.115 | -5.527e-07 | 8.6957 | -0.8477 | 0.1189 | -1.4755e-06 | 8.4104
I2 | A | 15 | 1 | H | -0.2224 | 0.1157 | -4.822e-07 | 8.643 | -0.926 | 0.1202 | -1.4484e-06 | 8.3195
I2 | A | 16 | 1 | H | -0.2612 | 0.1159 | -4.903e-07 | 8.6281 | -0.9265 | 0.1199 | -1.4637e-06 | 8.3403
I2 | A | 17 | 1 | H | -0.2334 | 0.1166 | -4.707e-07 | 8.5763 | -0.9266 | 0.121 | -1.4363e-06 | 8.2645
I2 | A | 18 | 1 | H | -0.2627 | 0.1152 | -4.711e-07 | 8.6806 | -0.9434 | 0.1191 | -1.4017e-06 | 8.3963
I2 | A | 19 | 1 | H | -0.262 | 0.1159 | -4.479e-07 | 8.6281 | -0.946 | 0.1202 | -1.3902e-06 | 8.3195
I2 | A | 20 | 1 | H | -0.2953 | 0.1154 | -4.552e-07 | 8.6655 | -0.9583 | 0.1192 | -1.3963e-06 | 8.3893
I2 | A | 21 | 1 | H | -0.2789 | 0.1166 | -4.249e-07 | 8.5763 | -0.9596 | 0.1209 | -1.3875e-06 | 8.2713
I2 | A | 22 | 1 | H | -0.3286 | 0.1155 | -3.73e-07 | 8.658 | -0.981 | 0.1193 | -1.2879e-06 | 8.3822
I2 | A | 23 | 1 | H | -0.2831 | 0.1177 | -4.597e-07 | 8.4962 | -0.9633 | 0.1222 | -1.4798e-06 | 8.1833
I2 | A | 24 | 1 | H | -0.3125 | 0.1177 | -4.472e-07 | 8.4962 | -0.9397 | 0.1217 | -1.4192e-06 | 8.2169
I2 | A | 25 | 1 | H | -0.277 | 0.1193 | -4.836e-07 | 8.3822 | -0.9569 | 0.1238 | -1.4756e-06 | 8.0775
I2 | A | 26 | 1 | H | -0.2767 | 0.119 | -5.048e-07 | 8.4034 | -0.9394 | 0.123 | -1.4746e-06 | 8.1301
I2 | A | 27 | 1 | H | -0.276 | 0.1168 | -4.431e-07 | 8.5616 | -0.9265 | 0.1213 | -1.4479e-06 | 8.244
I2 | A | 28 | 1 | H | -0.2991 | 0.1165 | -4.726e-07 | 8.5837 | -0.9652 | 0.1203 | -1.391e-06 | 8.3126
I2 | A | 29 | 1 | H | -0.2845 | 0.1173 | -4.703e-07 | 8.5251 | -0.9941 | 0.1217 | -1.4273e-06 | 8.2169
I2 | A | 30 | 1 | H | -0.3542 | 0.1166 | -4.376e-07 | 8.5763 | -1.0127 | 0.1205 | -1.3881e-06 | 8.2988
I2 | A | 31 | 1 | H | -0.297 | 0.1206 | -4.787e-07 | 8.2919 | -0.9873 | 0.1251 | -1.4959e-06 | 7.9936
I2 | A | 32 | 1 | H | -0.318 | 0.1201 | -4.992e-07 | 8.3264 | -0.9665 | 0.1242 | -1.5229e-06 | 8.0515
I2 | A | 1 | 2 | H | -0.2035 | 0.113 | -4.325e-07 | 8.8496 | -0.9113 | 0.1175 | -1.3782e-06 | 8.5106
I2 | A | 2 | 2 | H | -0.2367 | 0.1139 | -4.657e-07 | 8.7796 | -0.9161 | 0.1179 | -1.4261e-06 | 8.4818
I2 | A | 3 | 2 | H | -0.176 | 0.1144 | -4.847e-07 | 8.7413 | -0.8987 | 0.1188 | -1.4623e-06 | 8.4175
I2 | A | 4 | 2 | H | -0.2267 | 0.1134 | -5.083e-07 | 8.8183 | -0.9147 | 0.1174 | -1.4338e-06 | 8.5179
I2 | A | 5 | 2 | H | -0.2414 | 0.1134 | -4.688e-07 | 8.8183 | -0.9476 | 0.1178 | -1.4073e-06 | 8.489
I2 | A | 6 | 2 | H | -0.229 | 0.1133 | -4.771e-07 | 8.8261 | -0.9309 | 0.1172 | -1.4041e-06 | 8.5324
I2 | A | 7 | 2 | H | -0.2376 | 0.1136 | -4.754e-07 | 8.8028 | -0.9641 | 0.118 | -1.4274e-06 | 8.4746
I2 | A | 8 | 2 | H | -0.2747 | 0.1114 | -4.409e-07 | 8.9767 | -0.9591 | 0.1152 | -1.3389e-06 | 8.6806
I2 | A | 9 | 2 | H | -0.2681 | 0.1114 | -4.298e-07 | 8.9767 | -0.9447 | 0.1156 | -1.3142e-06 | 8.6505
I2 | A | 10 | 2 | H | -0.2867 | 0.1113 | -4.373e-07 | 8.9847 | -0.9453 | 0.1151 | -1.3301e-06 | 8.6881
I2 | A | 11 | 2 | H | -0.2502 | 0.1133 | -4.816e-07 | 8.8261 | -0.9489 | 0.1176 | -1.4009e-06 | 8.5034
I2 | A | 12 | 2 | H | -0.289 | 0.1132 | -4.871e-07 | 8.8339 | -0.9328 | 0.1171 | -1.3966e-06 | 8.5397
I2 | A | 13 | 2 | H | -0.1903 | 0.1141 | -4.919e-07 | 8.7642 | -0.8932 | 0.1185 | -1.4333e-06 | 8.4388
I2 | A | 14 | 2 | H | -0.1992 | 0.115 | -5.433e-07 | 8.6957 | -0.8789 | 0.1189 | -1.462e-06 | 8.4104
I2 | A | 15 | 2 | H | -0.1696 | 0.1157 | -4.725e-07 | 8.643 | -0.859 | 0.1202 | -1.4723e-06 | 8.3195
I2 | A | 16 | 2 | H | -0.2293 | 0.116 | -4.893e-07 | 8.6207 | -0.8831 | 0.1199 | -1.444e-06 | 8.3403
I2 | A | 17 | 2 | H | -0.2222 | 0.1166 | -4.724e-07 | 8.5763 | -0.9084 | 0.121 | -1.4354e-06 | 8.2645
I2 | A | 18 | 2 | H | -0.2447 | 0.1152 | -4.681e-07 | 8.6806 | -0.9078 | 0.1191 | -1.4035e-06 | 8.3963
I2 | A | 19 | 2 | H | -0.2522 | 0.1159 | -4.525e-07 | 8.6281 | -0.9223 | 0.1203 | -1.4017e-06 | 8.3126
I2 | A | 20 | 2 | H | -0.2785 | 0.1154 | -4.58e-07 | 8.6655 | -0.9174 | 0.1192 | -1.39e-06 | 8.3893
I2 | A | 21 | 2 | H | -0.3232 | 0.1165 | -4.055e-07 | 8.5837 | -0.9867 | 0.1209 | -1.358e-06 | 8.2713
I2 | A | 22 | 2 | H | -0.3765 | 0.1155 | -3.692e-07 | 8.658 | -1.0098 | 0.1193 | -1.2697e-06 | 8.3822
I2 | A | 23 | 2 | H | -0.2438 | 0.1177 | -4.585e-07 | 8.4962 | -0.9458 | 0.1222 | -1.4627e-06 | 8.1833
I2 | A | 24 | 2 | H | -0.2748 | 0.1177 | -4.434e-07 | 8.4962 | -0.9568 | 0.1217 | -1.4168e-06 | 8.2169
I2 | A | 25 | 2 | H | -0.2393 | 0.1193 | -4.809e-07 | 8.3822 | -0.9244 | 0.1238 | -1.4752e-06 | 8.0775
I2 | A | 26 | 2 | H | -0.259 | 0.119 | -5.047e-07 | 8.4034 | -0.9227 | 0.123 | -1.4711e-06 | 8.1301
I2 | A | 27 | 2 | H | -0.2575 | 0.1169 | -4.506e-07 | 8.5543 | -0.8792 | 0.1212 | -1.4155e-06 | 8.2508
I2 | A | 28 | 2 | H | -0.2779 | 0.1165 | -4.678e-07 | 8.5837 | -0.9593 | 0.1204 | -1.4041e-06 | 8.3056
I2 | A | 29 | 2 | H | -0.2795 | 0.1173 | -4.69e-07 | 8.5251 | -0.9815 | 0.1217 | -1.4334e-06 | 8.2169
I2 | A | 30 | 2 | H | -0.3337 | 0.1166 | -4.362e-07 | 8.5763 | -1.005 | 0.1205 | -1.3966e-06 | 8.2988
I2 | A | 31 | 2 | H | -0.303 | 0.1206 | -4.905e-07 | 8.2919 | -0.9732 | 0.1251 | -1.4988e-06 | 7.9936
I2 | A | 32 | 2 | H | -0.3083 | 0.1202 | -5.026e-07 | 8.3195 | -0.951 | 0.1242 | -1.5167e-06 | 8.0515
I2 | B | 1 | 1 | H | -0.1913 | 0.1131 | -4.303e-07 | 8.8417 | -0.9168 | 0.1176 | -1.3862e-06 | 8.5034
I2 | B | 2 | 1 | H | -0.227 | 0.114 | -4.659e-07 | 8.7719 | -0.9177 | 0.118 | -1.4343e-06 | 8.4746
I2 | B | 3 | 1 | H | -0.205 | 0.1144 | -4.855e-07 | 8.7413 | -0.9178 | 0.1189 | -1.4603e-06 | 8.4104
I2 | B | 4 | 1 | H | -0.2139 | 0.1135 | -5.048e-07 | 8.8106 | -0.9199 | 0.1174 | -1.4383e-06 | 8.5179
I2 | B | 5 | 1 | H | -0.2346 | 0.1135 | -4.738e-07 | 8.8106 | -0.9494 | 0.1178 | -1.3946e-06 | 8.489
I2 | B | 6 | 1 | H | -0.2343 | 0.1134 | -4.785e-07 | 8.8183 | -0.9329 | 0.1172 | -1.4111e-06 | 8.5324
I2 | B | 7 | 1 | H | -0.2383 | 0.1137 | -4.75e-07 | 8.7951 | -0.9767 | 0.1181 | -1.4354e-06 | 8.4674
I2 | B | 8 | 1 | H | -0.2723 | 0.1115 | -4.422e-07 | 8.9686 | -0.9599 | 0.1152 | -1.3436e-06 | 8.6806
I2 | B | 9 | 1 | H | -0.2729 | 0.1115 | -4.321e-07 | 8.9686 | -0.9474 | 0.1157 | -1.322e-06 | 8.643
I2 | B | 10 | 1 | H | -0.2777 | 0.1114 | -4.337e-07 | 8.9767 | -0.9383 | 0.1152 | -1.3358e-06 | 8.6806
I2 | B | 11 | 1 | H | -0.2518 | 0.1134 | -4.851e-07 | 8.8183 | -0.9617 | 0.1176 | -1.397e-06 | 8.5034
I2 | B | 12 | 1 | H | -0.2849 | 0.1133 | -4.859e-07 | 8.8261 | -0.9401 | 0.1171 | -1.3925e-06 | 8.5397
I2 | B | 13 | 1 | H | -0.1517 | 0.1142 | -5.039e-07 | 8.7566 | -0.8667 | 0.1186 | -1.4426e-06 | 8.4317
I2 | B | 14 | 1 | H | -0.1616 | 0.1151 | -5.501e-07 | 8.6881 | -0.8488 | 0.119 | -1.4798e-06 | 8.4034
I2 | B | 15 | 1 | H | -0.2249 | 0.1158 | -4.783e-07 | 8.6356 | -0.9229 | 0.1202 | -1.4437e-06 | 8.3195
I2 | B | 16 | 1 | H | -0.2637 | 0.1161 | -4.953e-07 | 8.6133 | -0.9251 | 0.12 | -1.4596e-06 | 8.3333
I2 | B | 17 | 1 | H | -0.2349 | 0.1167 | -4.73e-07 | 8.569 | -0.9319 | 0.1211 | -1.4422e-06 | 8.2576
I2 | B | 18 | 1 | H | -0.2628 | 0.1153 | -4.708e-07 | 8.673 | -0.9392 | 0.1192 | -1.3999e-06 | 8.3893
I2 | B | 19 | 1 | H | -0.2673 | 0.116 | -4.508e-07 | 8.6207 | -0.9482 | 0.1203 | -1.3934e-06 | 8.3126
I2 | B | 20 | 1 | H | -0.2957 | 0.1155 | -4.578e-07 | 8.658 | -0.9572 | 0.1193 | -1.3983e-06 | 8.3822
I2 | B | 21 | 1 | H | -0.2807 | 0.1167 | -4.259e-07 | 8.569 | -0.961 | 0.121 | -1.3944e-06 | 8.2645
I2 | B | 22 | 1 | H | -0.3269 | 0.1156 | -3.738e-07 | 8.6505 | -0.9822 | 0.1194 | -1.293e-06 | 8.3752
I2 | B | 23 | 1 | H | -0.2821 | 0.1178 | -4.58e-07 | 8.489 | -0.9641 | 0.1223 | -1.4822e-06 | 8.1766
I2 | B | 24 | 1 | H | -0.3142 | 0.1178 | -4.46e-07 | 8.489 | -0.9428 | 0.1218 | -1.4262e-06 | 8.2102
I2 | B | 25 | 1 | H | -0.2808 | 0.1194 | -4.856e-07 | 8.3752 | -0.9607 | 0.1239 | -1.4831e-06 | 8.071
I2 | B | 26 | 1 | H | -0.2739 | 0.1191 | -5.012e-07 | 8.3963 | -0.9452 | 0.123 | -1.4728e-06 | 8.1301
I2 | B | 27 | 1 | H | -0.2814 | 0.1169 | -4.49e-07 | 8.5543 | -0.917 | 0.1213 | -1.4301e-06 | 8.244
I2 | B | 28 | 1 | H | -0.2996 | 0.1166 | -4.755e-07 | 8.5763 | -0.9649 | 0.1204 | -1.3886e-06 | 8.3056
I2 | B | 29 | 1 | H | -0.2862 | 0.1174 | -4.68e-07 | 8.5179 | -0.9932 | 0.1217 | -1.4278e-06 | 8.2169
I2 | B | 30 | 1 | H | -0.3554 | 0.1167 | -4.369e-07 | 8.569 | -1.0151 | 0.1205 | -1.3891e-06 | 8.2988
I2 | B | 31 | 1 | H | -0.3048 | 0.1207 | -4.899e-07 | 8.285 | -0.986 | 0.1251 | -1.4975e-06 | 7.9936
I2 | B | 32 | 1 | H | -0.3184 | 0.1202 | -4.988e-07 | 8.3195 | -0.9704 | 0.1243 | -1.5347e-06 | 8.0451
I2 | B | 1 | 2 | H | -0.2094 | 0.1131 | -4.355e-07 | 8.8417 | -0.9152 | 0.1176 | -1.3908e-06 | 8.5034
I2 | B | 2 | 2 | H | -0.2336 | 0.114 | -4.625e-07 | 8.7719 | -0.9167 | 0.118 | -1.4282e-06 | 8.4746
I2 | B | 3 | 2 | H | -0.1811 | 0.1145 | -4.884e-07 | 8.7336 | -0.8994 | 0.1189 | -1.4637e-06 | 8.4104
I2 | B | 4 | 2 | H | -0.2294 | 0.1135 | -5.099e-07 | 8.8106 | -0.9168 | 0.1175 | -1.4405e-06 | 8.5106
I2 | B | 5 | 2 | H | -0.249 | 0.1135 | -4.798e-07 | 8.8106 | -0.9452 | 0.1179 | -1.408e-06 | 8.4818
I2 | B | 6 | 2 | H | -0.2323 | 0.1134 | -4.788e-07 | 8.8183 | -0.925 | 0.1172 | -1.4005e-06 | 8.5324
I2 | B | 7 | 2 | H | -0.2315 | 0.1137 | -4.728e-07 | 8.7951 | -0.9599 | 0.118 | -1.4183e-06 | 8.4746
I2 | B | 8 | 2 | H | -0.2855 | 0.1115 | -4.534e-07 | 8.9686 | -0.9643 | 0.1152 | -1.341e-06 | 8.6806
I2 | B | 9 | 2 | H | -0.2657 | 0.1115 | -4.287e-07 | 8.9686 | -0.9437 | 0.1157 | -1.3193e-06 | 8.643
I2 | B | 10 | 2 | H | -0.28 | 0.1114 | -4.36e-07 | 8.9767 | -0.9428 | 0.1152 | -1.327e-06 | 8.6806
I2 | B | 11 | 2 | H | -0.2516 | 0.1134 | -4.856e-07 | 8.8183 | -0.9446 | 0.1176 | -1.3958e-06 | 8.5034
I2 | B | 12 | 2 | H | -0.2875 | 0.1133 | -4.828e-07 | 8.8261 | -0.9324 | 0.1172 | -1.3986e-06 | 8.5324
I2 | B | 13 | 2 | H | -0.1838 | 0.1142 | -4.874e-07 | 8.7566 | -0.8987 | 0.1185 | -1.4366e-06 | 8.4388
I2 | B | 14 | 2 | H | -0.2051 | 0.1151 | -5.514e-07 | 8.6881 | -0.8854 | 0.119 | -1.4734e-06 | 8.4034
I2 | B | 15 | 2 | H | -0.1899 | 0.1158 | -4.788e-07 | 8.6356 | -0.8496 | 0.1203 | -1.468e-06 | 8.3126
I2 | B | 16 | 2 | H | -0.2304 | 0.1161 | -4.908e-07 | 8.6133 | -0.8858 | 0.12 | -1.4482e-06 | 8.3333
I2 | B | 17 | 2 | H | -0.2233 | 0.1167 | -4.71e-07 | 8.569 | -0.9143 | 0.1211 | -1.4426e-06 | 8.2576
I2 | B | 18 | 2 | H | -0.2462 | 0.1153 | -4.693e-07 | 8.673 | -0.9107 | 0.1192 | -1.4065e-06 | 8.3893
I2 | B | 19 | 2 | H | -0.2549 | 0.116 | -4.516e-07 | 8.6207 | -0.9197 | 0.1204 | -1.4e-06 | 8.3056
I2 | B | 20 | 2 | H | -0.2799 | 0.1155 | -4.54e-07 | 8.658 | -0.9197 | 0.1193 | -1.3932e-06 | 8.3822
I2 | B | 21 | 2 | H | -0.3265 | 0.1167 | -4.106e-07 | 8.569 | -0.9876 | 0.1209 | -1.3588e-06 | 8.2713
I2 | B | 22 | 2 | H | -0.3815 | 0.1156 | -3.711e-07 | 8.6505 | -1.0158 | 0.1194 | -1.2764e-06 | 8.3752
I2 | B | 23 | 2 | H | -0.2388 | 0.1178 | -4.566e-07 | 8.489 | -0.9426 | 0.1223 | -1.4639e-06 | 8.1766
I2 | B | 24 | 2 | H | -0.2707 | 0.1178 | -4.423e-07 | 8.489 | -0.9639 | 0.1218 | -1.4288e-06 | 8.2102
I2 | B | 25 | 2 | H | -0.2357 | 0.1194 | -4.805e-07 | 8.3752 | -0.9352 | 0.1239 | -1.4883e-06 | 8.071
I2 | B | 26 | 2 | H | -0.2547 | 0.1191 | -5.054e-07 | 8.3963 | -0.9179 | 0.1231 | -1.4729e-06 | 8.1235
I2 | B | 27 | 2 | H | -0.26 | 0.117 | -4.557e-07 | 8.547 | -0.8751 | 0.1213 | -1.4094e-06 | 8.244
I2 | B | 28 | 2 | H | -0.2787 | 0.1166 | -4.725e-07 | 8.5763 | -0.9625 | 0.1205 | -1.4059e-06 | 8.2988
I2 | B | 29 | 2 | H | -0.2797 | 0.1174 | -4.656e-07 | 8.5179 | -0.9808 | 0.1218 | -1.4353e-06 | 8.2102
I2 | B | 30 | 2 | H | -0.334 | 0.1167 | -4.348e-07 | 8.569 | -1.0073 | 0.1206 | -1.3972e-06 | 8.2919
I2 | B | 31 | 2 | H | -0.2946 | 0.1207 | -4.849e-07 | 8.285 | -0.9787 | 0.1252 | -1.5063e-06 | 7.9872
I2 | B | 32 | 2 | H | -0.3016 | 0.1203 | -5.006e-07 | 8.3126 | -0.9518 | 0.1243 | -1.5237e-06 | 8.0451
I3 | A | 1 | 1 | H | 0.0367 | 0.0191 | -5.44e-08 | 52.356 | 0.0728 | 0.0189 | 7.36e-08 | 52.9101
I3 | A | 2 | 1 | H | 0.0406 | 0.0187 | -4.22e-08 | 53.4759 | 0.0765 | 0.0184 | 8.03e-08 | 54.3478
I3 | A | 3 | 1 | H | 0.037 | 0.0186 | -5.16e-08 | 53.7634 | 0.0732 | 0.0183 | 7.01e-08 | 54.6448
I3 | A | 4 | 1 | H | -0.242 | 0.0351 | -3.68e-08 | 28.49 | -0.1031 | 0.0419 | 7.311e-07 | 23.8663
I3 | A | 5 | 1 | H | 0.0437 | 0.0185 | -3.29e-08 | 54.0541 | 0.0822 | 0.0182 | 8.53e-08 | 54.9451
I3 | A | 6 | 1 | H | 0.0466 | 0.0184 | -3.72e-08 | 54.3478 | 0.0784 | 0.0182 | 8.26e-08 | 54.9451
I3 | A | 7 | 1 | H | 0.0421 | 0.0184 | -4.31e-08 | 54.3478 | 0.0822 | 0.0181 | 7.47e-08 | 55.2486
I3 | A | 8 | 1 | H | 0.05 | 0.0184 | -4.71e-08 | 54.3478 | 0.0863 | 0.0181 | 7.34e-08 | 55.2486
I3 | A | 9 | 1 | H | 0.0518 | 0.0185 | -4.96e-08 | 54.0541 | 0.0913 | 0.0182 | 7.1e-08 | 54.9451
I3 | A | 10 | 1 | H | 0.051 | 0.0187 | -5e-08 | 53.4759 | 0.0842 | 0.0184 | 7.09e-08 | 54.3478
I3 | A | 11 | 1 | H | 0.0476 | 0.0186 | -4.12e-08 | 53.7634 | 0.0877 | 0.0183 | 8.25e-08 | 54.6448
I3 | A | 12 | 1 | H | 0.0539 | 0.0184 | -3.66e-08 | 54.3478 | 0.0922 | 0.0181 | 8.28e-08 | 55.2486
I3 | A | 13 | 1 | H | 0.0555 | 0.0183 | -4.38e-08 | 54.6448 | 0.0957 | 0.018 | 7.75e-08 | 55.5556
I3 | A | 14 | 1 | H | 0.049 | 0.0186 | -5.51e-08 | 53.7634 | 0.0862 | 0.0183 | 6.98e-08 | 54.6448
I3 | A | 15 | 1 | H | 0.0542 | 0.0185 | -4.72e-08 | 54.0541 | 0.0934 | 0.0181 | 7.35e-08 | 55.2486
I3 | A | 16 | 1 | H | 0.0537 | 0.0186 | -3.98e-08 | 53.7634 | 0.091 | 0.0183 | 8.39e-08 | 54.6448
I3 | A | 17 | 1 | H | 0.0616 | 0.0187 | -3.4e-08 | 53.4759 | 0.1016 | 0.0184 | 9.15e-08 | 54.3478
I3 | A | 18 | 1 | H | 0.0698 | 0.0186 | -2.84e-08 | 53.7634 | 0.1059 | 0.0183 | 9.37e-08 | 54.6448
I3 | A | 19 | 1 | H | 0.0689 | 0.0187 | -4.03e-08 | 53.4759 | 0.11 | 0.0184 | 8.46e-08 | 54.3478
I3 | A | 20 | 1 | H | 0.0655 | 0.0188 | -4.26e-08 | 53.1915 | 0.1028 | 0.0184 | 8.38e-08 | 54.3478
I3 | A | 21 | 1 | H | 0.0564 | 0.0188 | -5.98e-08 | 53.1915 | 0.0965 | 0.0184 | 6.75e-08 | 54.3478
I3 | A | 22 | 1 | H | 0.0672 | 0.0189 | -5.38e-08 | 52.9101 | 0.1015 | 0.0186 | 7.28e-08 | 53.7634
I3 | A | 23 | 1 | H | 0.0666 | 0.0188 | -3.79e-08 | 53.1915 | 0.1017 | 0.0185 | 8.81e-08 | 54.0541
I3 | A | 24 | 1 | H | 0.0611 | 0.0189 | -4.66e-08 | 52.9101 | 0.0956 | 0.0186 | 7.98e-08 | 53.7634
I3 | A | 25 | 1 | H | 0.0661 | 0.0189 | -4.58e-08 | 52.9101 | 0.0971 | 0.0186 | 7.79e-08 | 53.7634
I3 | A | 26 | 1 | H | 0.0658 | 0.0189 | -4.49e-08 | 52.9101 | 0.0999 | 0.0185 | 8.18e-08 | 54.0541
I3 | A | 27 | 1 | H | 0.0602 | 0.0189 | -4.37e-08 | 52.9101 | 0.0905 | 0.0186 | 8.02e-08 | 53.7634
I3 | A | 28 | 1 | H | 0.0573 | 0.0189 | -5.01e-08 | 52.9101 | 0.0963 | 0.0186 | 7.94e-08 | 53.7634
I3 | A | 29 | 1 | H | 0.0597 | 0.019 | -4.79e-08 | 52.6316 | 0.0927 | 0.0187 | 8.22e-08 | 53.4759
I3 | A | 30 | 1 | H | 0.0686 | 0.0191 | -3.92e-08 | 52.356 | 0.1083 | 0.0188 | 9.02e-08 | 53.1915
I3 | A | 31 | 1 | H | 0.0682 | 0.0191 | -4.06e-08 | 52.356 | 0.1129 | 0.0187 | 9.13e-08 | 53.4759
I3 | A | 32 | 1 | H | 0.0771 | 0.0192 | -4e-08 | 52.0833 | 0.1126 | 0.0189 | 9.27e-08 | 52.9101
I3 | A | 1 | 2 | H | 0.0291 | 0.0191 | -5.35e-08 | 52.356 | 0.0552 | 0.0188 | 7.48e-08 | 53.1915
I3 | A | 2 | 2 | H | 0.0335 | 0.0187 | -4.2e-08 | 53.4759 | 0.0592 | 0.0184 | 8.06e-08 | 54.3478
I3 | A | 3 | 2 | H | 0.027 | 0.0186 | -5.13e-08 | 53.7634 | 0.0545 | 0.0183 | 7.1e-08 | 54.6448
I3 | A | 4 | 2 | H | 0.2862 | 0.0351 | -9.54e-08 | 28.49 | 0.3293 | 0.0416 | 9.766e-07 | 24.0385
I3 | A | 5 | 2 | H | 0.037 | 0.0185 | -3.26e-08 | 54.0541 | 0.0697 | 0.0182 | 8.77e-08 | 54.9451
I3 | A | 6 | 2 | H | 0.0354 | 0.0184 | -3.88e-08 | 54.3478 | 0.063 | 0.0182 | 8.15e-08 | 54.9451
I3 | A | 7 | 2 | H | 0.0349 | 0.0184 | -4.22e-08 | 54.3478 | 0.0663 | 0.0181 | 7.57e-08 | 55.2486
I3 | A | 8 | 2 | H | 0.0421 | 0.0184 | -4.72e-08 | 54.3478 | 0.0695 | 0.0181 | 7.35e-08 | 55.2486
I3 | A | 9 | 2 | H | 0.0441 | 0.0185 | -4.96e-08 | 54.0541 | 0.0719 | 0.0182 | 7.17e-08 | 54.9451
I3 | A | 10 | 2 | H | 0.0418 | 0.0187 | -5.08e-08 | 53.4759 | 0.0678 | 0.0184 | 7.11e-08 | 54.3478
I3 | A | 11 | 2 | H | 0.0406 | 0.0186 | -4.07e-08 | 53.7634 | 0.0705 | 0.0183 | 8.29e-08 | 54.6448
I3 | A | 12 | 2 | H | 0.0456 | 0.0184 | -3.72e-08 | 54.3478 | 0.078 | 0.0181 | 8.31e-08 | 55.2486
I3 | A | 13 | 2 | H | 0.0478 | 0.0183 | -4.31e-08 | 54.6448 | 0.0776 | 0.018 | 7.68e-08 | 55.5556
I3 | A | 14 | 2 | H | 0.0428 | 0.0186 | -5.4e-08 | 53.7634 | 0.0684 | 0.0183 | 6.91e-08 | 54.6448
I3 | A | 15 | 2 | H | 0.0455 | 0.0184 | -4.63e-08 | 54.3478 | 0.0775 | 0.0181 | 7.38e-08 | 55.2486
I3 | A | 16 | 2 | H | 0.0451 | 0.0186 | -4.12e-08 | 53.7634 | 0.0745 | 0.0183 | 8.32e-08 | 54.6448
I3 | A | 17 | 2 | H | 0.0528 | 0.0187 | -3.42e-08 | 53.4759 | 0.087 | 0.0184 | 9.16e-08 | 54.3478
I3 | A | 18 | 2 | H | 0.0582 | 0.0186 | -2.95e-08 | 53.7634 | 0.0901 | 0.0183 | 9.35e-08 | 54.6448
I3 | A | 19 | 2 | H | 0.0605 | 0.0187 | -3.96e-08 | 53.4759 | 0.0935 | 0.0184 | 8.46e-08 | 54.3478
I3 | A | 20 | 2 | H | 0.0549 | 0.0188 | -4.39e-08 | 53.1915 | 0.0864 | 0.0184 | 8.42e-08 | 54.3478
I3 | A | 21 | 2 | H | 0.0443 | 0.0188 | -5.92e-08 | 53.1915 | 0.0793 | 0.0184 | 6.85e-08 | 54.3478
I3 | A | 22 | 2 | H | 0.0504 | 0.0189 | -5.67e-08 | 52.9101 | 0.0819 | 0.0186 | 7.28e-08 | 53.7634
I3 | A | 23 | 2 | H | 0.0573 | 0.0189 | -3.86e-08 | 52.9101 | 0.0888 | 0.0185 | 8.99e-08 | 54.0541
I3 | A | 24 | 2 | H | 0.0526 | 0.0189 | -4.79e-08 | 52.9101 | 0.0784 | 0.0186 | 7.92e-08 | 53.7634
I3 | A | 25 | 2 | H | 0.0589 | 0.0189 | -4.5e-08 | 52.9101 | 0.0826 | 0.0186 | 7.79e-08 | 53.7634
I3 | A | 26 | 2 | H | 0.0564 | 0.0189 | -4.61e-08 | 52.9101 | 0.0856 | 0.0185 | 8.13e-08 | 54.0541
I3 | A | 27 | 2 | H | 0.0548 | 0.0189 | -4.32e-08 | 52.9101 | 0.0837 | 0.0186 | 8.28e-08 | 53.7634
I3 | A | 28 | 2 | H | 0.0536 | 0.0189 | -4.96e-08 | 52.9101 | 0.0842 | 0.0186 | 7.96e-08 | 53.7634
I3 | A | 29 | 2 | H | 0.054 | 0.019 | -4.65e-08 | 52.6316 | 0.0837 | 0.0186 | 8.44e-08 | 53.7634
I3 | A | 30 | 2 | H | 0.0587 | 0.0191 | -4.06e-08 | 52.356 | 0.096 | 0.0188 | 9.16e-08 | 53.1915
I3 | A | 31 | 2 | H | 0.0619 | 0.0191 | -3.92e-08 | 52.356 | 0.1002 | 0.0187 | 9.17e-08 | 53.4759
I3 | A | 32 | 2 | H | 0.0688 | 0.0192 | -4.03e-08 | 52.0833 | 0.0979 | 0.0189 | 9.21e-08 | 52.9101
I3 | A | 1 | 1 | H | 0.0358 | 0.0192 | -5.49e-08 | 52.0833 | 0.0736 | 0.0189 | 7.44e-08 | 52.9101
I3 | A | 2 | 1 | H | 0.0392 | 0.0187 | -4.2e-08 | 53.4759 | 0.076 | 0.0185 | 8.07e-08 | 54.0541
I3 | B | 3 | 1 | H | 0.0343 | 0.0186 | -5.26e-08 | 53.7634 | 0.0733 | 0.0184 | 7.11e-08 | 54.3478
I3 | B | 4 | 1 | H | -0.2423 | 0.0352 | -2.73e-08 | 28.4091 | 0.0075 | 0.0414 | 1.1588e-06 | 24.1546
I3 | B | 5 | 1 | H | 0.0443 | 0.0185 | -3.34e-08 | 54.0541 | 0.0825 | 0.0182 | 8.54e-08 | 54.9451
I3 | B | 6 | 1 | H | 0.0423 | 0.0185 | -3.85e-08 | 54.0541 | 0.0781 | 0.0182 | 8.14e-08 | 54.9451
I3 | B | 7 | 1 | H | 0.0421 | 0.0184 | -4.37e-08 | 54.3478 | 0.08 | 0.0181 | 7.4e-08 | 55.2486
I3 | B | 8 | 1 | H | 0.0503 | 0.0184 | -4.75e-08 | 54.3478 | 0.0851 | 0.0181 | 7.34e-08 | 55.2486
I3 | B | 9 | 1 | H | 0.052 | 0.0185 | -5.03e-08 | 54.0541 | 0.0899 | 0.0182 | 7e-08 | 54.9451
I3 | B | 10 | 1 | H | 0.0469 | 0.0187 | -5.21e-08 | 53.4759 | 0.0859 | 0.0184 | 7.21e-08 | 54.3478
I3 | B | 11 | 1 | H | 0.0462 | 0.0187 | -4.19e-08 | 53.4759 | 0.0894 | 0.0183 | 8.36e-08 | 54.6448
I3 | B | 12 | 1 | H | 0.0557 | 0.0185 | -3.66e-08 | 54.0541 | 0.0922 | 0.0181 | 8.28e-08 | 55.2486
I3 | B | 13 | 1 | H | 0.0555 | 0.0183 | -4.4e-08 | 54.6448 | 0.0941 | 0.018 | 7.67e-08 | 55.5556
I3 | B | 14 | 1 | H | 0.0501 | 0.0187 | -5.42e-08 | 53.4759 | 0.0864 | 0.0183 | 6.94e-08 | 54.6448
I3 | B | 15 | 1 | H | 0.0544 | 0.0185 | -4.75e-08 | 54.0541 | 0.0951 | 0.0181 | 7.44e-08 | 55.2486
I3 | B | 16 | 1 | H | 0.0528 | 0.0186 | -4.14e-08 | 53.7634 | 0.0907 | 0.0183 | 8.34e-08 | 54.6448
I3 | B | 17 | 1 | H | 0.0587 | 0.0188 | -3.56e-08 | 53.1915 | 0.1034 | 0.0184 | 9.15e-08 | 54.3478
I3 | B | 18 | 1 | H | 0.0688 | 0.0187 | -2.9e-08 | 53.4759 | 0.1063 | 0.0183 | 9.41e-08 | 54.6448
I3 | B | 19 | 1 | H | 0.0668 | 0.0187 | -4.13e-08 | 53.4759 | 0.1103 | 0.0184 | 8.42e-08 | 54.3478
I3 | B | 20 | 1 | H | 0.0636 | 0.0188 | -4.38e-08 | 53.1915 | 0.1042 | 0.0185 | 8.34e-08 | 54.0541
I3 | B | 21 | 1 | H | 0.0561 | 0.0188 | -5.92e-08 | 53.1915 | 0.0955 | 0.0185 | 6.68e-08 | 54.0541
I3 | B | 22 | 1 | H | 0.0642 | 0.019 | -5.56e-08 | 52.6316 | 0.1029 | 0.0186 | 7.37e-08 | 53.7634
I3 | B | 23 | 1 | H | 0.0677 | 0.0189 | -3.93e-08 | 52.9101 | 0.1045 | 0.0185 | 8.81e-08 | 54.0541
I3 | B | 24 | 1 | H | 0.0641 | 0.0189 | -4.65e-08 | 52.9101 | 0.0968 | 0.0186 | 7.98e-08 | 53.7634
I3 | B | 25 | 1 | H | 0.0638 | 0.0189 | -4.66e-08 | 52.9101 | 0.0988 | 0.0186 | 7.87e-08 | 53.7634
I3 | B | 26 | 1 | H | 0.0676 | 0.0189 | -4.5e-08 | 52.9101 | 0.0998 | 0.0186 | 8.28e-08 | 53.7634
I3 | B | 27 | 1 | H | 0.0605 | 0.0189 | -4.46e-08 | 52.9101 | 0.0922 | 0.0186 | 8.15e-08 | 53.7634
I3 | B | 28 | 1 | H | 0.0564 | 0.019 | -5.12e-08 | 52.6316 | 0.0993 | 0.0187 | 8.11e-08 | 53.4759
I3 | B | 29 | 1 | H | 0.0637 | 0.019 | -4.64e-08 | 52.6316 | 0.0943 | 0.0187 | 8.31e-08 | 53.4759
I3 | B | 30 | 1 | H | 0.0693 | 0.0191 | -3.96e-08 | 52.356 | 0.11 | 0.0188 | 9.18e-08 | 53.1915
I3 | B | 31 | 1 | H | 0.0689 | 0.0191 | -4.03e-08 | 52.356 | 0.1142 | 0.0188 | 9.08e-08 | 53.1915
I3 | B | 32 | 1 | H | 0.0769 | 0.0192 | -4.11e-08 | 52.0833 | 0.1132 | 0.0189 | 9.36e-08 | 52.9101
I3 | B | 1 | 2 | H | 0.0286 | 0.0192 | -5.4e-08 | 52.0833 | 0.054 | 0.0189 | 7.49e-08 | 52.9101
I3 | B | 2 | 2 | H | 0.0302 | 0.0187 | -4.35e-08 | 53.4759 | 0.0569 | 0.0185 | 8.05e-08 | 54.0541
I3 | B | 3 | 2 | H | 0.0254 | 0.0186 | -5.21e-08 | 53.7634 | 0.0536 | 0.0184 | 7.16e-08 | 54.3478
I3 | B | 4 | 2 | H | 0.248 | 0.0352 | -7.41e-08 | 28.4091 | 0.3099 | 0.0417 | 9.563e-07 | 23.9808
I3 | B | 5 | 2 | H | 0.0346 | 0.0185 | -3.32e-08 | 54.0541 | 0.0671 | 0.0182 | 8.6e-08 | 54.9451
I3 | B | 6 | 2 | H | 0.0357 | 0.0185 | -3.88e-08 | 54.0541 | 0.0594 | 0.0182 | 8.04e-08 | 54.9451
I3 | B | 7 | 2 | H | 0.0334 | 0.0184 | -4.3e-08 | 54.3478 | 0.0654 | 0.0181 | 7.47e-08 | 55.2486
I3 | B | 8 | 2 | H | 0.0422 | 0.0184 | -4.73e-08 | 54.3478 | 0.0668 | 0.0181 | 7.26e-08 | 55.2486
I3 | B | 9 | 2 | H | 0.044 | 0.0185 | -5.01e-08 | 54.0541 | 0.0743 | 0.0182 | 7.21e-08 | 54.9451
I3 | B | 10 | 2 | H | 0.0375 | 0.0187 | -5.29e-08 | 53.4759 | 0.0648 | 0.0184 | 7.02e-08 | 54.3478
I3 | B | 11 | 2 | H | 0.0419 | 0.0186 | -3.99e-08 | 53.7634 | 0.0733 | 0.0183 | 8.39e-08 | 54.6448
I3 | B | 12 | 2 | H | 0.0481 | 0.0185 | -3.59e-08 | 54.0541 | 0.074 | 0.0181 | 8.18e-08 | 55.2486
I3 | B | 13 | 2 | H | 0.0481 | 0.0183 | -4.32e-08 | 54.6448 | 0.0753 | 0.018 | 7.63e-08 | 55.5556
I3 | B | 14 | 2 | H | 0.0417 | 0.0187 | -5.49e-08 | 53.4759 | 0.0666 | 0.0184 | 6.87e-08 | 54.3478
I3 | B | 15 | 2 | H | 0.0442 | 0.0185 | -4.79e-08 | 54.0541 | 0.0751 | 0.0181 | 7.37e-08 | 55.2486
I3 | B | 16 | 2 | H | 0.044 | 0.0186 | -4.19e-08 | 53.7634 | 0.0733 | 0.0183 | 8.27e-08 | 54.6448
I3 | B | 17 | 2 | H | 0.0509 | 0.0188 | -3.43e-08 | 53.1915 | 0.0868 | 0.0184 | 9.23e-08 | 54.3478
I3 | B | 18 | 2 | H | 0.0598 | 0.0187 | -2.94e-08 | 53.4759 | 0.0874 | 0.0183 | 9.24e-08 | 54.6448
I3 | B | 19 | 2 | H | 0.0596 | 0.0187 | -4e-08 | 53.4759 | 0.0927 | 0.0184 | 8.45e-08 | 54.3478
I3 | B | 20 | 2 | H | 0.0536 | 0.0188 | -4.38e-08 | 53.1915 | 0.0845 | 0.0185 | 8.35e-08 | 54.0541
I3 | B | 21 | 2 | H | 0.0444 | 0.0188 | -5.92e-08 | 53.1915 | 0.0771 | 0.0185 | 6.89e-08 | 54.0541
I3 | B | 22 | 2 | H | 0.0515 | 0.019 | -5.66e-08 | 52.6316 | 0.0831 | 0.0186 | 7.37e-08 | 53.7634
I3 | B | 23 | 2 | H | 0.0579 | 0.0189 | -3.81e-08 | 52.9101 | 0.0857 | 0.0185 | 8.91e-08 | 54.0541
I3 | B | 24 | 2 | H | 0.0514 | 0.0189 | -4.77e-08 | 52.9101 | 0.0783 | 0.0186 | 7.99e-08 | 53.7634
I3 | B | 25 | 2 | H | 0.0592 | 0.0189 | -4.52e-08 | 52.9101 | 0.0842 | 0.0186 | 7.94e-08 | 53.7634
I3 | B | 26 | 2 | H | 0.0579 | 0.0189 | -4.65e-08 | 52.9101 | 0.0849 | 0.0186 | 8.1e-08 | 53.7634
I3 | B | 27 | 2 | H | 0.0565 | 0.0189 | -4.28e-08 | 52.9101 | 0.081 | 0.0186 | 8.14e-08 | 53.7634
I3 | B | 28 | 2 | H | 0.0525 | 0.019 | -5.07e-08 | 52.6316 | 0.0842 | 0.0187 | 8.04e-08 | 53.4759
I3 | B | 29 | 2 | H | 0.0527 | 0.019 | -4.65e-08 | 52.6316 | 0.0792 | 0.0187 | 8.33e-08 | 53.4759
I3 | B | 30 | 2 | H | 0.0585 | 0.0191 | -4.16e-08 | 52.356 | 0.0919 | 0.0188 | 9.01e-08 | 53.1915
I3 | B | 31 | 2 | H | 0.0634 | 0.0191 | -3.85e-08 | 52.356 | 0.1018 | 0.0188 | 9.21e-08 | 53.1915
I3 | B | 32 | 2 | H | 0.0705 | 0.0192 | -4.02e-08 | 52.0833 | 0.0964 | 0.0189 | 9.26e-08 | 52.9101
M1 | A | 1 | 1 | H | -0.9384 | 0.0558 | -1.784e-07 | 17.9211 | -1.0815 | 0.0589 | -5.386e-07 | 16.9779
M1 | A | 2 | 1 | H | -0.8773 | 0.0558 | -1.662e-07 | 17.9211 | -0.9912 | 0.059 | -5.272e-07 | 16.9492
M1 | A | 3 | 1 | H | -0.9004 | 0.0554 | -1.734e-07 | 18.0505 | -1.0183 | 0.0586 | -5.341e-07 | 17.0648
M1 | A | 4 | 1 | H | -0.9185 | 0.0549 | -1.817e-07 | 18.2149 | -1.0362 | 0.058 | -5.429e-07 | 17.2414
M1 | A | 5 | 1 | H | -0.9481 | 0.0546 | -2.028e-07 | 18.315 | -1.0328 | 0.0576 | -5.35e-07 | 17.3611
M1 | A | 6 | 1 | H | -0.8959 | 0.0544 | -1.755e-07 | 18.3824 | -1.0341 | 0.0575 | -5.349e-07 | 17.3913
M1 | A | 7 | 1 | H | -0.9167 | 0.0546 | -1.955e-07 | 18.315 | -1.0386 | 0.0577 | -5.524e-07 | 17.331
M1 | A | 8 | 1 | H | -0.9482 | 0.0544 | -1.896e-07 | 18.3824 | -1.0811 | 0.0575 | -5.536e-07 | 17.3913
M1 | A | 9 | 1 | H | -0.9575 | 0.0543 | -1.993e-07 | 18.4162 | -1.0824 | 0.0573 | -5.545e-07 | 17.452
M1 | A | 10 | 1 | H | -1.0247 | 0.0544 | -2.237e-07 | 18.3824 | -1.1152 | 0.0574 | -5.685e-07 | 17.4216
M1 | A | 11 | 1 | H | -0.9726 | 0.0544 | -2.06e-07 | 18.3824 | -1.0876 | 0.0574 | -5.543e-07 | 17.4216
M1 | A | 12 | 1 | H | -0.9895 | 0.0547 | -2.318e-07 | 18.2815 | -1.1262 | 0.0578 | -5.875e-07 | 17.301
M1 | A | 13 | 1 | H | -0.9746 | 0.0547 | -2.409e-07 | 18.2815 | -1.1379 | 0.0579 | -6.211e-07 | 17.2712
M1 | A | 14 | 1 | H | -1.0403 | 0.0548 | -2.388e-07 | 18.2482 | -1.1745 | 0.0579 | -5.958e-07 | 17.2712
M1 | A | 15 | 1 | H | -1.014 | 0.0554 | -2.466e-07 | 18.0505 | -1.1506 | 0.0586 | -6.246e-07 | 17.0648
M1 | A | 16 | 1 | H | -1.0218 | 0.0558 | -2.442e-07 | 17.9211 | -1.1772 | 0.0589 | -6.432e-07 | 16.9779
M1 | B | 1 | 1 | H | -0.9175 | 0.057 | -1.739e-07 | 17.5439 | -1.0664 | 0.0603 | -5.472e-07 | 16.5837
M1 | B | 2 | 1 | H | -0.8623 | 0.0571 | -1.679e-07 | 17.5131 | -0.9794 | 0.0603 | -5.421e-07 | 16.5837
M1 | B | 3 | 1 | H | -0.9116 | 0.0567 | -1.863e-07 | 17.6367 | -1 | 0.0599 | -5.428e-07 | 16.6945
M1 | B | 4 | 1 | H | -0.9154 | 0.0562 | -1.914e-07 | 17.7936 | -1.0154 | 0.0593 | -5.454e-07 | 16.8634
M1 | B | 5 | 1 | H | -0.902 | 0.0557 | -1.859e-07 | 17.9533 | -1.0258 | 0.0589 | -5.536e-07 | 16.9779
M1 | B | 6 | 1 | H | -0.9181 | 0.0557 | -1.848e-07 | 17.9533 | -1.0388 | 0.0588 | -5.534e-07 | 17.0068
M1 | B | 7 | 1 | H | -0.9287 | 0.0558 | -1.996e-07 | 17.9211 | -1.0518 | 0.059 | -5.734e-07 | 16.9492
M1 | B | 8 | 1 | H | -0.9605 | 0.0556 | -2.029e-07 | 17.9856 | -1.0977 | 0.0588 | -5.84e-07 | 17.0068
M1 | B | 9 | 1 | H | -0.969 | 0.0554 | -2.072e-07 | 18.0505 | -1.1046 | 0.0586 | -5.903e-07 | 17.0648
M1 | B | 10 | 1 | H | -0.9969 | 0.0556 | -2.222e-07 | 17.9856 | -1.0973 | 0.0587 | -5.744e-07 | 17.0358
M1 | B | 11 | 1 | H | -0.9979 | 0.0556 | -2.291e-07 | 17.9856 | -1.0778 | 0.0587 | -5.707e-07 | 17.0358
M1 | B | 12 | 1 | H | -1.0203 | 0.0559 | -2.449e-07 | 17.8891 | -1.1191 | 0.0591 | -6.013e-07 | 16.9205
M1 | B | 13 | 1 | H | -0.9817 | 0.0559 | -2.408e-07 | 17.8891 | -1.1146 | 0.0591 | -6.224e-07 | 16.9205
M1 | B | 14 | 1 | H | -1.0546 | 0.056 | -2.443e-07 | 17.8571 | -1.1572 | 0.0592 | -6.04e-07 | 16.8919
M1 | B | 15 | 1 | H | -1.0157 | 0.0566 | -2.518e-07 | 17.6678 | -1.1337 | 0.0598 | -6.398e-07 | 16.7224
M1 | B | 16 | 1 | H | -1.051 | 0.057 | -2.717e-07 | 17.5439 | -1.1859 | 0.0602 | -6.707e-07 | 16.6113
M2 | A | 1 | 1 | H | 2.6326 | 0.044 | 8.052e-07 | 22.7273 | 3.1354 | 0.045 | 8.823e-07 | 22.2222
M2 | A | 2 | 1 | H | 2.6268 | 0.0442 | 8.053e-07 | 22.6244 | 3.1015 | 0.0452 | 8.604e-07 | 22.1239
M2 | A | 3 | 1 | H | 2.5993 | 0.044 | 8.065e-07 | 22.7273 | 3.1093 | 0.045 | 8.762e-07 | 22.2222
M2 | A | 4 | 1 | H | 2.6049 | 0.0439 | 7.955e-07 | 22.779 | 3.1065 | 0.0448 | 8.68e-07 | 22.3214
M2 | A | 5 | 1 | H | 2.6119 | 0.0438 | 7.982e-07 | 22.8311 | 3.1028 | 0.0448 | 8.549e-07 | 22.3214
M2 | A | 6 | 1 | H | 2.6083 | 0.0436 | 7.889e-07 | 22.9358 | 3.1256 | 0.0446 | 8.647e-07 | 22.4215
M2 | A | 7 | 1 | H | 2.6612 | 0.0434 | 7.972e-07 | 23.0415 | 3.1478 | 0.0444 | 8.653e-07 | 22.5225
M2 | A | 8 | 1 | H | 2.6573 | 0.0432 | 7.972e-07 | 23.1481 | 3.1424 | 0.0442 | 8.622e-07 | 22.6244
M2 | A | 9 | 1 | H | 2.6515 | 0.0432 | 7.931e-07 | 23.1481 | 3.1383 | 0.0442 | 8.593e-07 | 22.6244
M2 | A | 10 | 1 | H | 2.6727 | 0.0432 | 7.946e-07 | 23.1481 | 3.1661 | 0.0442 | 8.608e-07 | 22.6244
M2 | A | 11 | 1 | H | 2.6577 | 0.0435 | 7.95e-07 | 22.9885 | 3.1719 | 0.0445 | 8.751e-07 | 22.4719
M2 | A | 12 | 1 | H | 2.6256 | 0.0436 | 7.917e-07 | 22.9358 | 3.149 | 0.0445 | 8.768e-07 | 22.4719
M2 | A | 13 | 1 | H | 2.6512 | 0.0434 | 7.93e-07 | 23.0415 | 3.1532 | 0.0444 | 8.702e-07 | 22.5225
M2 | A | 14 | 1 | H | 2.6262 | 0.0434 | 7.829e-07 | 23.0415 | 3.1309 | 0.0444 | 8.607e-07 | 22.5225
M2 | A | 15 | 1 | H | 2.6358 | 0.0434 | 8.007e-07 | 23.0415 | 3.1298 | 0.0444 | 8.662e-07 | 22.5225
M2 | A | 16 | 1 | H | 2.6251 | 0.0433 | 7.831e-07 | 23.0947 | 3.144 | 0.0443 | 8.656e-07 | 22.5734
M2 | B | 1 | 1 | H | 2.6488 | 0.0444 | 8.313e-07 | 22.5225 | 3.133 | 0.0454 | 9.009e-07 | 22.0264
M2 | B | 2 | 1 | H | 2.6256 | 0.0446 | 8.209e-07 | 22.4215 | 3.1193 | 0.0456 | 8.888e-07 | 21.9298
M2 | B | 3 | 1 | H | 2.5968 | 0.0444 | 8.182e-07 | 22.5225 | 3.115 | 0.0454 | 8.959e-07 | 22.0264
M2 | B | 4 | 1 | H | 2.5988 | 0.0443 | 8.099e-07 | 22.5734 | 3.0989 | 0.0453 | 8.814e-07 | 22.0751
M2 | B | 5 | 1 | H | 2.6219 | 0.0442 | 8.192e-07 | 22.6244 | 3.1046 | 0.0452 | 8.709e-07 | 22.1239
M2 | B | 6 | 1 | H | 2.6146 | 0.044 | 8.063e-07 | 22.7273 | 3.1062 | 0.045 | 8.69e-07 | 22.2222
M2 | B | 7 | 1 | H | 2.6639 | 0.0438 | 8.156e-07 | 22.8311 | 3.1163 | 0.0449 | 8.617e-07 | 22.2717
M2 | B | 8 | 1 | H | 2.6577 | 0.0436 | 8.145e-07 | 22.9358 | 3.1343 | 0.0446 | 8.782e-07 | 22.4215
M2 | B | 9 | 1 | H | 2.6608 | 0.0435 | 8.164e-07 | 22.9885 | 3.15 | 0.0445 | 8.854e-07 | 22.4719
M2 | B | 10 | 1 | H | 2.6724 | 0.0436 | 8.113e-07 | 22.9358 | 3.1529 | 0.0446 | 8.699e-07 | 22.4215
M2 | B | 11 | 1 | H | 2.6652 | 0.0439 | 8.151e-07 | 22.779 | 3.132 | 0.0449 | 8.699e-07 | 22.2717
M2 | B | 12 | 1 | H | 2.6317 | 0.0439 | 8.16e-07 | 22.779 | 3.1293 | 0.0449 | 8.816e-07 | 22.2717
M2 | B | 13 | 1 | H | 2.6486 | 0.0438 | 8.122e-07 | 22.8311 | 3.1275 | 0.0448 | 8.705e-07 | 22.3214
M2 | B | 14 | 1 | H | 2.6295 | 0.0438 | 8.038e-07 | 22.8311 | 3.1351 | 0.0448 | 8.83e-07 | 22.3214
M2 | B | 15 | 1 | H | 2.6431 | 0.0437 | 8.155e-07 | 22.8833 | 3.1413 | 0.0447 | 8.914e-07 | 22.3714
M2 | B | 16 | 1 | H | 2.6368 | 0.0437 | 8.072e-07 | 22.8833 | 3.1483 | 0.0447 | 8.865e-07 | 22.3714
M3 | A | 1 | 1 | H | 0.1449 | 0.0384 | -2.417e-07 | 26.0417 | 0.8281 | 0.0394 | -4.039e-07 | 25.3807
M3 | A | 2 | 1 | H | 0.1224 | 0.0386 | -2.726e-07 | 25.9067 | 0.8263 | 0.0396 | -4.284e-07 | 25.2525
M3 | A | 3 | 1 | H | 0.1168 | 0.0385 | -2.707e-07 | 25.974 | 0.7798 | 0.0395 | -4.419e-07 | 25.3165
M3 | A | 4 | 1 | H | 0.0839 | 0.0384 | -2.711e-07 | 26.0417 | 0.7668 | 0.0394 | -4.323e-07 | 25.3807
M3 | A | 5 | 1 | H | 0.0797 | 0.0383 | -2.754e-07 | 26.1097 | 0.7484 | 0.0394 | -4.429e-07 | 25.3807
M3 | A | 6 | 1 | H | 0.0625 | 0.0384 | -2.891e-07 | 26.0417 | 0.7428 | 0.0394 | -4.53e-07 | 25.3807
M3 | A | 7 | 1 | H | 0.054 | 0.0385 | -2.995e-07 | 25.974 | 0.7419 | 0.0395 | -4.603e-07 | 25.3165
M3 | A | 8 | 1 | H | 0.0444 | 0.0385 | -3.116e-07 | 25.974 | 0.7236 | 0.0396 | -4.789e-07 | 25.2525
M3 | A | 9 | 1 | H | 0.0559 | 0.0386 | -3.117e-07 | 25.9067 | 0.7099 | 0.0397 | -4.861e-07 | 25.1889
M3 | A | 10 | 1 | H | 0.014 | 0.0387 | -3.279e-07 | 25.8398 | 0.6741 | 0.0397 | -4.985e-07 | 25.1889
M3 | A | 11 | 1 | H | -0.0049 | 0.0387 | -3.346e-07 | 25.8398 | 0.6703 | 0.0397 | -5.066e-07 | 25.1889
M3 | A | 12 | 1 | H | 0.0161 | 0.0388 | -3.383e-07 | 25.7732 | 0.6814 | 0.0399 | -5.115e-07 | 25.0627
M3 | A | 13 | 1 | H | 0.0142 | 0.0391 | -3.467e-07 | 25.5754 | 0.6754 | 0.0401 | -5.221e-07 | 24.9377
M3 | A | 14 | 1 | H | -0.0416 | 0.0392 | -3.575e-07 | 25.5102 | 0.6425 | 0.0402 | -5.284e-07 | 24.8756
M3 | A | 15 | 1 | H | -0.049 | 0.0393 | -3.608e-07 | 25.4453 | 0.6196 | 0.0404 | -5.38e-07 | 24.7525
M3 | A | 16 | 1 | H | -0.0505 | 0.0396 | -3.737e-07 | 25.2525 | 0.6257 | 0.0406 | -5.496e-07 | 24.6305
M3 | B | 1 | 1 | H | 0.1429 | 0.0385 | -2.459e-07 | 25.974 | 0.8383 | 0.0396 | -4.019e-07 | 25.2525
M3 | B | 2 | 1 | H | 0.1284 | 0.0387 | -2.705e-07 | 25.8398 | 0.8369 | 0.0397 | -4.276e-07 | 25.1889
M3 | B | 3 | 1 | H | 0.1054 | 0.0386 | -2.778e-07 | 25.9067 | 0.7757 | 0.0397 | -4.483e-07 | 25.1889
M3 | B | 4 | 1 | H | 0.0966 | 0.0385 | -2.683e-07 | 25.974 | 0.7648 | 0.0395 | -4.381e-07 | 25.3165
M3 | B | 5 | 1 | H | 0.0833 | 0.0385 | -2.777e-07 | 25.974 | 0.7456 | 0.0395 | -4.494e-07 | 25.3165
M3 | B | 6 | 1 | H | 0.0676 | 0.0385 | -2.896e-07 | 25.974 | 0.7406 | 0.0396 | -4.594e-07 | 25.2525
M3 | B | 7 | 1 | H | 0.0645 | 0.0387 | -2.981e-07 | 25.8398 | 0.7367 | 0.0397 | -4.668e-07 | 25.1889
M3 | B | 8 | 1 | H | 0.056 | 0.0387 | -3.092e-07 | 25.8398 | 0.7365 | 0.0397 | -4.782e-07 | 25.1889
M3 | B | 9 | 1 | H | 0.0332 | 0.0388 | -3.23e-07 | 25.7732 | 0.6923 | 0.0398 | -4.965e-07 | 25.1256
M3 | B | 10 | 1 | H | 0.0155 | 0.0388 | -3.293e-07 | 25.7732 | 0.6853 | 0.0399 | -4.971e-07 | 25.0627
M3 | B | 11 | 1 | H | -0.0016 | 0.0388 | -3.368e-07 | 25.7732 | 0.6934 | 0.0398 | -4.989e-07 | 25.1256
M3 | B | 12 | 1 | H | 0.0052 | 0.039 | -3.459e-07 | 25.641 | 0.6795 | 0.04 | -5.183e-07 | 25
M3 | B | 13 | 1 | H | 0.0046 | 0.0392 | -3.53e-07 | 25.5102 | 0.6747 | 0.0403 | -5.298e-07 | 24.8139
M3 | B | 14 | 1 | H | -0.0294 | 0.0393 | -3.542e-07 | 25.4453 | 0.6508 | 0.0404 | -5.281e-07 | 24.7525
M3 | B | 15 | 1 | H | -0.0528 | 0.0395 | -3.647e-07 | 25.3165 | 0.6111 | 0.0405 | -5.44e-07 | 24.6914
M3 | B | 16 | 1 | H | -0.0538 | 0.0397 | -3.795e-07 | 25.1889 | 0.6297 | 0.0408 | -5.551e-07 | 24.5098
M4 | A | 1 | 1 | H | 0.3266 | 0.0281 | 2.243e-07 | 35.5872 | 1.1934 | 0.0273 | 9.315e-07 | 36.63
M4 | A | 2 | 1 | H | 0.3186 | 0.0282 | 2.243e-07 | 35.461 | 1.186 | 0.0274 | 9.314e-07 | 36.4964
M4 | A | 3 | 1 | H | 0.3263 | 0.0281 | 2.281e-07 | 35.5872 | 1.1849 | 0.0274 | 9.277e-07 | 36.4964
M4 | A | 4 | 1 | H | 0.3086 | 0.0281 | 2.38e-07 | 35.5872 | 1.1735 | 0.0274 | 9.432e-07 | 36.4964
M4 | A | 5 | 1 | H | 0.3078 | 0.0281 | 2.369e-07 | 35.5872 | 1.1688 | 0.0274 | 9.417e-07 | 36.4964
M4 | A | 6 | 1 | H | 0.3098 | 0.0281 | 2.384e-07 | 35.5872 | 1.1802 | 0.0274 | 9.5e-07 | 36.4964
M4 | A | 7 | 1 | H | 0.3178 | 0.0281 | 2.378e-07 | 35.5872 | 1.1848 | 0.0274 | 9.515e-07 | 36.4964
M4 | A | 8 | 1 | H | 0.3347 | 0.0283 | 2.49e-07 | 35.3357 | 1.2019 | 0.0275 | 9.687e-07 | 36.3636
M4 | A | 9 | 1 | H | 0.323 | 0.0282 | 2.527e-07 | 35.461 | 1.1913 | 0.0275 | 9.704e-07 | 36.3636
M4 | A | 10 | 1 | H | 0.3189 | 0.0284 | 2.506e-07 | 35.2113 | 1.1905 | 0.0277 | 9.811e-07 | 36.1011
M4 | A | 11 | 1 | H | 0.325 | 0.0283 | 2.518e-07 | 35.3357 | 1.1934 | 0.0276 | 9.803e-07 | 36.2319
M4 | A | 12 | 1 | H | 0.3292 | 0.0285 | 2.608e-07 | 35.0877 | 1.2009 | 0.0277 | 9.958e-07 | 36.1011
M4 | A | 13 | 1 | H | 0.3245 | 0.0286 | 2.584e-07 | 34.965 | 1.2121 | 0.0279 | 1.0096e-06 | 35.8423
M4 | A | 14 | 1 | H | 0.3148 | 0.0288 | 2.694e-07 | 34.7222 | 1.1999 | 0.028 | 1.0254e-06 | 35.7143
M4 | A | 15 | 1 | H | 0.3379 | 0.0288 | 2.796e-07 | 34.7222 | 1.1818 | 0.028 | 1.0205e-06 | 35.7143
M4 | A | 16 | 1 | H | 0.3318 | 0.0287 | 2.787e-07 | 34.8432 | 1.1945 | 0.0279 | 1.0244e-06 | 35.8423
M4 | B | 1 | 1 | H | 0.3325 | 0.028 | 2.256e-07 | 35.7143 | 1.1892 | 0.0273 | 9.269e-07 | 36.63
M4 | B | 2 | 1 | H | 0.3207 | 0.0281 | 2.235e-07 | 35.5872 | 1.1835 | 0.0274 | 9.283e-07 | 36.4964
M4 | B | 3 | 1 | H | 0.3221 | 0.0281 | 2.26e-07 | 35.5872 | 1.1795 | 0.0274 | 9.23e-07 | 36.4964
M4 | B | 4 | 1 | H | 0.3067 | 0.0281 | 2.348e-07 | 35.5872 | 1.1674 | 0.0274 | 9.383e-07 | 36.4964
M4 | B | 5 | 1 | H | 0.3069 | 0.028 | 2.376e-07 | 35.7143 | 1.1636 | 0.0273 | 9.37e-07 | 36.63
M4 | B | 6 | 1 | H | 0.3122 | 0.0281 | 2.393e-07 | 35.5872 | 1.1754 | 0.0274 | 9.458e-07 | 36.4964
M4 | B | 7 | 1 | H | 0.3195 | 0.0281 | 2.386e-07 | 35.5872 | 1.1849 | 0.0274 | 9.479e-07 | 36.4964
M4 | B | 8 | 1 | H | 0.3291 | 0.0282 | 2.446e-07 | 35.461 | 1.1966 | 0.0275 | 9.632e-07 | 36.3636
M4 | B | 9 | 1 | H | 0.3173 | 0.0282 | 2.467e-07 | 35.461 | 1.191 | 0.0275 | 9.68e-07 | 36.3636
M4 | B | 10 | 1 | H | 0.3222 | 0.0284 | 2.524e-07 | 35.2113 | 1.1877 | 0.0277 | 9.775e-07 | 36.1011
M4 | B | 11 | 1 | H | 0.3248 | 0.0283 | 2.517e-07 | 35.3357 | 1.1923 | 0.0276 | 9.765e-07 | 36.2319
M4 | B | 12 | 1 | H | 0.335 | 0.0284 | 2.632e-07 | 35.2113 | 1.1927 | 0.0277 | 9.901e-07 | 36.1011
M4 | B | 13 | 1 | H | 0.3252 | 0.0286 | 2.589e-07 | 34.965 | 1.2052 | 0.0279 | 1.0039e-06 | 35.8423
M4 | B | 14 | 1 | H | 0.3167 | 0.0287 | 2.699e-07 | 34.8432 | 1.1984 | 0.028 | 1.0225e-06 | 35.7143
M4 | B | 15 | 1 | H | 0.3347 | 0.0287 | 2.758e-07 | 34.8432 | 1.1879 | 0.028 | 1.0199e-06 | 35.7143
M4 | B | 16 | 1 | H | 0.3309 | 0.0286 | 2.793e-07 | 34.965 | 1.1949 | 0.0279 | 1.0235e-06 | 35.8423
M5 | A | 1 | 1 | H | -0.0014 | 0.0208 | -1.347e-07 | 48.0769 | 0.0882 | 0.0215 | -1.785e-07 | 46.5116
M5 | A | 2 | 1 | H | -0.0045 | 0.021 | -1.534e-07 | 47.619 | 0.089 | 0.0216 | -1.977e-07 | 46.2963
M5 | A | 3 | 1 | H | 0.0027 | 0.0209 | -1.524e-07 | 47.8469 | 0.0794 | 0.0216 | -2.007e-07 | 46.2963
M5 | A | 4 | 1 | H | -0.0181 | 0.0209 | -1.619e-07 | 47.8469 | 0.0735 | 0.0215 | -2.058e-07 | 46.5116
M5 | A | 5 | 1 | H | -0.0111 | 0.0208 | -1.596e-07 | 48.0769 | 0.068 | 0.0214 | -2.046e-07 | 46.729
M5 | A | 6 | 1 | H | -0.0196 | 0.0209 | -1.666e-07 | 47.8469 | 0.069 | 0.0215 | -2.127e-07 | 46.5116
M5 | A | 7 | 1 | H | -0.0205 | 0.0209 | -1.734e-07 | 47.8469 | 0.0733 | 0.0216 | -2.147e-07 | 46.2963
M5 | A | 8 | 1 | H | -0.0188 | 0.0208 | -1.725e-07 | 48.0769 | 0.0728 | 0.0214 | -2.175e-07 | 46.729
M5 | A | 9 | 1 | H | -0.0209 | 0.0209 | -1.776e-07 | 47.8469 | 0.0664 | 0.0215 | -2.231e-07 | 46.5116
M5 | A | 10 | 1 | H | -0.0262 | 0.0209 | -1.811e-07 | 47.8469 | 0.0601 | 0.0216 | -2.264e-07 | 46.2963
M5 | A | 11 | 1 | H | -0.024 | 0.0211 | -1.816e-07 | 47.3934 | 0.0563 | 0.0217 | -2.295e-07 | 46.0829
M5 | A | 12 | 1 | H | -0.0334 | 0.021 | -1.82e-07 | 47.619 | 0.0509 | 0.0217 | -2.333e-07 | 46.0829
M5 | A | 13 | 1 | H | -0.0324 | 0.0211 | -1.907e-07 | 47.3934 | 0.0526 | 0.0217 | -2.42e-07 | 46.0829
M5 | A | 14 | 1 | H | -0.0423 | 0.0211 | -1.961e-07 | 47.3934 | 0.0468 | 0.0218 | -2.408e-07 | 45.8716
M5 | A | 15 | 1 | H | -0.0448 | 0.0213 | -2.009e-07 | 46.9484 | 0.0344 | 0.0219 | -2.539e-07 | 45.6621
M5 | A | 16 | 1 | H | -0.0483 | 0.0212 | -2.024e-07 | 47.1698 | 0.0307 | 0.0219 | -2.581e-07 | 45.6621
M5 | B | 1 | 1 | H | -0.0016 | 0.0207 | -1.342e-07 | 48.3092 | 0.0844 | 0.0214 | -1.788e-07 | 46.729
M5 | B | 2 | 1 | H | -0.0126 | 0.0209 | -1.566e-07 | 47.8469 | 0.0837 | 0.0215 | -1.975e-07 | 46.5116
M5 | B | 3 | 1 | H | -0.0077 | 0.0208 | -1.554e-07 | 48.0769 | 0.0781 | 0.0215 | -2.028e-07 | 46.5116
M5 | B | 4 | 1 | H | -0.0154 | 0.0208 | -1.579e-07 | 48.0769 | 0.0691 | 0.0214 | -2.076e-07 | 46.729
M5 | B | 5 | 1 | H | -0.0137 | 0.0207 | -1.58e-07 | 48.3092 | 0.0655 | 0.0213 | -2.053e-07 | 46.9484
M5 | B | 6 | 1 | H | -0.0224 | 0.0208 | -1.675e-07 | 48.0769 | 0.0695 | 0.0214 | -2.106e-07 | 46.729
M5 | B | 7 | 1 | H | -0.0201 | 0.0208 | -1.719e-07 | 48.0769 | 0.0636 | 0.0215 | -2.19e-07 | 46.5116
M5 | B | 8 | 1 | H | -0.0146 | 0.0207 | -1.686e-07 | 48.3092 | 0.0728 | 0.0214 | -2.156e-07 | 46.729
M5 | B | 9 | 1 | H | -0.0258 | 0.0208 | -1.785e-07 | 48.0769 | 0.0571 | 0.0215 | -2.27e-07 | 46.5116
M5 | B | 10 | 1 | H | -0.0264 | 0.0208 | -1.797e-07 | 48.0769 | 0.0573 | 0.0215 | -2.281e-07 | 46.5116
M5 | B | 11 | 1 | H | -0.0253 | 0.021 | -1.814e-07 | 47.619 | 0.0506 | 0.0216 | -2.325e-07 | 46.2963
M5 | B | 12 | 1 | H | -0.0399 | 0.021 | -1.865e-07 | 47.619 | 0.0472 | 0.0216 | -2.328e-07 | 46.2963
M5 | B | 13 | 1 | H | -0.037 | 0.021 | -1.914e-07 | 47.619 | 0.0572 | 0.0216 | -2.364e-07 | 46.2963
M5 | B | 14 | 1 | H | -0.0396 | 0.021 | -1.932e-07 | 47.619 | 0.0451 | 0.0217 | -2.414e-07 | 46.0829
M5 | B | 15 | 1 | H | -0.0413 | 0.0212 | -1.969e-07 | 47.1698 | 0.029 | 0.0219 | -2.539e-07 | 45.6621
M5 | B | 16 | 1 | H | -0.0506 | 0.0211 | -2.015e-07 | 47.3934 | 0.0387 | 0.0218 | -2.484e-07 | 45.8716
M6 | A | 1 | 1 | H | 0.0701 | 0.0129 | -1.39e-08 | 77.5194 | 0.0868 | 0.0133 | -3.64e-08 | 75.188
M6 | A | 2 | 1 | H | 0.0569 | 0.013 | -2.89e-08 | 76.9231 | 0.078 | 0.0134 | -5.14e-08 | 74.6269
M6 | A | 3 | 1 | H | 0.0574 | 0.013 | -2.61e-08 | 76.9231 | 0.0796 | 0.0134 | -4.73e-08 | 74.6269
M6 | A | 4 | 1 | H | 0.0518 | 0.0129 | -2.49e-08 | 77.5194 | 0.0811 | 0.0133 | -4.59e-08 | 75.188
M6 | A | 5 | 1 | H | 0.0563 | 0.013 | -2.15e-08 | 76.9231 | 0.0748 | 0.0133 | -4.52e-08 | 75.188
M6 | A | 6 | 1 | H | 0.0576 | 0.0131 | -2.58e-08 | 76.3359 | 0.0744 | 0.0134 | -5.05e-08 | 74.6269
M6 | A | 7 | 1 | H | 0.0522 | 0.013 | -2.86e-08 | 76.9231 | 0.0788 | 0.0134 | -4.78e-08 | 74.6269
M6 | A | 8 | 1 | H | 0.056 | 0.013 | -2.41e-08 | 76.9231 | 0.0828 | 0.0134 | -4.4e-08 | 74.6269
M6 | A | 9 | 1 | H | 0.0535 | 0.0131 | -2.25e-08 | 76.3359 | 0.0756 | 0.0135 | -4.34e-08 | 74.0741
M6 | A | 10 | 1 | H | 0.0545 | 0.0132 | -2.75e-08 | 75.7576 | 0.0788 | 0.0135 | -4.84e-08 | 74.0741
M6 | A | 11 | 1 | H | 0.046 | 0.0132 | -2.77e-08 | 75.7576 | 0.0769 | 0.0135 | -4.62e-08 | 74.0741
M6 | A | 12 | 1 | H | 0.0447 | 0.0132 | -3.08e-08 | 75.7576 | 0.0762 | 0.0135 | -4.81e-08 | 74.0741
M6 | A | 13 | 1 | H | 0.0548 | 0.0133 | -2.98e-08 | 75.188 | 0.0767 | 0.0137 | -5.01e-08 | 72.9927
M6 | A | 14 | 1 | H | 0.05 | 0.0133 | -2.72e-08 | 75.188 | 0.0745 | 0.0136 | -4.66e-08 | 73.5294
M6 | A | 15 | 1 | H | 0.0436 | 0.0133 | -2.86e-08 | 75.188 | 0.0777 | 0.0136 | -4.4e-08 | 73.5294
M6 | A | 16 | 1 | H | 0.0458 | 0.0133 | -2.71e-08 | 75.188 | 0.0709 | 0.0136 | -4.65e-08 | 73.5294
M6 | B | 1 | 1 | H | 0.0691 | 0.0129 | -1.46e-08 | 77.5194 | 0.0901 | 0.0132 | -3.39e-08 | 75.7576
M6 | B | 2 | 1 | H | 0.0556 | 0.013 | -2.99e-08 | 76.9231 | 0.0744 | 0.0133 | -5.2e-08 | 75.188
M6 | B | 3 | 1 | H | 0.0572 | 0.013 | -2.52e-08 | 76.9231 | 0.079 | 0.0133 | -4.84e-08 | 75.188
M6 | B | 4 | 1 | H | 0.0489 | 0.0129 | -2.69e-08 | 77.5194 | 0.0751 | 0.0132 | -4.8e-08 | 75.7576
M6 | B | 5 | 1 | H | 0.0514 | 0.0129 | -2.33e-08 | 77.5194 | 0.0761 | 0.0133 | -4.35e-08 | 75.188
M6 | B | 6 | 1 | H | 0.058 | 0.013 | -2.62e-08 | 76.9231 | 0.0753 | 0.0134 | -5e-08 | 74.6269
M6 | B | 7 | 1 | H | 0.052 | 0.013 | -2.82e-08 | 76.9231 | 0.0779 | 0.0133 | -4.91e-08 | 75.188
M6 | B | 8 | 1 | H | 0.0571 | 0.013 | -2.37e-08 | 76.9231 | 0.0815 | 0.0133 | -4.45e-08 | 75.188
M6 | B | 9 | 1 | H | 0.0537 | 0.0131 | -2.32e-08 | 76.3359 | 0.0747 | 0.0134 | -4.5e-08 | 74.6269
M6 | B | 10 | 1 | H | 0.0554 | 0.0131 | -2.66e-08 | 76.3359 | 0.0817 | 0.0135 | -4.69e-08 | 74.0741
M6 | B | 11 | 1 | H | 0.0507 | 0.0131 | -2.46e-08 | 76.3359 | 0.0763 | 0.0135 | -4.67e-08 | 74.0741
M6 | B | 12 | 1 | H | 0.0498 | 0.0131 | -2.81e-08 | 76.3359 | 0.0764 | 0.0135 | -4.78e-08 | 74.0741
M6 | B | 13 | 1 | H | 0.0591 | 0.0133 | -2.82e-08 | 75.188 | 0.079 | 0.0136 | -4.8e-08 | 73.5294
M6 | B | 14 | 1 | H | 0.0503 | 0.0132 | -2.74e-08 | 75.7576 | 0.0746 | 0.0136 | -4.64e-08 | 73.5294
M6 | B | 15 | 1 | H | 0.0458 | 0.0132 | -2.75e-08 | 75.7576 | 0.0761 | 0.0135 | -4.52e-08 | 74.0741
M6 | B | 16 | 1 | H | 0.0399 | 0.0133 | -2.93e-08 | 75.188 | 0.0735 | 0.0136 | -4.53e-08 | 73.5294
M7 | A | 1 | 1 | H | 0.0412 | 0.01 | -2.41e-08 | 100 | 0.085 | 0.0102 | -1.74e-08 | 98.0392
M7 | A | 2 | 1 | H | 0.0351 | 0.0101 | -3.11e-08 | 99.0099 | 0.0817 | 0.0102 | -2.39e-08 | 98.0392
M7 | A | 3 | 1 | H | 0.0324 | 0.01 | -2.94e-08 | 100 | 0.0766 | 0.0102 | -2.29e-08 | 98.0392
M7 | A | 4 | 1 | H | 0.032 | 0.01 | -2.98e-08 | 100 | 0.0794 | 0.0102 | -2.26e-08 | 98.0392
M7 | A | 5 | 1 | H | 0.0301 | 0.0101 | -3.04e-08 | 99.0099 | 0.0831 | 0.0102 | -2.13e-08 | 98.0392
M7 | A | 6 | 1 | H | 0.0338 | 0.01 | -2.95e-08 | 100 | 0.087 | 0.0101 | -2.1e-08 | 99.0099
M7 | A | 7 | 1 | H | 0.0321 | 0.0101 | -2.87e-08 | 99.0099 | 0.0851 | 0.0102 | -2.01e-08 | 98.0392
M7 | A | 8 | 1 | H | 0.036 | 0.0101 | -2.92e-08 | 99.0099 | 0.0912 | 0.0102 | -1.87e-08 | 98.0392
M7 | A | 9 | 1 | H | 0.0422 | 0.0101 | -2.89e-08 | 99.0099 | 0.0986 | 0.0102 | -1.86e-08 | 98.0392
M7 | A | 10 | 1 | H | 0.0384 | 0.0101 | -2.68e-08 | 99.0099 | 0.088 | 0.0102 | -1.88e-08 | 98.0392
M7 | A | 11 | 1 | H | 0.0336 | 0.0101 | -2.81e-08 | 99.0099 | 0.089 | 0.0102 | -1.82e-08 | 98.0392
M7 | A | 12 | 1 | H | 0.0342 | 0.0101 | -2.58e-08 | 99.0099 | 0.088 | 0.0102 | -1.64e-08 | 98.0392
M7 | A | 13 | 1 | H | 0.036 | 0.0101 | -2.78e-08 | 99.0099 | 0.0979 | 0.0102 | -1.51e-08 | 98.0392
M7 | A | 14 | 1 | H | 0.0356 | 0.0101 | -2.75e-08 | 99.0099 | 0.0904 | 0.0102 | -1.58e-08 | 98.0392
M7 | A | 15 | 1 | H | 0.0291 | 0.0102 | -2.71e-08 | 98.0392 | 0.1002 | 0.0103 | -1.2e-08 | 97.0874
M7 | A | 16 | 1 | H | 0.0327 | 0.0101 | -2.59e-08 | 99.0099 | 0.0915 | 0.0102 | -1.48e-08 | 98.0392
M7 | B | 1 | 1 | H | 0.0406 | 0.0101 | -2.43e-08 | 99.0099 | 0.0847 | 0.0102 | -1.73e-08 | 98.0392
M7 | B | 2 | 1 | H | 0.0326 | 0.0101 | -3.17e-08 | 99.0099 | 0.0807 | 0.0102 | -2.48e-08 | 98.0392
M7 | B | 3 | 1 | H | 0.0354 | 0.01 | -2.85e-08 | 100 | 0.0759 | 0.0102 | -2.33e-08 | 98.0392
M7 | B | 4 | 1 | H | 0.0382 | 0.01 | -2.79e-08 | 100 | 0.0794 | 0.0102 | -2.3e-08 | 98.0392
M7 | B | 5 | 1 | H | 0.0331 | 0.0101 | -2.93e-08 | 99.0099 | 0.0827 | 0.0102 | -2.16e-08 | 98.0392
M7 | B | 6 | 1 | H | 0.0326 | 0.01 | -3.02e-08 | 100 | 0.0823 | 0.0101 | -2.27e-08 | 99.0099
M7 | B | 7 | 1 | H | 0.0337 | 0.0101 | -2.86e-08 | 99.0099 | 0.0836 | 0.0102 | -2.1e-08 | 98.0392
M7 | B | 8 | 1 | H | 0.0368 | 0.0101 | -2.91e-08 | 99.0099 | 0.0924 | 0.0102 | -1.92e-08 | 98.0392
M7 | B | 9 | 1 | H | 0.0407 | 0.0101 | -2.97e-08 | 99.0099 | 0.0995 | 0.0102 | -1.81e-08 | 98.0392
M7 | B | 10 | 1 | H | 0.0368 | 0.0101 | -2.78e-08 | 99.0099 | 0.0853 | 0.0102 | -1.94e-08 | 98.0392
M7 | B | 11 | 1 | H | 0.0335 | 0.0101 | -2.84e-08 | 99.0099 | 0.0852 | 0.0102 | -1.99e-08 | 98.0392
M7 | B | 12 | 1 | H | 0.0323 | 0.0101 | -2.69e-08 | 99.0099 | 0.0864 | 0.0102 | -1.69e-08 | 98.0392
M7 | B | 13 | 1 | H | 0.0349 | 0.0102 | -2.81e-08 | 98.0392 | 0.0998 | 0.0102 | -1.47e-08 | 98.0392
M7 | B | 14 | 1 | H | 0.0381 | 0.0101 | -2.67e-08 | 99.0099 | 0.0891 | 0.0102 | -1.63e-08 | 98.0392
M7 | B | 15 | 1 | H | 0.0309 | 0.0102 | -2.66e-08 | 98.0392 | 0.0977 | 0.0103 | -1.31e-08 | 97.0874
M7 | B | 16 | 1 | H | 0.0346 | 0.0101 | -2.54e-08 | 99.0099 | 0.0881 | 0.0102 | -1.58e-08 | 98.0392
M8 | A | 1 | 1 | H | 0.5577 | 0.0284 | 2.411e-07 | 35.2113 | 0.9723 | 0.0276 | 5.83e-07 | 36.2319
M8 | A | 2 | 1 | H | 0.5313 | 0.0284 | 2.374e-07 | 35.2113 | 0.941 | 0.0277 | 5.812e-07 | 36.1011
M8 | A | 3 | 1 | H | 0.5841 | 0.0287 | 2.429e-07 | 34.8432 | 1.0102 | 0.0279 | 5.953e-07 | 35.8423
M8 | A | 4 | 1 | H | 0.6374 | 0.0292 | 2.534e-07 | 34.2466 | 1.0539 | 0.0284 | 6.164e-07 | 35.2113
M8 | A | 5 | 1 | H | 0.6399 | 0.0292 | 2.547e-07 | 34.2466 | 1.0533 | 0.0284 | 6.173e-07 | 35.2113
M8 | A | 6 | 1 | H | 0.6762 | 0.0296 | 2.536e-07 | 33.7838 | 1.1129 | 0.0288 | 6.294e-07 | 34.7222
M8 | A | 7 | 1 | H | 0.641 | 0.0298 | 2.579e-07 | 33.557 | 1.064 | 0.029 | 6.35e-07 | 34.4828
M8 | A | 8 | 1 | H | 0.6001 | 0.0299 | 2.568e-07 | 33.4448 | 1.0252 | 0.0291 | 6.408e-07 | 34.3643
M8 | A | 9 | 1 | H | 0.6438 | 0.0303 | 2.552e-07 | 33.0033 | 1.0709 | 0.0294 | 6.431e-07 | 34.0136
M8 | A | 10 | 1 | H | 0.6616 | 0.0301 | 2.528e-07 | 33.2226 | 1.0858 | 0.0293 | 6.39e-07 | 34.1297
M8 | A | 11 | 1 | H | 0.5997 | 0.0305 | 2.569e-07 | 32.7869 | 1.0211 | 0.0296 | 6.51e-07 | 33.7838
M8 | A | 12 | 1 | H | 0.6103 | 0.0306 | 2.569e-07 | 32.6797 | 1.0275 | 0.0297 | 6.564e-07 | 33.67
M8 | A | 13 | 1 | H | 0.5859 | 0.031 | 2.581e-07 | 32.2581 | 1.0031 | 0.0301 | 6.671e-07 | 33.2226
M8 | A | 14 | 1 | H | 0.6223 | 0.0309 | 2.53e-07 | 32.3625 | 1.0436 | 0.03 | 6.59e-07 | 33.3333
M8 | A | 15 | 1 | H | 0.6959 | 0.0313 | 2.654e-07 | 31.9489 | 1.1252 | 0.0304 | 6.8e-07 | 32.8947
M8 | A | 16 | 1 | H | 0.7447 | 0.0314 | 2.546e-07 | 31.8471 | 1.1786 | 0.0305 | 6.716e-07 | 32.7869
M8 | B | 1 | 1 | H | 0.555 | 0.0285 | 2.426e-07 | 35.0877 | 0.9706 | 0.0277 | 5.857e-07 | 36.1011
M8 | B | 2 | 1 | H | 0.5288 | 0.0285 | 2.38e-07 | 35.0877 | 0.9429 | 0.0278 | 5.869e-07 | 35.9712
M8 | B | 3 | 1 | H | 0.5833 | 0.0288 | 2.457e-07 | 34.7222 | 1.0073 | 0.028 | 5.981e-07 | 35.7143
M8 | B | 4 | 1 | H | 0.6381 | 0.0293 | 2.548e-07 | 34.1297 | 1.0521 | 0.0285 | 6.201e-07 | 35.0877
M8 | B | 5 | 1 | H | 0.6364 | 0.0293 | 2.549e-07 | 34.1297 | 1.0532 | 0.0285 | 6.215e-07 | 35.0877
M8 | B | 6 | 1 | H | 0.6838 | 0.0297 | 2.573e-07 | 33.67 | 1.1136 | 0.0289 | 6.344e-07 | 34.6021
M8 | B | 7 | 1 | H | 0.6404 | 0.0299 | 2.594e-07 | 33.4448 | 1.0683 | 0.0291 | 6.402e-07 | 34.3643
M8 | B | 8 | 1 | H | 0.6002 | 0.03 | 2.586e-07 | 33.3333 | 1.0244 | 0.0292 | 6.449e-07 | 34.2466
M8 | B | 9 | 1 | H | 0.6439 | 0.0304 | 2.577e-07 | 32.8947 | 1.0727 | 0.0295 | 6.497e-07 | 33.8983
M8 | B | 10 | 1 | H | 0.6613 | 0.0302 | 2.545e-07 | 33.1126 | 1.0861 | 0.0294 | 6.425e-07 | 34.0136
M8 | B | 11 | 1 | H | 0.602 | 0.0306 | 2.599e-07 | 32.6797 | 1.0242 | 0.0297 | 6.579e-07 | 33.67
M8 | B | 12 | 1 | H | 0.6129 | 0.0307 | 2.604e-07 | 32.5733 | 1.0245 | 0.0298 | 6.596e-07 | 33.557
M8 | B | 13 | 1 | H | 0.5818 | 0.0311 | 2.59e-07 | 32.1543 | 1.0047 | 0.0302 | 6.72e-07 | 33.1126
M8 | B | 14 | 1 | H | 0.6211 | 0.031 | 2.536e-07 | 32.2581 | 1.0456 | 0.0301 | 6.646e-07 | 33.2226
M8 | B | 15 | 1 | H | 0.6982 | 0.0314 | 2.68e-07 | 31.8471 | 1.1274 | 0.0305 | 6.854e-07 | 32.7869
M8 | B | 16 | 1 | H | 0.7465 | 0.0315 | 2.561e-07 | 31.746 | 1.1786 | 0.0306 | 6.76e-07 | 32.6797
M9 | A | 1 | 1 | H | 0.3588 | 0.0232 | 1.79e-07 | 43.1034 | 0.5055 | 0.0227 | 2.701e-07 | 44.0529
M9 | A | 2 | 1 | H | 0.3696 | 0.0232 | 1.835e-07 | 43.1034 | 0.5126 | 0.0228 | 2.754e-07 | 43.8596
M9 | A | 3 | 1 | H | 0.406 | 0.0234 | 1.842e-07 | 42.735 | 0.5577 | 0.0229 | 2.772e-07 | 43.6681
M9 | A | 4 | 1 | H | 0.4577 | 0.024 | 1.924e-07 | 41.6667 | 0.6162 | 0.0235 | 2.926e-07 | 42.5532
M9 | A | 5 | 1 | H | 0.4377 | 0.0239 | 1.968e-07 | 41.841 | 0.5881 | 0.0234 | 2.942e-07 | 42.735
M9 | A | 6 | 1 | H | 0.494 | 0.0242 | 1.967e-07 | 41.3223 | 0.6533 | 0.0237 | 2.985e-07 | 42.1941
M9 | A | 7 | 1 | H | 0.448 | 0.0246 | 2.035e-07 | 40.6504 | 0.5997 | 0.0241 | 3.07e-07 | 41.4938
M9 | A | 8 | 1 | H | 0.3996 | 0.0249 | 2.049e-07 | 40.1606 | 0.5564 | 0.0244 | 3.162e-07 | 40.9836
M9 | A | 9 | 1 | H | 0.4626 | 0.0251 | 2.057e-07 | 39.8406 | 0.6256 | 0.0245 | 3.171e-07 | 40.8163
M9 | A | 10 | 1 | H | 0.4703 | 0.0254 | 2.127e-07 | 39.3701 | 0.6283 | 0.0249 | 3.229e-07 | 40.1606
M9 | A | 11 | 1 | H | 0.423 | 0.0255 | 2.149e-07 | 39.2157 | 0.5742 | 0.025 | 3.28e-07 | 40
M9 | A | 12 | 1 | H | 0.3999 | 0.0259 | 2.291e-07 | 38.61 | 0.5486 | 0.0254 | 3.434e-07 | 39.3701
M9 | A | 13 | 1 | H | 0.3785 | 0.0261 | 2.245e-07 | 38.3142 | 0.5261 | 0.0256 | 3.433e-07 | 39.0625
M9 | A | 14 | 1 | H | 0.4193 | 0.0265 | 2.357e-07 | 37.7358 | 0.5732 | 0.0259 | 3.585e-07 | 38.61
M9 | A | 15 | 1 | H | 0.49 | 0.0266 | 2.279e-07 | 37.594 | 0.6459 | 0.026 | 3.51e-07 | 38.4615
M9 | A | 16 | 1 | H | 0.5349 | 0.0272 | 2.342e-07 | 36.7647 | 0.7011 | 0.0266 | 3.666e-07 | 37.594
M9 | B | 1 | 1 | H | 0.358 | 0.0233 | 1.79e-07 | 42.9185 | 0.5076 | 0.0228 | 2.721e-07 | 43.8596
M9 | B | 2 | 1 | H | 0.3689 | 0.0233 | 1.845e-07 | 42.9185 | 0.513 | 0.0228 | 2.766e-07 | 43.8596
M9 | B | 3 | 1 | H | 0.4078 | 0.0235 | 1.841e-07 | 42.5532 | 0.5588 | 0.023 | 2.784e-07 | 43.4783
M9 | B | 4 | 1 | H | 0.4599 | 0.0241 | 1.951e-07 | 41.4938 | 0.6147 | 0.0236 | 2.938e-07 | 42.3729
M9 | B | 5 | 1 | H | 0.438 | 0.024 | 1.977e-07 | 41.6667 | 0.5887 | 0.0235 | 2.959e-07 | 42.5532
M9 | B | 6 | 1 | H | 0.497 | 0.0243 | 2.002e-07 | 41.1523 | 0.6542 | 0.0238 | 3.003e-07 | 42.0168
M9 | B | 7 | 1 | H | 0.4471 | 0.0247 | 2.035e-07 | 40.4858 | 0.6011 | 0.0242 | 3.086e-07 | 41.3223
M9 | B | 8 | 1 | H | 0.4021 | 0.0249 | 2.076e-07 | 40.1606 | 0.5552 | 0.0244 | 3.156e-07 | 40.9836
M9 | B | 9 | 1 | H | 0.4624 | 0.0251 | 2.063e-07 | 39.8406 | 0.6267 | 0.0246 | 3.194e-07 | 40.6504
M9 | B | 10 | 1 | H | 0.4687 | 0.0255 | 2.126e-07 | 39.2157 | 0.6279 | 0.025 | 3.225e-07 | 40
M9 | B | 11 | 1 | H | 0.4235 | 0.0255 | 2.152e-07 | 39.2157 | 0.575 | 0.025 | 3.303e-07 | 40
M9 | B | 12 | 1 | H | 0.3995 | 0.026 | 2.304e-07 | 38.4615 | 0.5483 | 0.0255 | 3.455e-07 | 39.2157
M9 | B | 13 | 1 | H | 0.3772 | 0.0262 | 2.229e-07 | 38.1679 | 0.5257 | 0.0256 | 3.45e-07 | 39.0625
M9 | B | 14 | 1 | H | 0.4192 | 0.0265 | 2.357e-07 | 37.7358 | 0.5732 | 0.026 | 3.593e-07 | 38.4615
M9 | B | 15 | 1 | H | 0.4907 | 0.0267 | 2.29e-07 | 37.4532 | 0.6455 | 0.0261 | 3.511e-07 | 38.3142
M9 | B | 16 | 1 | H | 0.5351 | 0.0273 | 2.356e-07 | 36.63 | 0.7021 | 0.0267 | 3.683e-07 | 37.4532
M10 | A | 1 | 1 | H | 0.0455 | 0.021 | -9.21e-08 | 47.619 | 0.0806 | 0.0206 | 6.82e-08 | 48.5437
M10 | A | 2 | 1 | H | 0.029 | 0.0207 | -8.63e-08 | 48.3092 | 0.0629 | 0.0203 | 6.99e-08 | 49.2611
M10 | A | 3 | 1 | H | 0.0519 | 0.0208 | -9.27e-08 | 48.0769 | 0.0924 | 0.0204 | 6.69e-08 | 49.0196
M10 | A | 4 | 1 | H | 0.0769 | 0.0209 | -9.01e-08 | 47.8469 | 0.1106 | 0.0205 | 6.89e-08 | 48.7805
M10 | A | 5 | 1 | H | 0.1038 | 0.021 | -9.23e-08 | 47.619 | 0.1542 | 0.0206 | 6.86e-08 | 48.5437
M10 | A | 6 | 1 | H | 0.1324 | 0.0213 | -9.65e-08 | 46.9484 | 0.1745 | 0.0209 | 6.96e-08 | 47.8469
M10 | A | 7 | 1 | H | 0.1407 | 0.0212 | -9.7e-08 | 47.1698 | 0.1879 | 0.0208 | 6.78e-08 | 48.0769
M10 | A | 8 | 1 | H | 0.1042 | 0.0216 | -9.57e-08 | 46.2963 | 0.1435 | 0.0212 | 7.4e-08 | 47.1698
M10 | A | 9 | 1 | H | 0.1026 | 0.0215 | -1e-07 | 46.5116 | 0.1462 | 0.021 | 6.82e-08 | 47.619
M10 | A | 10 | 1 | H | 0.14 | 0.0216 | -9.36e-08 | 46.2963 | 0.1864 | 0.0211 | 7.74e-08 | 47.3934
M10 | A | 11 | 1 | H | 0.1201 | 0.0215 | -9.61e-08 | 46.5116 | 0.1627 | 0.0211 | 7.44e-08 | 47.3934
M10 | A | 12 | 1 | H | 0.1276 | 0.0219 | -9.68e-08 | 45.6621 | 0.1695 | 0.0214 | 7.69e-08 | 46.729
M10 | A | 13 | 1 | H | 0.1103 | 0.0216 | -9.52e-08 | 46.2963 | 0.1513 | 0.0211 | 7.8e-08 | 47.3934
M10 | A | 14 | 1 | H | 0.0749 | 0.0219 | -9.47e-08 | 45.6621 | 0.1121 | 0.0215 | 8.24e-08 | 46.5116
M10 | A | 15 | 1 | H | 0.1112 | 0.0219 | -9.52e-08 | 45.6621 | 0.1585 | 0.0215 | 8.33e-08 | 46.5116
M10 | A | 16 | 1 | H | 0.1272 | 0.0222 | -9.74e-08 | 45.045 | 0.1813 | 0.0218 | 8.66e-08 | 45.8716
M10 | B | 1 | 1 | H | 0.0468 | 0.021 | -9.14e-08 | 47.619 | 0.0822 | 0.0206 | 6.85e-08 | 48.5437
M10 | B | 2 | 1 | H | 0.0269 | 0.0208 | -8.81e-08 | 48.0769 | 0.0631 | 0.0204 | 6.95e-08 | 49.0196
M10 | B | 3 | 1 | H | 0.0523 | 0.0209 | -9.24e-08 | 47.8469 | 0.0924 | 0.0204 | 6.7e-08 | 49.0196
M10 | B | 4 | 1 | H | 0.0765 | 0.0209 | -8.96e-08 | 47.8469 | 0.1089 | 0.0205 | 6.82e-08 | 48.7805
M10 | B | 5 | 1 | H | 0.104 | 0.021 | -9.25e-08 | 47.619 | 0.1554 | 0.0206 | 6.84e-08 | 48.5437
M10 | B | 6 | 1 | H | 0.134 | 0.0214 | -9.59e-08 | 46.729 | 0.1747 | 0.021 | 6.93e-08 | 47.619
M10 | B | 7 | 1 | H | 0.1403 | 0.0213 | -9.76e-08 | 46.9484 | 0.1858 | 0.0209 | 6.69e-08 | 47.8469
M10 | B | 8 | 1 | H | 0.104 | 0.0217 | -9.6e-08 | 46.0829 | 0.1451 | 0.0212 | 7.43e-08 | 47.1698
M10 | B | 9 | 1 | H | 0.1026 | 0.0215 | -1.008e-07 | 46.5116 | 0.148 | 0.0211 | 6.99e-08 | 47.3934
M10 | B | 10 | 1 | H | 0.1402 | 0.0216 | -9.32e-08 | 46.2963 | 0.1851 | 0.0212 | 7.69e-08 | 47.1698
M10 | B | 11 | 1 | H | 0.1218 | 0.0216 | -9.59e-08 | 46.2963 | 0.1625 | 0.0212 | 7.42e-08 | 47.1698
M10 | B | 12 | 1 | H | 0.1266 | 0.0219 | -9.75e-08 | 45.6621 | 0.1706 | 0.0215 | 7.75e-08 | 46.5116
M10 | B | 13 | 1 | H | 0.112 | 0.0216 | -9.45e-08 | 46.2963 | 0.1516 | 0.0212 | 7.76e-08 | 47.1698
M10 | B | 14 | 1 | H | 0.0732 | 0.022 | -9.62e-08 | 45.4545 | 0.1131 | 0.0215 | 8.4e-08 | 46.5116
M10 | B | 15 | 1 | H | 0.111 | 0.022 | -9.53e-08 | 45.4545 | 0.1554 | 0.0215 | 8.2e-08 | 46.5116
M10 | B | 16 | 1 | H | 0.1289 | 0.0222 | -9.68e-08 | 45.045 | 0.1822 | 0.0218 | 8.66e-08 | 45.8716
M11 | A | 1 | 1 | H | 0.0126 | 0.0089 | -2.69e-08 | 112.3596 | 0.0198 | 0.0088 | -2.71e-08 | 113.6364
M11 | A | 2 | 1 | H | 0.0067 | 0.0085 | -2.74e-08 | 117.6471 | 0.0138 | 0.0085 | -2.72e-08 | 117.6471
M11 | A | 3 | 1 | H | 0.0157 | 0.0089 | -2.89e-08 | 112.3596 | 0.0226 | 0.0088 | -2.91e-08 | 113.6364
M11 | A | 4 | 1 | H | 0.0317 | 0.0089 | -3.16e-08 | 112.3596 | 0.0448 | 0.0088 | -3.15e-08 | 113.6364
M11 | A | 5 | 1 | H | 0.0403 | 0.0086 | -3.08e-08 | 116.2791 | 0.0544 | 0.0085 | -3.11e-08 | 117.6471
M11 | A | 6 | 1 | H | 0.034 | 0.0086 | -3.24e-08 | 116.2791 | 0.0472 | 0.0086 | -3.28e-08 | 116.2791
M11 | A | 7 | 1 | H | 0.034 | 0.0087 | -3.5e-08 | 114.9425 | 0.0487 | 0.0086 | -3.57e-08 | 116.2791
M11 | A | 8 | 1 | H | 0.02 | 0.009 | -3.6e-08 | 111.1111 | 0.0327 | 0.0089 | -3.68e-08 | 112.3596
M11 | A | 9 | 1 | H | 0.0246 | 0.0089 | -3.68e-08 | 112.3596 | 0.0373 | 0.0088 | -3.7e-08 | 113.6364
M11 | A | 10 | 1 | H | 0.0354 | 0.0088 | -3.87e-08 | 113.6364 | 0.0503 | 0.0087 | -3.9e-08 | 114.9425
M11 | A | 11 | 1 | H | 0.0206 | 0.0086 | -3.71e-08 | 116.2791 | 0.0309 | 0.0085 | -3.76e-08 | 117.6471
M11 | A | 12 | 1 | H | 0.0092 | 0.0086 | -3.84e-08 | 116.2791 | 0.0214 | 0.0086 | -3.87e-08 | 116.2791
M11 | A | 13 | 1 | H | -0.0001 | 0.0086 | -3.93e-08 | 116.2791 | 0.0078 | 0.0085 | -4.01e-08 | 117.6471
M11 | A | 14 | 1 | H | -0.0057 | 0.0087 | -4.03e-08 | 114.9425 | 0.0034 | 0.0086 | -4.1e-08 | 116.2791
M11 | A | 15 | 1 | H | 0.0251 | 0.0088 | -4.4e-08 | 113.6364 | 0.0385 | 0.0087 | -4.39e-08 | 114.9425
M11 | A | 16 | 1 | H | 0.0259 | 0.0089 | -4.64e-08 | 112.3596 | 0.04 | 0.0089 | -4.7e-08 | 112.3596
M11 | B | 1 | 1 | H | 0.0122 | 0.0089 | -2.73e-08 | 112.3596 | 0.0195 | 0.0088 | -2.73e-08 | 113.6364
M11 | B | 2 | 1 | H | 0.0068 | 0.0085 | -2.74e-08 | 117.6471 | 0.0138 | 0.0085 | -2.75e-08 | 117.6471
M11 | B | 3 | 1 | H | 0.0156 | 0.0089 | -2.93e-08 | 112.3596 | 0.0225 | 0.0088 | -2.92e-08 | 113.6364
M11 | B | 4 | 1 | H | 0.0332 | 0.0089 | -3.13e-08 | 112.3596 | 0.0447 | 0.0088 | -3.16e-08 | 113.6364
M11 | B | 5 | 1 | H | 0.0401 | 0.0086 | -3.11e-08 | 116.2791 | 0.0542 | 0.0085 | -3.11e-08 | 117.6471
M11 | B | 6 | 1 | H | 0.0341 | 0.0086 | -3.26e-08 | 116.2791 | 0.0472 | 0.0086 | -3.29e-08 | 116.2791
M11 | B | 7 | 1 | H | 0.0341 | 0.0087 | -3.51e-08 | 114.9425 | 0.0476 | 0.0086 | -3.6e-08 | 116.2791
M11 | B | 8 | 1 | H | 0.0194 | 0.009 | -3.63e-08 | 111.1111 | 0.0328 | 0.0089 | -3.69e-08 | 112.3596
M11 | B | 9 | 1 | H | 0.025 | 0.0089 | -3.69e-08 | 112.3596 | 0.0376 | 0.0088 | -3.71e-08 | 113.6364
M11 | B | 10 | 1 | H | 0.0352 | 0.0088 | -3.88e-08 | 113.6364 | 0.0503 | 0.0087 | -3.91e-08 | 114.9425
M11 | B | 11 | 1 | H | 0.0204 | 0.0086 | -3.73e-08 | 116.2791 | 0.0311 | 0.0085 | -3.76e-08 | 117.6471
M11 | B | 12 | 1 | H | 0.0087 | 0.0087 | -3.86e-08 | 114.9425 | 0.0214 | 0.0086 | -3.89e-08 | 116.2791
M11 | B | 13 | 1 | H | -0.0006 | 0.0086 | -3.95e-08 | 116.2791 | 0.0082 | 0.0085 | -3.98e-08 | 117.6471
M11 | B | 14 | 1 | H | -0.0055 | 0.0087 | -4.03e-08 | 114.9425 | 0.0033 | 0.0086 | -4.11e-08 | 116.2791
M11 | B | 15 | 1 | H | 0.0249 | 0.0088 | -4.42e-08 | 113.6364 | 0.0384 | 0.0087 | -4.41e-08 | 114.9425
M11 | B | 16 | 1 | H | 0.0258 | 0.0089 | -4.67e-08 | 112.3596 | 0.0398 | 0.0089 | -4.73e-08 | 112.3596
M1 | A | 1 | 1 | L | 16.2069 | 0.1871 | 3.1607e-05 | 5.3447 | 20.9437 | 0.1826 | 4.4307e-05 | 5.4765
M1 | A | 2 | 1 | L | 17.1191 | 0.185 | 3.3005e-05 | 5.4054 | 20.6639 | 0.1835 | 4.3701e-05 | 5.4496
M1 | A | 3 | 1 | L | 16.6503 | 0.185 | 3.1878e-05 | 5.4054 | 22.3735 | 0.1775 | 4.6553e-05 | 5.6338
M1 | A | 4 | 1 | L | 16.5936 | 0.1829 | 3.1066e-05 | 5.4675 | 20.7099 | 0.1799 | 4.2443e-05 | 5.5586
M1 | A | 5 | 1 | L | 15.357 | 0.1857 | 2.8577e-05 | 5.385 | 20.0179 | 0.1811 | 4.1054e-05 | 5.5218
M1 | A | 6 | 1 | L | 15.2112 | 0.1852 | 2.8493e-05 | 5.3996 | 20.0911 | 0.1801 | 4.1065e-05 | 5.5525
M1 | A | 7 | 1 | L | 15.1113 | 0.1857 | 2.7909e-05 | 5.385 | 19.3069 | 0.1826 | 3.9101e-05 | 5.4765
M1 | A | 8 | 1 | L | 16.777 | 0.1798 | 3.0609e-05 | 5.5617 | 20.7249 | 0.1773 | 4.1264e-05 | 5.6402
M1 | A | 9 | 1 | L | 16.5262 | 0.1802 | 2.9829e-05 | 5.5494 | 19.9595 | 0.1786 | 4.0044e-05 | 5.5991
M1 | A | 10 | 1 | L | 16.1942 | 0.181 | 2.9725e-05 | 5.5249 | 20.1025 | 0.1783 | 4.0694e-05 | 5.6085
M1 | A | 11 | 1 | L | 16.6186 | 0.1802 | 3.0857e-05 | 5.5494 | 21.5525 | 0.1744 | 4.3983e-05 | 5.7339
M1 | A | 12 | 1 | L | 15.8444 | 0.1833 | 2.9672e-05 | 5.4555 | 20.5574 | 0.1787 | 4.1776e-05 | 5.596
M1 | A | 13 | 1 | L | 14.5067 | 0.1877 | 2.7125e-05 | 5.3277 | 19.6483 | 0.1822 | 3.9665e-05 | 5.4885
M1 | A | 14 | 1 | L | 15.2661 | 0.1854 | 2.904e-05 | 5.3937 | 20.3517 | 0.18 | 4.151e-05 | 5.5556
M1 | A | 15 | 1 | L | 16.2469 | 0.185 | 3.0997e-05 | 5.4054 | 20.2551 | 0.1823 | 4.2117e-05 | 5.4855
M1 | A | 16 | 1 | L | 16.5014 | 0.1853 | 3.1937e-05 | 5.3967 | 20.7521 | 0.1813 | 4.4267e-05 | 5.5157
M1 | B | 1 | 1 | L | 16.6124 | 0.1904 | 3.3717e-05 | 5.2521 | 21.1619 | 0.186 | 4.7223e-05 | 5.3763
M1 | B | 2 | 1 | L | 15.5408 | 0.1935 | 3.1683e-05 | 5.168 | 20.5067 | 0.1881 | 4.5653e-05 | 5.3163
M1 | B | 3 | 1 | L | 15.8436 | 0.1912 | 3.2001e-05 | 5.2301 | 22.0229 | 0.1824 | 4.8338e-05 | 5.4825
M1 | B | 4 | 1 | L | 16.084 | 0.1881 | 3.1768e-05 | 5.3163 | 22.3066 | 0.1792 | 4.7873e-05 | 5.5804
M1 | B | 5 | 1 | L | 16.9136 | 0.1856 | 3.2747e-05 | 5.3879 | 21.5911 | 0.1805 | 4.6297e-05 | 5.5402
M1 | B | 6 | 1 | L | 16.5449 | 0.186 | 3.1777e-05 | 5.3763 | 20.5775 | 0.183 | 4.3574e-05 | 5.4645
M1 | B | 7 | 1 | L | 16.472 | 0.1861 | 3.1573e-05 | 5.3735 | 21.4005 | 0.1807 | 4.5099e-05 | 5.534
M1 | B | 8 | 1 | L | 16.3144 | 0.1851 | 3.108e-05 | 5.4025 | 21.6713 | 0.1788 | 4.4681e-05 | 5.5928
M1 | B | 9 | 1 | L | 15.7879 | 0.1862 | 2.9737e-05 | 5.3706 | 21.5526 | 0.1783 | 4.4726e-05 | 5.6085
M1 | B | 10 | 1 | L | 14.94 | 0.188 | 2.9195e-05 | 5.3191 | 19.5334 | 0.184 | 4.1331e-05 | 5.4348
M1 | B | 11 | 1 | L | 15.3635 | 0.1873 | 3.0206e-05 | 5.339 | 20.3663 | 0.1817 | 4.352e-05 | 5.5036
M1 | B | 12 | 1 | L | 16.3834 | 0.1859 | 3.17e-05 | 5.3792 | 19.754 | 0.185 | 4.194e-05 | 5.4054
M1 | B | 13 | 1 | L | 16.4016 | 0.1868 | 3.1336e-05 | 5.3533 | 21.1515 | 0.182 | 4.4404e-05 | 5.4945
M1 | B | 14 | 1 | L | 16.0082 | 0.1876 | 3.1181e-05 | 5.3305 | 21.9089 | 0.1795 | 4.6571e-05 | 5.571
M1 | B | 15 | 1 | L | 16.0448 | 0.1894 | 3.2024e-05 | 5.2798 | 22.349 | 0.1801 | 4.8447e-05 | 5.5525
M1 | B | 16 | 1 | L | 16.0971 | 0.1903 | 3.2612e-05 | 5.2549 | 20.6136 | 0.186 | 4.5583e-05 | 5.3763
M2 | A | 1 | 1 | L | 3.4047 | 0.2294 | 1.474e-06 | 4.3592 | 3.1103 | 0.2357 | 2.312e-06 | 4.2427
M2 | A | 2 | 1 | L | 3.0722 | 0.2308 | 1.228e-06 | 4.3328 | 3.03 | 0.2366 | 2.279e-06 | 4.2265
M2 | A | 3 | 1 | L | 3.3296 | 0.2286 | 1.557e-06 | 4.3745 | 2.9088 | 0.235 | 2.29e-06 | 4.2553
M2 | A | 4 | 1 | L | 3.2382 | 0.2285 | 1.515e-06 | 4.3764 | 3.0529 | 0.2345 | 2.354e-06 | 4.2644
M2 | A | 5 | 1 | L | 3.3633 | 0.2277 | 1.532e-06 | 4.3917 | 3.0986 | 0.2338 | 2.409e-06 | 4.2772
M2 | A | 6 | 1 | L | 3.0845 | 0.2274 | 1.275e-06 | 4.3975 | 2.8658 | 0.2335 | 2.088e-06 | 4.2827
M2 | A | 7 | 1 | L | 3.0979 | 0.2262 | 1.242e-06 | 4.4209 | 3.3085 | 0.2314 | 2.471e-06 | 4.3215
M2 | A | 8 | 1 | L | 2.9769 | 0.2256 | 1.262e-06 | 4.4326 | 3.0921 | 0.2308 | 2.486e-06 | 4.3328
M2 | A | 9 | 1 | L | 2.9382 | 0.2253 | 1.314e-06 | 4.4385 | 3.0879 | 0.2306 | 2.482e-06 | 4.3365
M2 | A | 10 | 1 | L | 3.0722 | 0.2257 | 1.189e-06 | 4.4307 | 2.948 | 0.2314 | 2.195e-06 | 4.3215
M2 | A | 11 | 1 | L | 2.9978 | 0.2266 | 1.388e-06 | 4.4131 | 2.9987 | 0.2321 | 2.449e-06 | 4.3085
M2 | A | 12 | 1 | L | 3.1915 | 0.2268 | 1.368e-06 | 4.4092 | 3.2646 | 0.2322 | 2.478e-06 | 4.3066
M2 | A | 13 | 1 | L | 3.3048 | 0.2258 | 1.321e-06 | 4.4287 | 3.3966 | 0.2311 | 2.539e-06 | 4.3271
M2 | A | 14 | 1 | L | 3.1635 | 0.2262 | 1.382e-06 | 4.4209 | 3.4387 | 0.2313 | 2.693e-06 | 4.3234
M2 | A | 15 | 1 | L | 3.2869 | 0.2265 | 1.416e-06 | 4.415 | 3.0681 | 0.2325 | 2.319e-06 | 4.3011
M2 | A | 16 | 1 | L | 3.2985 | 0.2257 | 1.459e-06 | 4.4307 | 3.2554 | 0.2314 | 2.56e-06 | 4.3215
M2 | B | 1 | 1 | L | 3.2248 | 0.2319 | 1.388e-06 | 4.3122 | 3.085 | 0.238 | 2.379e-06 | 4.2017
M2 | B | 2 | 1 | L | 3.3283 | 0.2324 | 1.467e-06 | 4.3029 | 2.9715 | 0.239 | 2.261e-06 | 4.1841
M2 | B | 3 | 1 | L | 3.1149 | 0.2312 | 1.381e-06 | 4.3253 | 2.9867 | 0.2371 | 2.433e-06 | 4.2176
M2 | B | 4 | 1 | L | 3.2556 | 0.2305 | 1.547e-06 | 4.3384 | 2.8778 | 0.2371 | 2.242e-06 | 4.2176
M2 | B | 5 | 1 | L | 3.2682 | 0.23 | 1.463e-06 | 4.3478 | 2.9764 | 0.2362 | 2.373e-06 | 4.2337
M2 | B | 6 | 1 | L | 3.3616 | 0.2288 | 1.564e-06 | 4.3706 | 3.1148 | 0.2351 | 2.397e-06 | 4.2535
M2 | B | 7 | 1 | L | 3.1335 | 0.2281 | 1.355e-06 | 4.384 | 3.3122 | 0.2335 | 2.537e-06 | 4.2827
M2 | B | 8 | 1 | L | 2.8565 | 0.2278 | 1.165e-06 | 4.3898 | 3.1715 | 0.2328 | 2.563e-06 | 4.2955
M2 | B | 9 | 1 | L | 3.176 | 0.2268 | 1.511e-06 | 4.4092 | 3.0147 | 0.2328 | 2.449e-06 | 4.2955
M2 | B | 10 | 1 | L | 3.047 | 0.2276 | 1.241e-06 | 4.3937 | 2.917 | 0.2335 | 2.233e-06 | 4.2827
M2 | B | 11 | 1 | L | 3.191 | 0.2282 | 1.601e-06 | 4.3821 | 3.2616 | 0.2336 | 2.774e-06 | 4.2808
M2 | B | 12 | 1 | L | 3.4743 | 0.2282 | 1.608e-06 | 4.3821 | 3.3544 | 0.2341 | 2.589e-06 | 4.2717
M2 | B | 13 | 1 | L | 3.3777 | 0.2275 | 1.494e-06 | 4.3956 | 3.3417 | 0.2333 | 2.531e-06 | 4.2863
M2 | B | 14 | 1 | L | 3.4448 | 0.2276 | 1.652e-06 | 4.3937 | 3.3914 | 0.2335 | 2.712e-06 | 4.2827
M2 | B | 15 | 1 | L | 3.3236 | 0.2283 | 1.5e-06 | 4.3802 | 3.1915 | 0.2343 | 2.44e-06 | 4.268
M2 | B | 16 | 1 | L | 3.501 | 0.2273 | 1.671e-06 | 4.3995 | 3.1144 | 0.2337 | 2.472e-06 | 4.279
M3 | A | 1 | 1 | L | -3.9074 | 0.2393 | -1.718e-06 | 4.1789 | -3.159 | 0.2427 | -1.098e-06 | 4.1203
M3 | A | 2 | 1 | L | -3.9054 | 0.24 | -1.788e-06 | 4.1667 | -3.1929 | 0.2434 | -1.236e-06 | 4.1085
M3 | A | 3 | 1 | L | -4.0657 | 0.2394 | -1.808e-06 | 4.1771 | -3.4155 | 0.2428 | -1.262e-06 | 4.1186
M3 | A | 4 | 1 | L | -4.1631 | 0.2386 | -1.75e-06 | 4.1911 | -3.9668 | 0.2427 | -1.45e-06 | 4.1203
M3 | A | 5 | 1 | L | -4.1296 | 0.2382 | -1.749e-06 | 4.1982 | -4.1671 | 0.2427 | -1.627e-06 | 4.1203
M3 | A | 6 | 1 | L | -4.0043 | 0.2385 | -1.731e-06 | 4.1929 | -3.7001 | 0.2423 | -1.402e-06 | 4.1271
M3 | A | 7 | 1 | L | -3.9692 | 0.2388 | -1.737e-06 | 4.1876 | -3.3019 | 0.2421 | -1.203e-06 | 4.1305
M3 | A | 8 | 1 | L | -3.9673 | 0.2392 | -1.777e-06 | 4.1806 | -3.2681 | 0.2424 | -1.201e-06 | 4.1254
M3 | A | 9 | 1 | L | -4.1165 | 0.2396 | -1.781e-06 | 4.1736 | -3.4231 | 0.2429 | -1.229e-06 | 4.1169
M3 | A | 10 | 1 | L | -4.0464 | 0.2392 | -1.755e-06 | 4.1806 | -3.4208 | 0.2426 | -1.222e-06 | 4.122
M3 | A | 11 | 1 | L | -4.0469 | 0.2392 | -1.697e-06 | 4.1806 | -3.3296 | 0.2423 | -1.09e-06 | 4.1271
M3 | A | 12 | 1 | L | -3.9093 | 0.2399 | -1.83e-06 | 4.1684 | -3.2735 | 0.2433 | -1.293e-06 | 4.1102
M3 | A | 13 | 1 | L | -3.9459 | 0.2413 | -1.884e-06 | 4.1442 | -3.3748 | 0.2447 | -1.346e-06 | 4.0866
M3 | A | 14 | 1 | L | -4.1041 | 0.2423 | -1.85e-06 | 4.1271 | -3.5254 | 0.2458 | -1.308e-06 | 4.0683
M3 | A | 15 | 1 | L | -4.1514 | 0.2434 | -1.868e-06 | 4.1085 | -3.6143 | 0.2469 | -1.336e-06 | 4.0502
M3 | A | 16 | 1 | L | -4.2008 | 0.2443 | -1.897e-06 | 4.0933 | -3.5891 | 0.2477 | -1.326e-06 | 4.0371
M3 | B | 1 | 1 | L | -3.9258 | 0.2402 | -1.73e-06 | 4.1632 | -3.1875 | 0.2437 | -1.119e-06 | 4.1034
M3 | B | 2 | 1 | L | -3.9254 | 0.2408 | -1.803e-06 | 4.1528 | -3.2272 | 0.2444 | -1.262e-06 | 4.0917
M3 | B | 3 | 1 | L | -4.0759 | 0.2403 | -1.826e-06 | 4.1615 | -3.4105 | 0.2437 | -1.27e-06 | 4.1034
M3 | B | 4 | 1 | L | -4.1523 | 0.2395 | -1.758e-06 | 4.1754 | -3.5712 | 0.2429 | -1.207e-06 | 4.1169
M3 | B | 5 | 1 | L | -4.09 | 0.239 | -1.745e-06 | 4.1841 | -4.0146 | 0.2433 | -1.538e-06 | 4.1102
M3 | B | 6 | 1 | L | -4.0061 | 0.2393 | -1.749e-06 | 4.1789 | -3.3145 | 0.2426 | -1.173e-06 | 4.122
M3 | B | 7 | 1 | L | -3.9572 | 0.2396 | -1.743e-06 | 4.1736 | -3.3173 | 0.243 | -1.23e-06 | 4.1152
M3 | B | 8 | 1 | L | -3.9855 | 0.2401 | -1.793e-06 | 4.1649 | -3.2795 | 0.2433 | -1.219e-06 | 4.1102
M3 | B | 9 | 1 | L | -4.1168 | 0.2404 | -1.789e-06 | 4.1597 | -3.4113 | 0.2437 | -1.234e-06 | 4.1034
M3 | B | 10 | 1 | L | -4.051 | 0.24 | -1.762e-06 | 4.1667 | -3.4292 | 0.2435 | -1.239e-06 | 4.1068
M3 | B | 11 | 1 | L | -4.0304 | 0.2399 | -1.703e-06 | 4.1684 | -3.3166 | 0.2431 | -1.101e-06 | 4.1135
M3 | B | 12 | 1 | L | -3.9223 | 0.2407 | -1.839e-06 | 4.1545 | -3.2769 | 0.2442 | -1.314e-06 | 4.095
M3 | B | 13 | 1 | L | -3.9588 | 0.2421 | -1.897e-06 | 4.1305 | -3.377 | 0.2455 | -1.345e-06 | 4.0733
M3 | B | 14 | 1 | L | -4.0908 | 0.2431 | -1.843e-06 | 4.1135 | -3.5513 | 0.2467 | -1.336e-06 | 4.0535
M3 | B | 15 | 1 | L | -4.0992 | 0.2441 | -1.847e-06 | 4.0967 | -3.6226 | 0.2478 | -1.341e-06 | 4.0355
M3 | B | 16 | 1 | L | -4.1637 | 0.245 | -1.89e-06 | 4.0816 | -3.5632 | 0.2485 | -1.341e-06 | 4.0241
M4 | A | 1 | 1 | L | 2.2027 | 0.2205 | 6.01e-07 | 4.5351 | 3.527 | 0.2234 | 3.71e-07 | 4.4763
M4 | A | 2 | 1 | L | 2.0667 | 0.2211 | 5.48e-07 | 4.5228 | 3.4141 | 0.224 | 2.92e-07 | 4.4643
M4 | A | 3 | 1 | L | 1.9907 | 0.2215 | 5.49e-07 | 4.5147 | 3.3611 | 0.2244 | 2.81e-07 | 4.4563
M4 | A | 4 | 1 | L | 1.9387 | 0.2214 | 6.24e-07 | 4.5167 | 3.3011 | 0.2243 | 3.7e-07 | 4.4583
M4 | A | 5 | 1 | L | 1.955 | 0.2211 | 6.08e-07 | 4.5228 | 3.2423 | 0.2241 | 3.33e-07 | 4.4623
M4 | A | 6 | 1 | L | 1.9808 | 0.2212 | 6.15e-07 | 4.5208 | 3.3047 | 0.2241 | 3.45e-07 | 4.4623
M4 | A | 7 | 1 | L | 2.0769 | 0.2217 | 5.96e-07 | 4.5106 | 3.3172 | 0.2248 | 2.77e-07 | 4.4484
M4 | A | 8 | 1 | L | 2.101 | 0.2221 | 6.06e-07 | 4.5025 | 3.4503 | 0.225 | 3.16e-07 | 4.4444
M4 | A | 9 | 1 | L | 1.9743 | 0.2221 | 6.35e-07 | 4.5025 | 3.2786 | 0.2251 | 3.24e-07 | 4.4425
M4 | A | 10 | 1 | L | 2.0311 | 0.223 | 6.32e-07 | 4.4843 | 3.3567 | 0.226 | 3.45e-07 | 4.4248
M4 | A | 11 | 1 | L | 1.9881 | 0.223 | 5.78e-07 | 4.4843 | 3.4336 | 0.2258 | 3.12e-07 | 4.4287
M4 | A | 12 | 1 | L | 2.0149 | 0.2245 | 5.88e-07 | 4.4543 | 3.3692 | 0.2274 | 3.29e-07 | 4.3975
M4 | A | 13 | 1 | L | 2.0462 | 0.2257 | 5.92e-07 | 4.4307 | 3.3866 | 0.2286 | 3.15e-07 | 4.3745
M4 | A | 14 | 1 | L | 1.8498 | 0.2264 | 6.25e-07 | 4.417 | 3.2011 | 0.2293 | 3.58e-07 | 4.3611
M4 | A | 15 | 1 | L | 1.9159 | 0.2265 | 6.65e-07 | 4.415 | 3.1472 | 0.2296 | 3.28e-07 | 4.3554
M4 | A | 16 | 1 | L | 1.8743 | 0.2261 | 6.54e-07 | 4.4228 | 3.0993 | 0.2293 | 3.14e-07 | 4.3611
M4 | B | 1 | 1 | L | 2.2151 | 0.2203 | 6.11e-07 | 4.5393 | 3.5269 | 0.2233 | 3.74e-07 | 4.4783
M4 | B | 2 | 1 | L | 2.0507 | 0.221 | 5.42e-07 | 4.5249 | 3.4252 | 0.2239 | 2.98e-07 | 4.4663
M4 | B | 3 | 1 | L | 1.9931 | 0.2213 | 5.43e-07 | 4.5188 | 3.3819 | 0.2242 | 2.95e-07 | 4.4603
M4 | B | 4 | 1 | L | 1.9474 | 0.2212 | 6.27e-07 | 4.5208 | 3.2942 | 0.2241 | 3.68e-07 | 4.4623
M4 | B | 5 | 1 | L | 1.9593 | 0.2209 | 6.13e-07 | 4.5269 | 3.2473 | 0.2239 | 3.38e-07 | 4.4663
M4 | B | 6 | 1 | L | 1.9827 | 0.221 | 6.12e-07 | 4.5249 | 3.2967 | 0.2239 | 3.41e-07 | 4.4663
M4 | B | 7 | 1 | L | 2.0726 | 0.2215 | 5.89e-07 | 4.5147 | 3.3301 | 0.2246 | 2.87e-07 | 4.4524
M4 | B | 8 | 1 | L | 2.1067 | 0.2218 | 6.03e-07 | 4.5086 | 3.4577 | 0.2248 | 3.21e-07 | 4.4484
M4 | B | 9 | 1 | L | 1.9765 | 0.2219 | 6.33e-07 | 4.5065 | 3.2856 | 0.2249 | 3.27e-07 | 4.4464
M4 | B | 10 | 1 | L | 2.0167 | 0.2228 | 6.23e-07 | 4.4883 | 3.3801 | 0.2258 | 3.55e-07 | 4.4287
M4 | B | 11 | 1 | L | 1.9856 | 0.2227 | 5.84e-07 | 4.4903 | 3.4481 | 0.2256 | 3.14e-07 | 4.4326
M4 | B | 12 | 1 | L | 2.0184 | 0.2243 | 5.85e-07 | 4.4583 | 3.3653 | 0.2272 | 3.31e-07 | 4.4014
M4 | B | 13 | 1 | L | 2.0529 | 0.2254 | 5.98e-07 | 4.4366 | 3.3971 | 0.2284 | 3.21e-07 | 4.3783
M4 | B | 14 | 1 | L | 1.8578 | 0.2261 | 6.28e-07 | 4.4228 | 3.2068 | 0.2291 | 3.6e-07 | 4.3649
M4 | B | 15 | 1 | L | 1.9041 | 0.2263 | 6.52e-07 | 4.4189 | 3.1512 | 0.2295 | 3.29e-07 | 4.3573
M4 | B | 16 | 1 | L | 1.8752 | 0.2259 | 6.55e-07 | 4.4267 | 3.1117 | 0.229 | 3.22e-07 | 4.3668
M5 | A | 1 | 1 | L | 3.6538 | 0.2016 | -3e-09 | 4.9603 | 4.9315 | 0.2054 | -5.34e-07 | 4.8685
M5 | A | 2 | 1 | L | 3.6802 | 0.2027 | -1.41e-07 | 4.9334 | 4.9021 | 0.2065 | -7.28e-07 | 4.8426
M5 | A | 3 | 1 | L | 3.6783 | 0.202 | -1.18e-07 | 4.9505 | 4.9802 | 0.2055 | -6.78e-07 | 4.8662
M5 | A | 4 | 1 | L | 3.4041 | 0.2021 | -3.33e-07 | 4.948 | 4.8027 | 0.2052 | -7.14e-07 | 4.8733
M5 | A | 5 | 1 | L | 3.3398 | 0.202 | -3.68e-07 | 4.9505 | 4.8681 | 0.2049 | -6.83e-07 | 4.8804
M5 | A | 6 | 1 | L | 3.5975 | 0.2016 | -1.13e-07 | 4.9603 | 4.8714 | 0.2051 | -6.93e-07 | 4.8757
M5 | A | 7 | 1 | L | 3.7032 | 0.2025 | -1.64e-07 | 4.9383 | 4.9369 | 0.206 | -7.45e-07 | 4.8544
M5 | A | 8 | 1 | L | 3.6484 | 0.2016 | -1.54e-07 | 4.9603 | 4.9411 | 0.205 | -7.04e-07 | 4.878
M5 | A | 9 | 1 | L | 3.3567 | 0.2022 | -4.32e-07 | 4.9456 | 4.9427 | 0.2051 | -7.29e-07 | 4.8757
M5 | A | 10 | 1 | L | 3.686 | 0.2021 | -1.52e-07 | 4.948 | 4.9002 | 0.2056 | -7.32e-07 | 4.8638
M5 | A | 11 | 1 | L | 3.6809 | 0.2023 | -1.11e-07 | 4.9432 | 4.9038 | 0.2058 | -6.93e-07 | 4.8591
M5 | A | 12 | 1 | L | 3.4688 | 0.2029 | -2.66e-07 | 4.9285 | 4.8242 | 0.2061 | -7.27e-07 | 4.852
M5 | A | 13 | 1 | L | 3.6357 | 0.2031 | -2.01e-07 | 4.9237 | 4.9 | 0.2066 | -7.49e-07 | 4.8403
M5 | A | 14 | 1 | L | 3.5991 | 0.2037 | -1.53e-07 | 4.9092 | 4.861 | 0.2073 | -7.53e-07 | 4.8239
M5 | A | 15 | 1 | L | 3.5561 | 0.2041 | -1.49e-07 | 4.8996 | 4.8465 | 0.2077 | -7.3e-07 | 4.8146
M5 | A | 16 | 1 | L | 3.6555 | 0.2037 | -1.22e-07 | 4.9092 | 4.8008 | 0.2073 | -7.32e-07 | 4.8239
M5 | B | 1 | 1 | L | 3.347 | 0.2014 | -2.92e-07 | 4.9652 | 4.9516 | 0.2044 | -5.19e-07 | 4.8924
M5 | B | 2 | 1 | L | 3.5672 | 0.2021 | -2.64e-07 | 4.948 | 4.9078 | 0.2056 | -7.11e-07 | 4.8638
M5 | B | 3 | 1 | L | 3.5082 | 0.2014 | -2.7e-07 | 4.9652 | 5.0054 | 0.2045 | -6.63e-07 | 4.89
M5 | B | 4 | 1 | L | 3.4694 | 0.201 | -2.7e-07 | 4.9751 | 4.8288 | 0.2042 | -6.89e-07 | 4.8972
M5 | B | 5 | 1 | L | 3.494 | 0.2007 | -2.26e-07 | 4.9826 | 4.8555 | 0.204 | -6.91e-07 | 4.902
M5 | B | 6 | 1 | L | 3.5567 | 0.2007 | -1.33e-07 | 4.9826 | 4.8379 | 0.2043 | -7.09e-07 | 4.8948
M5 | B | 7 | 1 | L | 3.7175 | 0.2015 | -1.54e-07 | 4.9628 | 4.9499 | 0.2051 | -7.36e-07 | 4.8757
M5 | B | 8 | 1 | L | 3.6167 | 0.2007 | -1.9e-07 | 4.9826 | 4.9599 | 0.2041 | -6.89e-07 | 4.8996
M5 | B | 9 | 1 | L | 3.5716 | 0.2008 | -2.39e-07 | 4.9801 | 4.9349 | 0.2042 | -7.34e-07 | 4.8972
M5 | B | 10 | 1 | L | 3.3251 | 0.202 | -4.94e-07 | 4.9505 | 4.8888 | 0.2047 | -7.39e-07 | 4.8852
M5 | B | 11 | 1 | L | 3.3888 | 0.2019 | -3.7e-07 | 4.9529 | 4.8969 | 0.2049 | -6.9e-07 | 4.8804
M5 | B | 12 | 1 | L | 3.4139 | 0.2021 | -3.28e-07 | 4.948 | 4.7853 | 0.2052 | -7.32e-07 | 4.8733
M5 | B | 13 | 1 | L | 3.7019 | 0.2021 | -1.65e-07 | 4.948 | 4.9016 | 0.2056 | -7.4e-07 | 4.8638
M5 | B | 14 | 1 | L | 3.6304 | 0.2027 | -1.42e-07 | 4.9334 | 4.843 | 0.2064 | -7.5e-07 | 4.845
M5 | B | 15 | 1 | L | 3.6104 | 0.2031 | -1.22e-07 | 4.9237 | 4.8651 | 0.2067 | -7.12e-07 | 4.8379
M5 | B | 16 | 1 | L | 3.4508 | 0.2032 | -3.05e-07 | 4.9213 | 4.8095 | 0.2064 | -7.22e-07 | 4.845
M7 | A | 1 | 1 | L | -0.5772 | 0.1149 | -6.82e-07 | 8.7032 | -1.3689 | 0.1189 | -1.35e-06 | 8.4104
M7 | A | 2 | 1 | L | -0.5344 | 0.115 | -7.78e-07 | 8.6957 | -1.4002 | 0.1192 | -1.496e-06 | 8.3893
M7 | A | 3 | 1 | L | -0.6075 | 0.1146 | -7.65e-07 | 8.726 | -1.3792 | 0.1186 | -1.462e-06 | 8.4317
M7 | A | 4 | 1 | L | -0.6175 | 0.1147 | -7.77e-07 | 8.7184 | -1.3852 | 0.1186 | -1.474e-06 | 8.4317
M7 | A | 5 | 1 | L | -0.6801 | 0.1149 | -7.84e-07 | 8.7032 | -1.4808 | 0.1189 | -1.492e-06 | 8.4104
M7 | A | 6 | 1 | L | -0.6354 | 0.1148 | -8.02e-07 | 8.7108 | -1.3685 | 0.1187 | -1.501e-06 | 8.4246
M7 | A | 7 | 1 | L | -0.652 | 0.115 | -8.15e-07 | 8.6957 | -1.4211 | 0.119 | -1.515e-06 | 8.4034
M7 | A | 8 | 1 | L | -0.6715 | 0.1152 | -8.28e-07 | 8.6806 | -1.4442 | 0.1192 | -1.53e-06 | 8.3893
M7 | A | 9 | 1 | L | -0.637 | 0.1154 | -8.47e-07 | 8.6655 | -1.4108 | 0.1193 | -1.547e-06 | 8.3822
M7 | A | 10 | 1 | L | -0.6869 | 0.1154 | -8.43e-07 | 8.6655 | -1.4666 | 0.1194 | -1.545e-06 | 8.3752
M7 | A | 11 | 1 | L | -0.7566 | 0.1154 | -8.57e-07 | 8.6655 | -1.5384 | 0.1193 | -1.563e-06 | 8.3822
M7 | A | 12 | 1 | L | -0.7658 | 0.1155 | -8.48e-07 | 8.658 | -1.5378 | 0.1195 | -1.549e-06 | 8.3682
M7 | A | 13 | 1 | L | -0.7195 | 0.1159 | -8.7e-07 | 8.6281 | -1.5081 | 0.1198 | -1.576e-06 | 8.3472
M7 | A | 14 | 1 | L | -0.7521 | 0.116 | -8.68e-07 | 8.6207 | -1.5139 | 0.1201 | -1.589e-06 | 8.3264
M7 | A | 15 | 1 | L | -0.8287 | 0.1163 | -8.66e-07 | 8.5985 | -1.6254 | 0.1203 | -1.588e-06 | 8.3126
M7 | A | 16 | 1 | L | -0.7697 | 0.116 | -8.89e-07 | 8.6207 | -1.5881 | 0.12 | -1.599e-06 | 8.3333
M7 | B | 1 | 1 | L | -0.5834 | 0.115 | -6.85e-07 | 8.6957 | -1.3737 | 0.119 | -1.353e-06 | 8.4034
M7 | B | 2 | 1 | L | -0.5365 | 0.1151 | -7.79e-07 | 8.6881 | -1.4007 | 0.1192 | -1.496e-06 | 8.3893
M7 | B | 3 | 1 | L | -0.605 | 0.1147 | -7.66e-07 | 8.7184 | -1.378 | 0.1187 | -1.466e-06 | 8.4246
M7 | B | 4 | 1 | L | -0.6191 | 0.1148 | -7.78e-07 | 8.7108 | -1.3896 | 0.1187 | -1.477e-06 | 8.4246
M7 | B | 5 | 1 | L | -0.6695 | 0.115 | -7.84e-07 | 8.6957 | -1.4761 | 0.119 | -1.493e-06 | 8.4034
M7 | B | 6 | 1 | L | -0.6361 | 0.1149 | -8.03e-07 | 8.7032 | -1.3653 | 0.1188 | -1.503e-06 | 8.4175
M7 | B | 7 | 1 | L | -0.644 | 0.1151 | -8.15e-07 | 8.6881 | -1.4181 | 0.1191 | -1.515e-06 | 8.3963
M7 | B | 8 | 1 | L | -0.6658 | 0.1153 | -8.27e-07 | 8.673 | -1.4414 | 0.1192 | -1.53e-06 | 8.3893
M7 | B | 9 | 1 | L | -0.6306 | 0.1155 | -8.46e-07 | 8.658 | -1.4194 | 0.1194 | -1.552e-06 | 8.3752
M7 | B | 10 | 1 | L | -0.6813 | 0.1155 | -8.43e-07 | 8.658 | -1.4603 | 0.1195 | -1.544e-06 | 8.3682
M7 | B | 11 | 1 | L | -0.7573 | 0.1155 | -8.59e-07 | 8.658 | -1.5414 | 0.1194 | -1.566e-06 | 8.3752
M7 | B | 12 | 1 | L | -0.7641 | 0.1156 | -8.49e-07 | 8.6505 | -1.541 | 0.1196 | -1.552e-06 | 8.3612
M7 | B | 13 | 1 | L | -0.7141 | 0.116 | -8.69e-07 | 8.6207 | -1.5173 | 0.1199 | -1.579e-06 | 8.3403
M7 | B | 14 | 1 | L | -0.7479 | 0.1162 | -8.68e-07 | 8.6059 | -1.5112 | 0.1201 | -1.591e-06 | 8.3264
M7 | B | 15 | 1 | L | -0.8314 | 0.1164 | -8.69e-07 | 8.5911 | -1.6233 | 0.1204 | -1.587e-06 | 8.3056
M7 | B | 16 | 1 | L | -0.769 | 0.1161 | -8.9e-07 | 8.6133 | -1.5861 | 0.1201 | -1.6e-06 | 8.3264

## DNB rad coeff
Band | HAM | Gain State | Agg Zone | Band Median |  | Band Average
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | c1 | c1 | c1 | c1
LGS | A | 1 | 1 | 1.2066e-06 | 1.2514e-06 | 1.2062e-06 | 1.2489e-06
LGS | A | 1 | 7 | 1.8468e-06 | 1.9273e-06 | 1.8459e-06 | 1.9252e-06
LGS | A | 1 | 11 | 2.6716e-06 | 2.7647e-06 | 2.6695e-06 | 2.7639e-06
LGS | A | 1 | 15 | 4.1312e-06 | 4.2752e-06 | 4.1247e-06 | 4.2692e-06
LGS | A | 1 | 20 | 6.1468e-06 | 6.3635e-06 | 6.1385e-06 | 6.352e-06
LGS | A | 1 | 26 | 9.7452e-06 | 1.0072e-05 | 9.7305e-06 | 1.0065e-05
LGS | A | 1 | 32 | 1.5425e-05 | 1.5934e-05 | 1.5419e-05 | 1.5938e-05
LGS | A | 1 | 34 | 6.2684e-07 | 6.4984e-07 | 6.2656e-07 | 6.4849e-07
LGS | B | 1 | 1 | 1.2041e-06 | 1.2497e-06 | 1.2039e-06 | 1.2474e-06
LGS | B | 1 | 7 | 1.8433e-06 | 1.9247e-06 | 1.8424e-06 | 1.9231e-06
LGS | B | 1 | 11 | 2.6669e-06 | 2.7615e-06 | 2.6646e-06 | 2.7608e-06
LGS | B | 1 | 15 | 4.1235e-06 | 4.2701e-06 | 4.117e-06 | 4.2646e-06
LGS | B | 1 | 20 | 6.1352e-06 | 6.3566e-06 | 6.1275e-06 | 6.3452e-06
LGS | B | 1 | 26 | 9.7265e-06 | 1.006e-05 | 9.7127e-06 | 1.0054e-05
LGS | B | 1 | 32 | 1.5395e-05 | 1.5917e-05 | 1.5389e-05 | 1.5921e-05
LGS | B | 1 | 34 | 6.2563e-07 | 6.4911e-07 | 6.254e-07 | 6.4775e-07
MGS | A | 2 | 1 | 3.553e-09 | 3.6013e-09 | 3.5611e-09 | 3.6099e-09
MGS | A | 2 | 7 | 5.426e-09 | 5.5462e-09 | 5.441e-09 | 5.5618e-09
MGS | A | 2 | 11 | 7.9449e-09 | 8.164e-09 | 7.9686e-09 | 8.1884e-09
MGS | A | 2 | 15 | 1.2297e-08 | 1.263e-08 | 1.2335e-08 | 1.2669e-08
MGS | A | 2 | 20 | 1.835e-08 | 1.8848e-08 | 1.8398e-08 | 1.8895e-08
MGS | A | 2 | 26 | 2.9111e-08 | 3.0088e-08 | 2.9195e-08 | 3.0172e-08
MGS | A | 2 | 32 | 4.5798e-08 | 4.7376e-08 | 4.5904e-08 | 4.7493e-08
MGS | A | 2 | 34 | 1.8544e-09 | 1.8794e-09 | 1.8587e-09 | 1.884e-09
MGS | B | 2 | 1 | 3.5461e-09 | 3.5948e-09 | 3.5541e-09 | 3.6032e-09
MGS | B | 2 | 7 | 5.4151e-09 | 5.5358e-09 | 5.4298e-09 | 5.5512e-09
MGS | B | 2 | 11 | 7.9279e-09 | 8.1472e-09 | 7.9517e-09 | 8.1717e-09
MGS | B | 2 | 15 | 1.2597e-08 | 1.2899e-08 | 1.2706e-08 | 1.3049e-08
MGS | B | 2 | 20 | 1.7833e-08 | 1.8381e-08 | 1.831e-08 | 1.8812e-08
MGS | B | 2 | 26 | 2.9071e-08 | 3.0027e-08 | 2.9139e-08 | 3.0111e-08
MGS | B | 2 | 32 | 4.5697e-08 | 4.7278e-08 | 4.5802e-08 | 4.7397e-08
MGS | B | 2 | 34 | 1.8509e-09 | 1.8765e-09 | 1.8554e-09 | 1.8809e-09
HGA | A | 3 | 1 | 1.4214e-11 | 1.4379e-11 | 1.4233e-11 | 1.4387e-11
HGA | A | 3 | 7 | 2.1894e-11 | 2.2185e-11 | 2.1929e-11 | 2.2155e-11
HGA | A | 3 | 11 | 3.1725e-11 | 3.1955e-11 | 3.1744e-11 | 3.195e-11
HGA | A | 3 | 15 | 4.9204e-11 | 4.9321e-11 | 4.9174e-11 | 4.9382e-11
HGA | A | 3 | 20 | 7.336e-11 | 7.3617e-11 | 7.3366e-11 | 7.3686e-11
HGA | A | 3 | 26 | 1.1668e-10 | 1.1724e-10 | 1.167e-10 | 1.1724e-10
HGA | A | 3 | 32 | 1.8307e-10 | 1.8492e-10 | 1.8339e-10 | 1.8432e-10
HGA | A | 3 | 34 | 7.3477e-12 | 7.4382e-12 | 7.3561e-12 | 7.4403e-12
HGA | B | 3 | 1 | 1.421e-11 | 1.437e-11 | 1.4225e-11 | 1.4377e-11
HGA | B | 3 | 7 | 2.1875e-11 | 2.2169e-11 | 2.1908e-11 | 2.2137e-11
HGA | B | 3 | 11 | 3.17e-11 | 3.193e-11 | 3.1717e-11 | 3.1924e-11
HGA | B | 3 | 15 | 4.9154e-11 | 4.9264e-11 | 4.9123e-11 | 4.9326e-11
HGA | B | 3 | 20 | 7.3287e-11 | 7.3547e-11 | 7.3291e-11 | 7.3612e-11
HGA | B | 3 | 26 | 1.1654e-10 | 1.171e-10 | 1.1656e-10 | 1.171e-10
HGA | B | 3 | 32 | 1.8281e-10 | 1.8466e-10 | 1.7921e-10 | 1.8007e-10
HGA | B | 3 | 34 | 7.3435e-12 | 7.4345e-12 | 7.3523e-12 | 7.4367e-12
HGB | A | 3 | 1 | 1.4483e-11 | 1.4568e-11 | 1.4528e-11 | 1.4615e-11
HGB | A | 3 | 7 | 2.2273e-11 | 2.2427e-11 | 2.2363e-11 | 2.2519e-11
HGB | A | 3 | 11 | 3.2297e-11 | 3.233e-11 | 3.2414e-11 | 3.2456e-11
HGB | A | 3 | 15 | 4.9901e-11 | 4.9864e-11 | 5.0152e-11 | 5.0117e-11
HGB | A | 3 | 20 | 7.4631e-11 | 7.4585e-11 | 7.4862e-11 | 7.4831e-11
HGB | A | 3 | 26 | 1.1866e-10 | 1.1861e-10 | 1.1888e-10 | 1.1884e-10
HGB | A | 3 | 32 | 1.8622e-10 | 1.8681e-10 | 1.8642e-10 | 1.865e-10
HGB | A | 3 | 34 | 7.4866e-12 | 7.5323e-12 | 7.5037e-12 | 7.5516e-12
HGB | B | 3 | 1 | 1.4467e-11 | 1.4556e-11 | 1.4516e-11 | 1.4605e-11
HGB | B | 3 | 7 | 2.2249e-11 | 2.2402e-11 | 2.2339e-11 | 2.2499e-11
HGB | B | 3 | 11 | 3.2256e-11 | 3.2305e-11 | 3.2381e-11 | 3.2428e-11
HGB | B | 3 | 15 | 4.9841e-11 | 4.9812e-11 | 5.0091e-11 | 5.0063e-11
HGB | B | 3 | 20 | 7.454e-11 | 7.45e-11 | 7.4768e-11 | 7.4744e-11
HGB | B | 3 | 26 | 1.185e-10 | 1.1846e-10 | 1.1872e-10 | 1.1868e-10
HGB | B | 3 | 32 | 1.8594e-10 | 1.8656e-10 | 1.8614e-10 | 1.8623e-10
HGB | B | 3 | 34 | 7.4779e-12 | 7.5275e-12 | 7.4965e-12 | 7.5457e-12
Band | HAM | Gain State | Agg Zone | Detector | Detector Dependent |  |  | Notes:
[5 empty] | Eside A | Eside B |  | 1. This table provides VIIRS F2 DNB radiometric coefficients.
[5 empty] | c1 | c1 |  | 2. The results were derived from RC-1 using the SIS-100 source.
LGS | A | 1 | 1 | 1 | 1.2045e-06 | 1.2553e-06 |  | 3. Linear fits of the SIS-100 radiance vs dn were used to determine the radiometric coefficients.
LGS | A | 1 | 1 | 2 | 1.2072e-06 | 1.2553e-06 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
LGS | A | 1 | 1 | 3 | 1.2168e-06 | 1.2652e-06 |  | 5. The test was set to diagnostic mode.
LGS | A | 1 | 1 | 4 | 1.2128e-06 | 1.261e-06 |  | 6. Units of c1 are W cm-2 sr-1
LGS | A | 1 | 1 | 5 | 1.206e-06 | 1.2536e-06 |  | 7. Radiances used in fit from SIScal_DNB_20130923.txt
LGS | A | 1 | 1 | 6 | 1.2093e-06 | 1.2581e-06
LGS | A | 1 | 1 | 7 | 1.2109e-06 | 1.2589e-06
LGS | A | 1 | 1 | 8 | 1.2041e-06 | 1.2497e-06
LGS | A | 1 | 1 | 9 | 1.1937e-06 | 1.2292e-06
LGS | A | 1 | 1 | 10 | 1.1985e-06 | 1.2343e-06
LGS | A | 1 | 1 | 11 | 1.2017e-06 | 1.2384e-06
LGS | A | 1 | 1 | 12 | 1.1993e-06 | 1.2357e-06
LGS | A | 1 | 1 | 13 | 1.2067e-06 | 1.2438e-06
LGS | A | 1 | 1 | 14 | 1.2081e-06 | 1.246e-06
LGS | A | 1 | 1 | 15 | 1.2065e-06 | 1.2441e-06
LGS | A | 1 | 1 | 16 | 1.2134e-06 | 1.2532e-06
LGS | A | 1 | 7 | 1 | 1.8531e-06 | 1.9425e-06
LGS | A | 1 | 7 | 2 | 1.8599e-06 | 1.9488e-06
LGS | A | 1 | 7 | 3 | 1.8566e-06 | 1.9446e-06
LGS | A | 1 | 7 | 4 | 1.8503e-06 | 1.939e-06
LGS | A | 1 | 7 | 5 | 1.8506e-06 | 1.9382e-06
LGS | A | 1 | 7 | 6 | 1.8464e-06 | 1.9363e-06
LGS | A | 1 | 7 | 7 | 1.8531e-06 | 1.9404e-06
LGS | A | 1 | 7 | 8 | 1.8437e-06 | 1.9283e-06
LGS | A | 1 | 7 | 9 | 1.8285e-06 | 1.8974e-06
LGS | A | 1 | 7 | 10 | 1.833e-06 | 1.9021e-06
LGS | A | 1 | 7 | 11 | 1.8327e-06 | 1.9025e-06
LGS | A | 1 | 7 | 12 | 1.8418e-06 | 1.9123e-06
LGS | A | 1 | 7 | 13 | 1.8404e-06 | 1.9105e-06
LGS | A | 1 | 7 | 14 | 1.8436e-06 | 1.9153e-06
LGS | A | 1 | 7 | 15 | 1.8473e-06 | 1.9189e-06
LGS | A | 1 | 7 | 16 | 1.8534e-06 | 1.9262e-06
LGS | A | 1 | 11 | 1 | 2.6835e-06 | 2.7913e-06
LGS | A | 1 | 11 | 2 | 2.6893e-06 | 2.7956e-06
LGS | A | 1 | 11 | 3 | 2.6878e-06 | 2.7963e-06
LGS | A | 1 | 11 | 4 | 2.6721e-06 | 2.7788e-06
LGS | A | 1 | 11 | 5 | 2.6745e-06 | 2.7819e-06
LGS | A | 1 | 11 | 6 | 2.6816e-06 | 2.7918e-06
LGS | A | 1 | 11 | 7 | 2.6725e-06 | 2.7765e-06
LGS | A | 1 | 11 | 8 | 2.6711e-06 | 2.7732e-06
LGS | A | 1 | 11 | 9 | 2.6518e-06 | 2.7318e-06
LGS | A | 1 | 11 | 10 | 2.6403e-06 | 2.7199e-06
LGS | A | 1 | 11 | 11 | 2.6593e-06 | 2.7405e-06
LGS | A | 1 | 11 | 12 | 2.6627e-06 | 2.7453e-06
LGS | A | 1 | 11 | 13 | 2.656e-06 | 2.7377e-06
LGS | A | 1 | 11 | 14 | 2.6733e-06 | 2.7563e-06
LGS | A | 1 | 11 | 15 | 2.6689e-06 | 2.754e-06
LGS | A | 1 | 11 | 16 | 2.6674e-06 | 2.7515e-06
LGS | A | 1 | 15 | 1 | 4.1523e-06 | 4.3153e-06
LGS | A | 1 | 15 | 2 | 4.1551e-06 | 4.3208e-06
LGS | A | 1 | 15 | 3 | 4.1533e-06 | 4.3193e-06
LGS | A | 1 | 15 | 4 | 4.1324e-06 | 4.2954e-06
LGS | A | 1 | 15 | 5 | 4.1327e-06 | 4.2994e-06
LGS | A | 1 | 15 | 6 | 4.1317e-06 | 4.2982e-06
LGS | A | 1 | 15 | 7 | 4.1336e-06 | 4.2915e-06
LGS | A | 1 | 15 | 8 | 4.1348e-06 | 4.2917e-06
LGS | A | 1 | 15 | 9 | 4.1075e-06 | 4.2311e-06
LGS | A | 1 | 15 | 10 | 4.083e-06 | 4.2051e-06
LGS | A | 1 | 15 | 11 | 4.0945e-06 | 4.2171e-06
LGS | A | 1 | 15 | 12 | 4.1053e-06 | 4.2315e-06
LGS | A | 1 | 15 | 13 | 4.104e-06 | 4.2312e-06
LGS | A | 1 | 15 | 14 | 4.124e-06 | 4.2499e-06
LGS | A | 1 | 15 | 15 | 4.1307e-06 | 4.2589e-06
LGS | A | 1 | 15 | 16 | 4.1198e-06 | 4.2505e-06
LGS | A | 1 | 20 | 1 | 6.1827e-06 | 6.4282e-06
LGS | A | 1 | 20 | 2 | 6.1688e-06 | 6.4142e-06
LGS | A | 1 | 20 | 3 | 6.1659e-06 | 6.4077e-06
LGS | A | 1 | 20 | 4 | 6.1568e-06 | 6.4011e-06
LGS | A | 1 | 20 | 5 | 6.1557e-06 | 6.4062e-06
LGS | A | 1 | 20 | 6 | 6.1519e-06 | 6.3946e-06
LGS | A | 1 | 20 | 7 | 6.1664e-06 | 6.399e-06
LGS | A | 1 | 20 | 8 | 6.1626e-06 | 6.395e-06
LGS | A | 1 | 20 | 9 | 6.1199e-06 | 6.3019e-06
LGS | A | 1 | 20 | 10 | 6.085e-06 | 6.2636e-06
LGS | A | 1 | 20 | 11 | 6.0725e-06 | 6.2543e-06
LGS | A | 1 | 20 | 12 | 6.1027e-06 | 6.2883e-06
LGS | A | 1 | 20 | 13 | 6.1249e-06 | 6.3121e-06
LGS | A | 1 | 20 | 14 | 6.1272e-06 | 6.3137e-06
LGS | A | 1 | 20 | 15 | 6.1319e-06 | 6.319e-06
LGS | A | 1 | 20 | 16 | 6.1417e-06 | 6.3324e-06
LGS | A | 1 | 26 | 1 | 9.8019e-06 | 1.0182e-05
LGS | A | 1 | 26 | 2 | 9.7954e-06 | 1.0173e-05
LGS | A | 1 | 26 | 3 | 9.7755e-06 | 1.015e-05
LGS | A | 1 | 26 | 4 | 9.7708e-06 | 1.0157e-05
LGS | A | 1 | 26 | 5 | 9.7719e-06 | 1.0163e-05
LGS | A | 1 | 26 | 6 | 9.7641e-06 | 1.0135e-05
LGS | A | 1 | 26 | 7 | 9.7513e-06 | 1.011e-05
LGS | A | 1 | 26 | 8 | 9.7564e-06 | 1.0117e-05
LGS | A | 1 | 26 | 9 | 9.6902e-06 | 9.9811e-06
LGS | A | 1 | 26 | 10 | 9.614e-06 | 9.8976e-06
LGS | A | 1 | 26 | 11 | 9.6517e-06 | 9.94e-06
LGS | A | 1 | 26 | 12 | 9.6682e-06 | 9.962e-06
LGS | A | 1 | 26 | 13 | 9.7022e-06 | 9.9998e-06
LGS | A | 1 | 26 | 14 | 9.7206e-06 | 1.0022e-05
LGS | A | 1 | 26 | 15 | 9.7153e-06 | 1.0012e-05
LGS | A | 1 | 26 | 16 | 9.7391e-06 | 1.0035e-05
LGS | A | 1 | 32 | 1 | 1.5444e-05 | 1.6054e-05
LGS | A | 1 | 32 | 2 | 1.5426e-05 | 1.6e-05
LGS | A | 1 | 32 | 3 | 1.5423e-05 | 1.6008e-05
LGS | A | 1 | 32 | 4 | 1.5478e-05 | 1.6084e-05
LGS | A | 1 | 32 | 5 | 1.5535e-05 | 1.6115e-05
LGS | A | 1 | 32 | 6 | 1.5519e-05 | 1.6106e-05
LGS | A | 1 | 32 | 7 | 1.5455e-05 | 1.6012e-05
LGS | A | 1 | 32 | 8 | 1.5429e-05 | 1.5994e-05
LGS | A | 1 | 32 | 9 | 1.5414e-05 | 1.586e-05
LGS | A | 1 | 32 | 10 | 1.5317e-05 | 1.5767e-05
LGS | A | 1 | 32 | 11 | 1.5394e-05 | 1.5848e-05
LGS | A | 1 | 32 | 12 | 1.5451e-05 | 1.5874e-05
LGS | A | 1 | 32 | 13 | 1.5387e-05 | 1.5856e-05
LGS | A | 1 | 32 | 14 | 1.5336e-05 | 1.579e-05
LGS | A | 1 | 32 | 15 | 1.5358e-05 | 1.582e-05
LGS | A | 1 | 32 | 16 | 1.5332e-05 | 1.5818e-05
LGS | A | 1 | 34 | 1 | 6.2578e-07 | 6.5199e-07
LGS | A | 1 | 34 | 2 | 6.2711e-07 | 6.5171e-07
LGS | A | 1 | 34 | 3 | 6.3202e-07 | 6.5682e-07
LGS | A | 1 | 34 | 4 | 6.2987e-07 | 6.5451e-07
LGS | A | 1 | 34 | 5 | 6.2634e-07 | 6.5085e-07
LGS | A | 1 | 34 | 6 | 6.2808e-07 | 6.5314e-07
LGS | A | 1 | 34 | 7 | 6.2888e-07 | 6.5364e-07
LGS | A | 1 | 34 | 8 | 6.2533e-07 | 6.4882e-07
LGS | A | 1 | 34 | 9 | 6.2003e-07 | 6.3831e-07
LGS | A | 1 | 34 | 10 | 6.2257e-07 | 6.4086e-07
LGS | A | 1 | 34 | 11 | 6.2417e-07 | 6.431e-07
LGS | A | 1 | 34 | 12 | 6.2287e-07 | 6.4168e-07
LGS | A | 1 | 34 | 13 | 6.2673e-07 | 6.4588e-07
LGS | A | 1 | 34 | 14 | 6.2762e-07 | 6.4707e-07
LGS | A | 1 | 34 | 15 | 6.2694e-07 | 6.4634e-07
LGS | A | 1 | 34 | 16 | 6.3059e-07 | 6.5108e-07
LGS | B | 1 | 1 | 1 | 1.2022e-06 | 1.2543e-06
LGS | B | 1 | 1 | 2 | 1.2051e-06 | 1.2536e-06
LGS | B | 1 | 1 | 3 | 1.2146e-06 | 1.2637e-06
LGS | B | 1 | 1 | 4 | 1.2106e-06 | 1.2593e-06
LGS | B | 1 | 1 | 5 | 1.2036e-06 | 1.2524e-06
LGS | B | 1 | 1 | 6 | 1.2068e-06 | 1.257e-06
LGS | B | 1 | 1 | 7 | 1.2086e-06 | 1.2575e-06
LGS | B | 1 | 1 | 8 | 1.2019e-06 | 1.2482e-06
LGS | B | 1 | 1 | 9 | 1.1914e-06 | 1.2278e-06
LGS | B | 1 | 1 | 10 | 1.1963e-06 | 1.2329e-06
LGS | B | 1 | 1 | 11 | 1.1994e-06 | 1.237e-06
LGS | B | 1 | 1 | 12 | 1.197e-06 | 1.2342e-06
LGS | B | 1 | 1 | 13 | 1.2042e-06 | 1.2422e-06
LGS | B | 1 | 1 | 14 | 1.2058e-06 | 1.2442e-06
LGS | B | 1 | 1 | 15 | 1.2041e-06 | 1.2425e-06
LGS | B | 1 | 1 | 16 | 1.2111e-06 | 1.2513e-06
LGS | B | 1 | 7 | 1 | 1.8499e-06 | 1.9403e-06
LGS | B | 1 | 7 | 2 | 1.8565e-06 | 1.9468e-06
LGS | B | 1 | 7 | 3 | 1.8532e-06 | 1.9426e-06
LGS | B | 1 | 7 | 4 | 1.8465e-06 | 1.9372e-06
LGS | B | 1 | 7 | 5 | 1.8471e-06 | 1.9361e-06
LGS | B | 1 | 7 | 6 | 1.8428e-06 | 1.9347e-06
LGS | B | 1 | 7 | 7 | 1.8498e-06 | 1.9383e-06
LGS | B | 1 | 7 | 8 | 1.8403e-06 | 1.9257e-06
LGS | B | 1 | 7 | 9 | 1.825e-06 | 1.8953e-06
LGS | B | 1 | 7 | 10 | 1.8293e-06 | 1.8997e-06
LGS | B | 1 | 7 | 11 | 1.8295e-06 | 1.9007e-06
LGS | B | 1 | 7 | 12 | 1.8383e-06 | 1.9103e-06
LGS | B | 1 | 7 | 13 | 1.8368e-06 | 1.9084e-06
LGS | B | 1 | 7 | 14 | 1.8402e-06 | 1.9128e-06
LGS | B | 1 | 7 | 15 | 1.8439e-06 | 1.9166e-06
LGS | B | 1 | 7 | 16 | 1.8499e-06 | 1.9237e-06
LGS | B | 1 | 11 | 1 | 2.6787e-06 | 2.7876e-06
LGS | B | 1 | 11 | 2 | 2.6842e-06 | 2.7926e-06
LGS | B | 1 | 11 | 3 | 2.6828e-06 | 2.7933e-06
LGS | B | 1 | 11 | 4 | 2.6674e-06 | 2.776e-06
LGS | B | 1 | 11 | 5 | 2.6698e-06 | 2.779e-06
LGS | B | 1 | 11 | 6 | 2.6765e-06 | 2.7892e-06
LGS | B | 1 | 11 | 7 | 2.6677e-06 | 2.7734e-06
LGS | B | 1 | 11 | 8 | 2.6663e-06 | 2.7701e-06
LGS | B | 1 | 11 | 9 | 2.6468e-06 | 2.7291e-06
LGS | B | 1 | 11 | 10 | 2.6355e-06 | 2.7171e-06
LGS | B | 1 | 11 | 11 | 2.6542e-06 | 2.7371e-06
LGS | B | 1 | 11 | 12 | 2.6576e-06 | 2.7424e-06
LGS | B | 1 | 11 | 13 | 2.6508e-06 | 2.7346e-06
LGS | B | 1 | 11 | 14 | 2.6683e-06 | 2.7529e-06
LGS | B | 1 | 11 | 15 | 2.6638e-06 | 2.7499e-06
LGS | B | 1 | 11 | 16 | 2.6625e-06 | 2.7476e-06
LGS | B | 1 | 15 | 1 | 4.1444e-06 | 4.3109e-06
LGS | B | 1 | 15 | 2 | 4.1476e-06 | 4.3165e-06
LGS | B | 1 | 15 | 3 | 4.145e-06 | 4.3152e-06
LGS | B | 1 | 15 | 4 | 4.1253e-06 | 4.2911e-06
LGS | B | 1 | 15 | 5 | 4.1251e-06 | 4.2964e-06
LGS | B | 1 | 15 | 6 | 4.1241e-06 | 4.294e-06
LGS | B | 1 | 15 | 7 | 4.1261e-06 | 4.2864e-06
LGS | B | 1 | 15 | 8 | 4.1274e-06 | 4.2864e-06
LGS | B | 1 | 15 | 9 | 4.0997e-06 | 4.2259e-06
LGS | B | 1 | 15 | 10 | 4.0753e-06 | 4.2008e-06
LGS | B | 1 | 15 | 11 | 4.0872e-06 | 4.2128e-06
LGS | B | 1 | 15 | 12 | 4.0974e-06 | 4.2265e-06
LGS | B | 1 | 15 | 13 | 4.0962e-06 | 4.2255e-06
LGS | B | 1 | 15 | 14 | 4.1163e-06 | 4.2453e-06
LGS | B | 1 | 15 | 15 | 4.1229e-06 | 4.2538e-06
LGS | B | 1 | 15 | 16 | 4.1121e-06 | 4.2456e-06
LGS | B | 1 | 20 | 1 | 6.1712e-06 | 6.4209e-06
LGS | B | 1 | 20 | 2 | 6.1575e-06 | 6.4084e-06
LGS | B | 1 | 20 | 3 | 6.1541e-06 | 6.4011e-06
LGS | B | 1 | 20 | 4 | 6.1461e-06 | 6.3942e-06
LGS | B | 1 | 20 | 5 | 6.144e-06 | 6.4012e-06
LGS | B | 1 | 20 | 6 | 6.14e-06 | 6.3887e-06
LGS | B | 1 | 20 | 7 | 6.1558e-06 | 6.3906e-06
LGS | B | 1 | 20 | 8 | 6.153e-06 | 6.3877e-06
LGS | B | 1 | 20 | 9 | 6.1086e-06 | 6.2963e-06
LGS | B | 1 | 20 | 10 | 6.0738e-06 | 6.2565e-06
LGS | B | 1 | 20 | 11 | 6.061e-06 | 6.2469e-06
LGS | B | 1 | 20 | 12 | 6.0934e-06 | 6.282e-06
LGS | B | 1 | 20 | 13 | 6.1139e-06 | 6.3057e-06
LGS | B | 1 | 20 | 14 | 6.1158e-06 | 6.3068e-06
LGS | B | 1 | 20 | 15 | 6.1214e-06 | 6.3107e-06
LGS | B | 1 | 20 | 16 | 6.1304e-06 | 6.3256e-06
LGS | B | 1 | 26 | 1 | 9.7828e-06 | 1.0173e-05
LGS | B | 1 | 26 | 2 | 9.7779e-06 | 1.0162e-05
LGS | B | 1 | 26 | 3 | 9.7569e-06 | 1.014e-05
LGS | B | 1 | 26 | 4 | 9.7535e-06 | 1.0148e-05
LGS | B | 1 | 26 | 5 | 9.7558e-06 | 1.0153e-05
LGS | B | 1 | 26 | 6 | 9.7473e-06 | 1.0126e-05
LGS | B | 1 | 26 | 7 | 9.7317e-06 | 1.0098e-05
LGS | B | 1 | 26 | 8 | 9.7405e-06 | 1.0106e-05
LGS | B | 1 | 26 | 9 | 9.6727e-06 | 9.9704e-06
LGS | B | 1 | 26 | 10 | 9.5957e-06 | 9.8868e-06
LGS | B | 1 | 26 | 11 | 9.6342e-06 | 9.9275e-06
LGS | B | 1 | 26 | 12 | 9.6506e-06 | 9.9499e-06
LGS | B | 1 | 26 | 13 | 9.683e-06 | 9.9872e-06
LGS | B | 1 | 26 | 14 | 9.7021e-06 | 1.0009e-05
LGS | B | 1 | 26 | 15 | 9.6976e-06 | 1e-05
LGS | B | 1 | 26 | 16 | 9.7213e-06 | 1.0022e-05
LGS | B | 1 | 32 | 1 | 1.5418e-05 | 1.6035e-05
LGS | B | 1 | 32 | 2 | 1.5396e-05 | 1.5985e-05
LGS | B | 1 | 32 | 3 | 1.5394e-05 | 1.5999e-05
LGS | B | 1 | 32 | 4 | 1.5449e-05 | 1.6072e-05
LGS | B | 1 | 32 | 5 | 1.5505e-05 | 1.6093e-05
LGS | B | 1 | 32 | 6 | 1.5488e-05 | 1.6092e-05
LGS | B | 1 | 32 | 7 | 1.5425e-05 | 1.599e-05
LGS | B | 1 | 32 | 8 | 1.5402e-05 | 1.5978e-05
LGS | B | 1 | 32 | 9 | 1.5382e-05 | 1.5847e-05
LGS | B | 1 | 32 | 10 | 1.5286e-05 | 1.5745e-05
LGS | B | 1 | 32 | 11 | 1.5361e-05 | 1.5826e-05
LGS | B | 1 | 32 | 12 | 1.542e-05 | 1.5856e-05
LGS | B | 1 | 32 | 13 | 1.5359e-05 | 1.5839e-05
LGS | B | 1 | 32 | 14 | 1.5308e-05 | 1.5775e-05
LGS | B | 1 | 32 | 15 | 1.5333e-05 | 1.5806e-05
LGS | B | 1 | 32 | 16 | 1.5302e-05 | 1.5798e-05
LGS | B | 1 | 34 | 1 | 6.247e-07 | 6.5138e-07
LGS | B | 1 | 34 | 2 | 6.26e-07 | 6.5106e-07
LGS | B | 1 | 34 | 3 | 6.3082e-07 | 6.5612e-07
LGS | B | 1 | 34 | 4 | 6.2876e-07 | 6.5375e-07
LGS | B | 1 | 34 | 5 | 6.2513e-07 | 6.5014e-07
LGS | B | 1 | 34 | 6 | 6.2687e-07 | 6.5255e-07
LGS | B | 1 | 34 | 7 | 6.2783e-07 | 6.5295e-07
LGS | B | 1 | 34 | 8 | 6.2417e-07 | 6.4808e-07
LGS | B | 1 | 34 | 9 | 6.1879e-07 | 6.3754e-07
LGS | B | 1 | 34 | 10 | 6.2138e-07 | 6.4016e-07
LGS | B | 1 | 34 | 11 | 6.2303e-07 | 6.4242e-07
LGS | B | 1 | 34 | 12 | 6.2178e-07 | 6.4093e-07
LGS | B | 1 | 34 | 13 | 6.2551e-07 | 6.4508e-07
LGS | B | 1 | 34 | 14 | 6.2654e-07 | 6.462e-07
LGS | B | 1 | 34 | 15 | 6.2576e-07 | 6.4539e-07
LGS | B | 1 | 34 | 16 | 6.2934e-07 | 6.5022e-07
MGS | A | 2 | 1 | 1 | 3.6565e-09 | 3.6964e-09
MGS | A | 2 | 1 | 2 | 3.6856e-09 | 3.7245e-09
MGS | A | 2 | 1 | 3 | 3.6888e-09 | 3.7276e-09
MGS | A | 2 | 1 | 4 | 3.6906e-09 | 3.7297e-09
MGS | A | 2 | 1 | 5 | 3.69e-09 | 3.729e-09
MGS | A | 2 | 1 | 6 | 3.6826e-09 | 3.7219e-09
MGS | A | 2 | 1 | 7 | 3.6536e-09 | 3.6917e-09
MGS | A | 2 | 1 | 8 | 3.65e-09 | 3.6878e-09
MGS | A | 2 | 1 | 9 | 3.4388e-09 | 3.4971e-09
MGS | A | 2 | 1 | 10 | 3.4341e-09 | 3.4922e-09
MGS | A | 2 | 1 | 11 | 3.4495e-09 | 3.5082e-09
MGS | A | 2 | 1 | 12 | 3.4554e-09 | 3.5144e-09
MGS | A | 2 | 1 | 13 | 3.4559e-09 | 3.5149e-09
MGS | A | 2 | 1 | 14 | 3.4523e-09 | 3.5111e-09
MGS | A | 2 | 1 | 15 | 3.4491e-09 | 3.5085e-09
MGS | A | 2 | 1 | 16 | 3.4439e-09 | 3.5035e-09
MGS | A | 2 | 7 | 1 | 5.627e-09 | 5.7362e-09
MGS | A | 2 | 7 | 2 | 5.6284e-09 | 5.7364e-09
MGS | A | 2 | 7 | 3 | 5.639e-09 | 5.7479e-09
MGS | A | 2 | 7 | 4 | 5.6306e-09 | 5.7388e-09
MGS | A | 2 | 7 | 5 | 5.6322e-09 | 5.7403e-09
MGS | A | 2 | 7 | 6 | 5.6097e-09 | 5.7168e-09
MGS | A | 2 | 7 | 7 | 5.5786e-09 | 5.6848e-09
MGS | A | 2 | 7 | 8 | 5.5709e-09 | 5.6771e-09
MGS | A | 2 | 7 | 9 | 5.2494e-09 | 5.3839e-09
MGS | A | 2 | 7 | 10 | 5.2499e-09 | 5.3829e-09
MGS | A | 2 | 7 | 11 | 5.2604e-09 | 5.3929e-09
MGS | A | 2 | 7 | 12 | 5.2769e-09 | 5.4113e-09
MGS | A | 2 | 7 | 13 | 5.2783e-09 | 5.4124e-09
MGS | A | 2 | 7 | 14 | 5.2812e-09 | 5.4153e-09
MGS | A | 2 | 7 | 15 | 5.2709e-09 | 5.4043e-09
MGS | A | 2 | 7 | 16 | 5.2732e-09 | 5.4072e-09
MGS | A | 2 | 11 | 1 | 8.2629e-09 | 8.4665e-09
MGS | A | 2 | 11 | 2 | 8.251e-09 | 8.4537e-09
MGS | A | 2 | 11 | 3 | 8.2635e-09 | 8.466e-09
MGS | A | 2 | 11 | 4 | 8.2644e-09 | 8.4658e-09
MGS | A | 2 | 11 | 5 | 8.2357e-09 | 8.4381e-09
MGS | A | 2 | 11 | 6 | 8.2033e-09 | 8.403e-09
MGS | A | 2 | 11 | 7 | 8.1775e-09 | 8.3767e-09
MGS | A | 2 | 11 | 8 | 8.1574e-09 | 8.3558e-09
MGS | A | 2 | 11 | 9 | 7.6851e-09 | 7.9231e-09
MGS | A | 2 | 11 | 10 | 7.6943e-09 | 7.9312e-09
MGS | A | 2 | 11 | 11 | 7.6925e-09 | 7.9275e-09
MGS | A | 2 | 11 | 12 | 7.7083e-09 | 7.9463e-09
MGS | A | 2 | 11 | 13 | 7.7319e-09 | 7.9715e-09
MGS | A | 2 | 11 | 14 | 7.7325e-09 | 7.9723e-09
MGS | A | 2 | 11 | 15 | 7.713e-09 | 7.9523e-09
MGS | A | 2 | 11 | 16 | 7.7247e-09 | 7.9644e-09
MGS | A | 2 | 15 | 1 | 1.2803e-08 | 1.3113e-08
MGS | A | 2 | 15 | 2 | 1.2783e-08 | 1.3091e-08
MGS | A | 2 | 15 | 3 | 1.2765e-08 | 1.3072e-08
MGS | A | 2 | 15 | 4 | 1.2781e-08 | 1.3088e-08
MGS | A | 2 | 15 | 5 | 1.2754e-08 | 1.306e-08
MGS | A | 2 | 15 | 6 | 1.2699e-08 | 1.3002e-08
MGS | A | 2 | 15 | 7 | 1.2646e-08 | 1.2946e-08
MGS | A | 2 | 15 | 8 | 1.2624e-08 | 1.2924e-08
MGS | A | 2 | 15 | 9 | 1.1894e-08 | 1.2257e-08
MGS | A | 2 | 15 | 10 | 1.1906e-08 | 1.2268e-08
MGS | A | 2 | 15 | 11 | 1.1921e-08 | 1.2281e-08
MGS | A | 2 | 15 | 12 | 1.1943e-08 | 1.2306e-08
MGS | A | 2 | 15 | 13 | 1.1959e-08 | 1.2324e-08
MGS | A | 2 | 15 | 14 | 1.1947e-08 | 1.2312e-08
MGS | A | 2 | 15 | 15 | 1.196e-08 | 1.2325e-08
MGS | A | 2 | 15 | 16 | 1.197e-08 | 1.2336e-08
MGS | A | 2 | 20 | 1 | 1.9112e-08 | 1.9569e-08
MGS | A | 2 | 20 | 2 | 1.9079e-08 | 1.9536e-08
MGS | A | 2 | 20 | 3 | 1.9069e-08 | 1.9524e-08
MGS | A | 2 | 20 | 4 | 1.9048e-08 | 1.9506e-08
MGS | A | 2 | 20 | 5 | 1.8993e-08 | 1.9447e-08
MGS | A | 2 | 20 | 6 | 1.8888e-08 | 1.9337e-08
MGS | A | 2 | 20 | 7 | 1.8827e-08 | 1.9277e-08
MGS | A | 2 | 20 | 8 | 1.8869e-08 | 1.9318e-08
MGS | A | 2 | 20 | 9 | 1.7782e-08 | 1.8321e-08
MGS | A | 2 | 20 | 10 | 1.7734e-08 | 1.8278e-08
MGS | A | 2 | 20 | 11 | 1.775e-08 | 1.8287e-08
MGS | A | 2 | 20 | 12 | 1.7797e-08 | 1.8334e-08
MGS | A | 2 | 20 | 13 | 1.7831e-08 | 1.8374e-08
MGS | A | 2 | 20 | 14 | 1.7849e-08 | 1.8391e-08
MGS | A | 2 | 20 | 15 | 1.786e-08 | 1.8405e-08
MGS | A | 2 | 20 | 16 | 1.7872e-08 | 1.842e-08
MGS | A | 2 | 26 | 1 | 3.0318e-08 | 3.1284e-08
MGS | A | 2 | 26 | 2 | 3.0283e-08 | 3.1241e-08
MGS | A | 2 | 26 | 3 | 3.0253e-08 | 3.1212e-08
MGS | A | 2 | 26 | 4 | 3.0185e-08 | 3.1146e-08
MGS | A | 2 | 26 | 5 | 3.0107e-08 | 3.1053e-08
MGS | A | 2 | 26 | 6 | 2.9969e-08 | 3.0912e-08
MGS | A | 2 | 26 | 7 | 2.9884e-08 | 3.0821e-08
MGS | A | 2 | 26 | 8 | 2.9946e-08 | 3.0889e-08
MGS | A | 2 | 26 | 9 | 2.8258e-08 | 2.9269e-08
MGS | A | 2 | 26 | 10 | 2.8203e-08 | 2.9214e-08
MGS | A | 2 | 26 | 11 | 2.8219e-08 | 2.921e-08
MGS | A | 2 | 26 | 12 | 2.827e-08 | 2.9252e-08
MGS | A | 2 | 26 | 13 | 2.8289e-08 | 2.927e-08
MGS | A | 2 | 26 | 14 | 2.8325e-08 | 2.9305e-08
MGS | A | 2 | 26 | 15 | 2.8338e-08 | 2.9316e-08
MGS | A | 2 | 26 | 16 | 2.8277e-08 | 2.9355e-08
MGS | A | 2 | 32 | 1 | 4.7729e-08 | 4.9233e-08
MGS | A | 2 | 32 | 2 | 4.7634e-08 | 4.914e-08
MGS | A | 2 | 32 | 3 | 4.775e-08 | 4.9263e-08
MGS | A | 2 | 32 | 4 | 4.7498e-08 | 4.8997e-08
MGS | A | 2 | 32 | 5 | 4.7249e-08 | 4.8731e-08
MGS | A | 2 | 32 | 6 | 4.7093e-08 | 4.8568e-08
MGS | A | 2 | 32 | 7 | 4.6998e-08 | 4.8467e-08
MGS | A | 2 | 32 | 8 | 4.7219e-08 | 4.8709e-08
MGS | A | 2 | 32 | 9 | 4.4448e-08 | 4.6144e-08
MGS | A | 2 | 32 | 10 | 4.4277e-08 | 4.5974e-08
MGS | A | 2 | 32 | 11 | 4.4281e-08 | 4.5961e-08
MGS | A | 2 | 32 | 12 | 4.4306e-08 | 4.5976e-08
MGS | A | 2 | 32 | 13 | 4.4426e-08 | 4.6103e-08
MGS | A | 2 | 32 | 14 | 4.4598e-08 | 4.6285e-08
MGS | A | 2 | 32 | 15 | 4.4456e-08 | 4.6142e-08
MGS | A | 2 | 32 | 16 | 4.4502e-08 | 4.6197e-08
MGS | A | 2 | 34 | 1 | 1.9093e-09 | 1.93e-09
MGS | A | 2 | 34 | 2 | 1.9255e-09 | 1.9455e-09
MGS | A | 2 | 34 | 3 | 1.9262e-09 | 1.9459e-09
MGS | A | 2 | 34 | 4 | 1.9269e-09 | 1.9471e-09
MGS | A | 2 | 34 | 5 | 1.9267e-09 | 1.9469e-09
MGS | A | 2 | 34 | 6 | 1.9231e-09 | 1.9428e-09
MGS | A | 2 | 34 | 7 | 1.907e-09 | 1.9264e-09
MGS | A | 2 | 34 | 8 | 1.9053e-09 | 1.9248e-09
MGS | A | 2 | 34 | 9 | 1.7947e-09 | 1.8252e-09
MGS | A | 2 | 34 | 10 | 1.7923e-09 | 1.8212e-09
MGS | A | 2 | 34 | 11 | 1.7996e-09 | 1.8306e-09
MGS | A | 2 | 34 | 12 | 1.8034e-09 | 1.8341e-09
MGS | A | 2 | 34 | 13 | 1.803e-09 | 1.8335e-09
MGS | A | 2 | 34 | 14 | 1.8003e-09 | 1.8324e-09
MGS | A | 2 | 34 | 15 | 1.7991e-09 | 1.8305e-09
MGS | A | 2 | 34 | 16 | 1.7962e-09 | 1.827e-09
MGS | B | 2 | 1 | 1 | 3.6501e-09 | 3.6898e-09
MGS | B | 2 | 1 | 2 | 3.6785e-09 | 3.7181e-09
MGS | B | 2 | 1 | 3 | 3.6811e-09 | 3.7201e-09
MGS | B | 2 | 1 | 4 | 3.6832e-09 | 3.7229e-09
MGS | B | 2 | 1 | 5 | 3.6831e-09 | 3.7218e-09
MGS | B | 2 | 1 | 6 | 3.6756e-09 | 3.715e-09
MGS | B | 2 | 1 | 7 | 3.6463e-09 | 3.6845e-09
MGS | B | 2 | 1 | 8 | 3.6433e-09 | 3.6813e-09
MGS | B | 2 | 1 | 9 | 3.4321e-09 | 3.4911e-09
MGS | B | 2 | 1 | 10 | 3.4278e-09 | 3.486e-09
MGS | B | 2 | 1 | 11 | 3.4429e-09 | 3.5014e-09
MGS | B | 2 | 1 | 12 | 3.4485e-09 | 3.5078e-09
MGS | B | 2 | 1 | 13 | 3.449e-09 | 3.5083e-09
MGS | B | 2 | 1 | 14 | 3.4453e-09 | 3.505e-09
MGS | B | 2 | 1 | 15 | 3.4421e-09 | 3.5016e-09
MGS | B | 2 | 1 | 16 | 3.4366e-09 | 3.4971e-09
MGS | B | 2 | 7 | 1 | 5.6156e-09 | 5.7256e-09
MGS | B | 2 | 7 | 2 | 5.6162e-09 | 5.7254e-09
MGS | B | 2 | 7 | 3 | 5.6275e-09 | 5.7363e-09
MGS | B | 2 | 7 | 4 | 5.6194e-09 | 5.728e-09
MGS | B | 2 | 7 | 5 | 5.6207e-09 | 5.7292e-09
MGS | B | 2 | 7 | 6 | 5.5983e-09 | 5.7056e-09
MGS | B | 2 | 7 | 7 | 5.567e-09 | 5.6739e-09
MGS | B | 2 | 7 | 8 | 5.5603e-09 | 5.6663e-09
MGS | B | 2 | 7 | 9 | 5.2387e-09 | 5.3734e-09
MGS | B | 2 | 7 | 10 | 5.239e-09 | 5.3729e-09
MGS | B | 2 | 7 | 11 | 5.2498e-09 | 5.3827e-09
MGS | B | 2 | 7 | 12 | 5.2654e-09 | 5.4005e-09
MGS | B | 2 | 7 | 13 | 5.2671e-09 | 5.4022e-09
MGS | B | 2 | 7 | 14 | 5.2699e-09 | 5.4053e-09
MGS | B | 2 | 7 | 15 | 5.26e-09 | 5.3945e-09
MGS | B | 2 | 7 | 16 | 5.2622e-09 | 5.3965e-09
MGS | B | 2 | 11 | 1 | 8.2462e-09 | 8.4479e-09
MGS | B | 2 | 11 | 2 | 8.2335e-09 | 8.4365e-09
MGS | B | 2 | 11 | 3 | 8.2462e-09 | 8.4489e-09
MGS | B | 2 | 11 | 4 | 8.2465e-09 | 8.4495e-09
MGS | B | 2 | 11 | 5 | 8.2184e-09 | 8.4209e-09
MGS | B | 2 | 11 | 6 | 8.1861e-09 | 8.3858e-09
MGS | B | 2 | 11 | 7 | 8.1602e-09 | 8.3592e-09
MGS | B | 2 | 11 | 8 | 8.1404e-09 | 8.3391e-09
MGS | B | 2 | 11 | 9 | 7.669e-09 | 7.907e-09
MGS | B | 2 | 11 | 10 | 7.6778e-09 | 7.9145e-09
MGS | B | 2 | 11 | 11 | 7.6756e-09 | 7.9123e-09
MGS | B | 2 | 11 | 12 | 7.6927e-09 | 7.9309e-09
MGS | B | 2 | 11 | 13 | 7.715e-09 | 7.9549e-09
MGS | B | 2 | 11 | 14 | 7.7155e-09 | 7.9553e-09
MGS | B | 2 | 11 | 15 | 7.6962e-09 | 7.9361e-09
MGS | B | 2 | 11 | 16 | 7.7076e-09 | 7.9479e-09
MGS | B | 2 | 15 | 1 | 1.2777e-08 | 1.3087e-08
MGS | B | 2 | 15 | 2 | 1.2756e-08 | 1.3063e-08
MGS | B | 2 | 15 | 3 | 1.274e-08 | 1.3046e-08
MGS | B | 2 | 15 | 4 | 1.2753e-08 | 1.3062e-08
MGS | B | 2 | 15 | 5 | 1.2726e-08 | 1.3035e-08
MGS | B | 2 | 15 | 6 | 1.2674e-08 | 1.2977e-08
MGS | B | 2 | 15 | 7 | 1.2619e-08 | 1.2921e-08
MGS | B | 2 | 15 | 8 | 1.2597e-08 | 1.2899e-08
MGS | B | 2 | 15 | 9 | 1.187e-08 | 1.2232e-08
MGS | B | 2 | 15 | 10 | 1.188e-08 | 1.2245e-08
MGS | B | 2 | 15 | 11 | 1.1896e-08 | 1.2256e-08
MGS | B | 2 | 15 | 12 | 1.1917e-08 | 1.228e-08
MGS | B | 2 | 15 | 13 | 1.1933e-08 | 1.23e-08
MGS | B | 2 | 15 | 14 | 1.1921e-08 | 1.2287e-08
MGS | B | 2 | 15 | 15 | 1.1934e-08 | 1.23e-08
MGS | B | 2 | 15 | 16 | 1.1945e-08 | 1.2312e-08
MGS | B | 2 | 20 | 1 | 1.9071e-08 | 1.9533e-08
MGS | B | 2 | 20 | 2 | 1.9035e-08 | 1.9498e-08
MGS | B | 2 | 20 | 3 | 1.9028e-08 | 1.9485e-08
MGS | B | 2 | 20 | 4 | 1.9005e-08 | 1.9464e-08
MGS | B | 2 | 20 | 5 | 1.8952e-08 | 1.9407e-08
MGS | B | 2 | 20 | 6 | 1.8848e-08 | 1.93e-08
MGS | B | 2 | 20 | 7 | 1.8786e-08 | 1.9238e-08
MGS | B | 2 | 20 | 8 | 1.8827e-08 | 1.9279e-08
MGS | B | 2 | 20 | 9 | 1.7742e-08 | 1.8286e-08
MGS | B | 2 | 20 | 10 | 1.7697e-08 | 1.824e-08
MGS | B | 2 | 20 | 11 | 1.7714e-08 | 1.8251e-08
MGS | B | 2 | 20 | 12 | 1.7758e-08 | 1.8298e-08
MGS | B | 2 | 20 | 13 | 1.7792e-08 | 1.8335e-08
MGS | B | 2 | 20 | 14 | 1.7808e-08 | 1.8355e-08
MGS | B | 2 | 20 | 15 | 1.7821e-08 | 1.8367e-08
MGS | B | 2 | 20 | 16 | 1.7833e-08 | 1.8381e-08
MGS | B | 2 | 26 | 1 | 3.0252e-08 | 3.1221e-08
MGS | B | 2 | 26 | 2 | 3.0217e-08 | 3.1181e-08
MGS | B | 2 | 26 | 3 | 3.0186e-08 | 3.115e-08
MGS | B | 2 | 26 | 4 | 3.012e-08 | 3.1084e-08
MGS | B | 2 | 26 | 5 | 3.004e-08 | 3.0993e-08
MGS | B | 2 | 26 | 6 | 2.9904e-08 | 3.0852e-08
MGS | B | 2 | 26 | 7 | 2.9819e-08 | 3.0759e-08
MGS | B | 2 | 26 | 8 | 2.9881e-08 | 3.0826e-08
MGS | B | 2 | 26 | 9 | 2.8197e-08 | 2.9211e-08
MGS | B | 2 | 26 | 10 | 2.8143e-08 | 2.9154e-08
MGS | B | 2 | 26 | 11 | 2.8159e-08 | 2.9152e-08
MGS | B | 2 | 26 | 12 | 2.821e-08 | 2.9193e-08
MGS | B | 2 | 26 | 13 | 2.8227e-08 | 2.9208e-08
MGS | B | 2 | 26 | 14 | 2.8267e-08 | 2.9247e-08
MGS | B | 2 | 26 | 15 | 2.8279e-08 | 2.9255e-08
MGS | B | 2 | 26 | 16 | 2.8323e-08 | 2.9295e-08
MGS | B | 2 | 32 | 1 | 4.7621e-08 | 4.9135e-08
MGS | B | 2 | 32 | 2 | 4.7529e-08 | 4.904e-08
MGS | B | 2 | 32 | 3 | 4.7641e-08 | 4.9165e-08
MGS | B | 2 | 32 | 4 | 4.7393e-08 | 4.8906e-08
MGS | B | 2 | 32 | 5 | 4.7147e-08 | 4.8629e-08
MGS | B | 2 | 32 | 6 | 4.6985e-08 | 4.8469e-08
MGS | B | 2 | 32 | 7 | 4.6893e-08 | 4.8366e-08
MGS | B | 2 | 32 | 8 | 4.7116e-08 | 4.8607e-08
MGS | B | 2 | 32 | 9 | 4.4349e-08 | 4.6046e-08
MGS | B | 2 | 32 | 10 | 4.4177e-08 | 4.5879e-08
MGS | B | 2 | 32 | 11 | 4.4183e-08 | 4.5869e-08
MGS | B | 2 | 32 | 12 | 4.421e-08 | 4.5881e-08
MGS | B | 2 | 32 | 13 | 4.4325e-08 | 4.6013e-08
MGS | B | 2 | 32 | 14 | 4.4501e-08 | 4.619e-08
MGS | B | 2 | 32 | 15 | 4.4356e-08 | 4.6049e-08
MGS | B | 2 | 32 | 16 | 4.4401e-08 | 4.61e-08
MGS | B | 2 | 34 | 1 | 1.9065e-09 | 1.9267e-09
MGS | B | 2 | 34 | 2 | 1.9215e-09 | 1.9424e-09
MGS | B | 2 | 34 | 3 | 1.9231e-09 | 1.943e-09
MGS | B | 2 | 34 | 4 | 1.9238e-09 | 1.9444e-09
MGS | B | 2 | 34 | 5 | 1.9233e-09 | 1.9431e-09
MGS | B | 2 | 34 | 6 | 1.9195e-09 | 1.9402e-09
MGS | B | 2 | 34 | 7 | 1.9037e-09 | 1.9231e-09
MGS | B | 2 | 34 | 8 | 1.9017e-09 | 1.922e-09
MGS | B | 2 | 34 | 9 | 1.7917e-09 | 1.8218e-09
MGS | B | 2 | 34 | 10 | 1.7889e-09 | 1.8188e-09
MGS | B | 2 | 34 | 11 | 1.7969e-09 | 1.8273e-09
MGS | B | 2 | 34 | 12 | 1.7996e-09 | 1.831e-09
MGS | B | 2 | 34 | 13 | 1.8001e-09 | 1.83e-09
MGS | B | 2 | 34 | 14 | 1.7974e-09 | 1.8286e-09
MGS | B | 2 | 34 | 15 | 1.7961e-09 | 1.8269e-09
MGS | B | 2 | 34 | 16 | 1.7929e-09 | 1.8242e-09
HGA | A | 3 | 1 | 1 | 1.4354e-11 | 1.4624e-11
HGA | A | 3 | 1 | 2 | 1.4439e-11 | 1.4702e-11
HGA | A | 3 | 1 | 3 | 1.4294e-11 | 1.4555e-11
HGA | A | 3 | 1 | 4 | 1.4221e-11 | 1.4486e-11
HGA | A | 3 | 1 | 5 | 1.4263e-11 | 1.4534e-11
HGA | A | 3 | 1 | 6 | 1.4194e-11 | 1.4457e-11
HGA | A | 3 | 1 | 7 | 1.4143e-11 | 1.4415e-11
HGA | A | 3 | 1 | 8 | 1.4201e-11 | 1.4467e-11
HGA | A | 3 | 1 | 9 | 1.4232e-11 | 1.4273e-11
HGA | A | 3 | 1 | 10 | 1.4133e-11 | 1.4171e-11
HGA | A | 3 | 1 | 11 | 1.4178e-11 | 1.4223e-11
HGA | A | 3 | 1 | 12 | 1.4208e-11 | 1.4248e-11
HGA | A | 3 | 1 | 13 | 1.4134e-11 | 1.417e-11
HGA | A | 3 | 1 | 14 | 1.4189e-11 | 1.4235e-11
HGA | A | 3 | 1 | 15 | 1.4311e-11 | 1.4343e-11
HGA | A | 3 | 1 | 16 | 1.424e-11 | 1.429e-11
HGA | A | 3 | 7 | 1 | 2.2226e-11 | 2.2638e-11
HGA | A | 3 | 7 | 2 | 2.2027e-11 | 2.2434e-11
HGA | A | 3 | 7 | 3 | 2.2034e-11 | 2.2441e-11
HGA | A | 3 | 7 | 4 | 2.1912e-11 | 2.2317e-11
HGA | A | 3 | 7 | 5 | 2.1924e-11 | 2.2333e-11
HGA | A | 3 | 7 | 6 | 2.1858e-11 | 2.226e-11
HGA | A | 3 | 7 | 7 | 2.1872e-11 | 2.2271e-11
HGA | A | 3 | 7 | 8 | 2.1862e-11 | 2.2259e-11
HGA | A | 3 | 7 | 9 | 2.1899e-11 | 2.1952e-11
HGA | A | 3 | 7 | 10 | 2.1856e-11 | 2.1912e-11
HGA | A | 3 | 7 | 11 | 2.1843e-11 | 2.1895e-11
HGA | A | 3 | 7 | 12 | 2.189e-11 | 2.1934e-11
HGA | A | 3 | 7 | 13 | 2.1814e-11 | 2.1857e-11
HGA | A | 3 | 7 | 14 | 2.1901e-11 | 2.1946e-11
HGA | A | 3 | 7 | 15 | 2.1876e-11 | 2.1918e-11
HGA | A | 3 | 7 | 16 | 2.2066e-11 | 2.2112e-11
HGA | A | 3 | 11 | 1 | 3.2111e-11 | 3.2566e-11
HGA | A | 3 | 11 | 2 | 3.1792e-11 | 3.225e-11
HGA | A | 3 | 11 | 3 | 3.1807e-11 | 3.2262e-11
HGA | A | 3 | 11 | 4 | 3.1873e-11 | 3.2322e-11
HGA | A | 3 | 11 | 5 | 3.1639e-11 | 3.2092e-11
HGA | A | 3 | 11 | 6 | 3.1616e-11 | 3.2064e-11
HGA | A | 3 | 11 | 7 | 3.1812e-11 | 3.227e-11
HGA | A | 3 | 11 | 8 | 3.1615e-11 | 3.2057e-11
HGA | A | 3 | 11 | 9 | 3.1684e-11 | 3.1642e-11
HGA | A | 3 | 11 | 10 | 3.1791e-11 | 3.1757e-11
HGA | A | 3 | 11 | 11 | 3.1606e-11 | 3.1565e-11
HGA | A | 3 | 11 | 12 | 3.1622e-11 | 3.1579e-11
HGA | A | 3 | 11 | 13 | 3.1766e-11 | 3.1731e-11
HGA | A | 3 | 11 | 14 | 3.1657e-11 | 3.1619e-11
HGA | A | 3 | 11 | 15 | 3.1616e-11 | 3.1567e-11
HGA | A | 3 | 11 | 16 | 3.1902e-11 | 3.1853e-11
HGA | A | 3 | 15 | 1 | 4.9571e-11 | 5.0164e-11
HGA | A | 3 | 15 | 2 | 4.9176e-11 | 4.9771e-11
HGA | A | 3 | 15 | 3 | 4.9096e-11 | 4.969e-11
HGA | A | 3 | 15 | 4 | 4.9278e-11 | 4.9867e-11
HGA | A | 3 | 15 | 5 | 4.9226e-11 | 4.9815e-11
HGA | A | 3 | 15 | 6 | 4.9205e-11 | 4.9795e-11
HGA | A | 3 | 15 | 7 | 4.9252e-11 | 4.9836e-11
HGA | A | 3 | 15 | 8 | 4.8925e-11 | 4.9513e-11
HGA | A | 3 | 15 | 9 | 4.9046e-11 | 4.8871e-11
HGA | A | 3 | 15 | 10 | 4.9225e-11 | 4.9053e-11
HGA | A | 3 | 15 | 11 | 4.9189e-11 | 4.9015e-11
HGA | A | 3 | 15 | 12 | 4.9202e-11 | 4.9033e-11
HGA | A | 3 | 15 | 13 | 4.9219e-11 | 4.9035e-11
HGA | A | 3 | 15 | 14 | 4.8919e-11 | 4.8745e-11
HGA | A | 3 | 15 | 15 | 4.894e-11 | 4.877e-11
HGA | A | 3 | 15 | 16 | 4.9321e-11 | 4.9129e-11
HGA | A | 3 | 20 | 1 | 7.3754e-11 | 7.4658e-11
HGA | A | 3 | 20 | 2 | 7.3623e-11 | 7.4517e-11
HGA | A | 3 | 20 | 3 | 7.3486e-11 | 7.4387e-11
HGA | A | 3 | 20 | 4 | 7.3376e-11 | 7.4279e-11
HGA | A | 3 | 20 | 5 | 7.3323e-11 | 7.4216e-11
HGA | A | 3 | 20 | 6 | 7.3167e-11 | 7.4057e-11
HGA | A | 3 | 20 | 7 | 7.3226e-11 | 7.4107e-11
HGA | A | 3 | 20 | 8 | 7.3171e-11 | 7.4052e-11
HGA | A | 3 | 20 | 9 | 7.3399e-11 | 7.3148e-11
HGA | A | 3 | 20 | 10 | 7.326e-11 | 7.3003e-11
HGA | A | 3 | 20 | 11 | 7.3202e-11 | 7.2947e-11
HGA | A | 3 | 20 | 12 | 7.3324e-11 | 7.3072e-11
HGA | A | 3 | 20 | 13 | 7.3381e-11 | 7.3126e-11
HGA | A | 3 | 20 | 14 | 7.3362e-11 | 7.3109e-11
HGA | A | 3 | 20 | 15 | 7.3357e-11 | 7.3113e-11
HGA | A | 3 | 20 | 16 | 7.3452e-11 | 7.3182e-11
HGA | A | 3 | 26 | 1 | 1.1706e-10 | 1.185e-10
HGA | A | 3 | 26 | 2 | 1.1679e-10 | 1.1825e-10
HGA | A | 3 | 26 | 3 | 1.166e-10 | 1.1803e-10
HGA | A | 3 | 26 | 4 | 1.165e-10 | 1.1796e-10
HGA | A | 3 | 26 | 5 | 1.1642e-10 | 1.1787e-10
HGA | A | 3 | 26 | 6 | 1.1682e-10 | 1.1824e-10
HGA | A | 3 | 26 | 7 | 1.1683e-10 | 1.183e-10
HGA | A | 3 | 26 | 8 | 1.1665e-10 | 1.1811e-10
HGA | A | 3 | 26 | 9 | 1.1699e-10 | 1.1661e-10
HGA | A | 3 | 26 | 10 | 1.1696e-10 | 1.1659e-10
HGA | A | 3 | 26 | 11 | 1.1681e-10 | 1.1644e-10
HGA | A | 3 | 26 | 12 | 1.1647e-10 | 1.1608e-10
HGA | A | 3 | 26 | 13 | 1.1651e-10 | 1.1616e-10
HGA | A | 3 | 26 | 14 | 1.1659e-10 | 1.1622e-10
HGA | A | 3 | 26 | 15 | 1.1651e-10 | 1.1615e-10
HGA | A | 3 | 26 | 16 | 1.1671e-10 | 1.1635e-10
HGA | A | 3 | 32 | 1 | 1.8352e-10 | 1.8591e-10
HGA | A | 3 | 32 | 2 | 1.8305e-10 | 1.8542e-10
HGA | A | 3 | 32 | 3 | 1.8531e-10 | 1.8769e-10
HGA | A | 3 | 32 | 4 | 1.8286e-10 | 1.8527e-10
HGA | A | 3 | 32 | 5 | 1.8246e-10 | 1.8479e-10
HGA | A | 3 | 32 | 6 | 1.8295e-10 | 1.853e-10
HGA | A | 3 | 32 | 7 | 1.8286e-10 | 1.8524e-10
HGA | A | 3 | 32 | 8 | 1.8349e-10 | 1.8586e-10
HGA | A | 3 | 32 | 9 | 1.8402e-10 | 1.8352e-10
HGA | A | 3 | 32 | 10 | 1.8308e-10 | 1.8257e-10
HGA | A | 3 | 32 | 11 | 1.8302e-10 | 1.8255e-10
HGA | A | 3 | 32 | 12 | 1.8272e-10 | 1.822e-10
HGA | A | 3 | 32 | 13 | 1.8277e-10 | 1.8227e-10
HGA | A | 3 | 32 | 14 | 1.8559e-10 | 1.8505e-10
HGA | A | 3 | 32 | 15 | 1.832e-10 | 1.8269e-10
HGA | A | 3 | 32 | 16 | 1.8328e-10 | 1.8277e-10
HGA | A | 3 | 34 | 1 | 7.3923e-12 | 7.5377e-12
HGA | A | 3 | 34 | 2 | 7.461e-12 | 7.6085e-12
HGA | A | 3 | 34 | 3 | 7.394e-12 | 7.5314e-12
HGA | A | 3 | 34 | 4 | 7.3568e-12 | 7.5003e-12
HGA | A | 3 | 34 | 5 | 7.3807e-12 | 7.5194e-12
HGA | A | 3 | 34 | 6 | 7.3452e-12 | 7.4865e-12
HGA | A | 3 | 34 | 7 | 7.3178e-12 | 7.4585e-12
HGA | A | 3 | 34 | 8 | 7.3502e-12 | 7.4869e-12
HGA | A | 3 | 34 | 9 | 7.3585e-12 | 7.3869e-12
HGA | A | 3 | 34 | 10 | 7.3086e-12 | 7.3362e-12
HGA | A | 3 | 34 | 11 | 7.3324e-12 | 7.3573e-12
HGA | A | 3 | 34 | 12 | 7.3423e-12 | 7.3715e-12
HGA | A | 3 | 34 | 13 | 7.3046e-12 | 7.331e-12
HGA | A | 3 | 34 | 14 | 7.3335e-12 | 7.3586e-12
HGA | A | 3 | 34 | 15 | 7.3908e-12 | 7.4179e-12
HGA | A | 3 | 34 | 16 | 7.3285e-12 | 7.3561e-12
HGA | B | 3 | 1 | 1 | 1.4338e-11 | 1.4612e-11
HGA | B | 3 | 1 | 2 | 1.4424e-11 | 1.4701e-11
HGA | B | 3 | 1 | 3 | 1.4281e-11 | 1.4551e-11
HGA | B | 3 | 1 | 4 | 1.4214e-11 | 1.4471e-11
HGA | B | 3 | 1 | 5 | 1.4256e-11 | 1.4517e-11
HGA | B | 3 | 1 | 6 | 1.4186e-11 | 1.4446e-11
HGA | B | 3 | 1 | 7 | 1.4135e-11 | 1.4397e-11
HGA | B | 3 | 1 | 8 | 1.4205e-11 | 1.4468e-11
HGA | B | 3 | 1 | 9 | 1.4226e-11 | 1.4264e-11
HGA | B | 3 | 1 | 10 | 1.4122e-11 | 1.4163e-11
HGA | B | 3 | 1 | 11 | 1.4172e-11 | 1.4206e-11
HGA | B | 3 | 1 | 12 | 1.42e-11 | 1.4238e-11
HGA | B | 3 | 1 | 13 | 1.412e-11 | 1.4158e-11
HGA | B | 3 | 1 | 14 | 1.4176e-11 | 1.4223e-11
HGA | B | 3 | 1 | 15 | 1.4306e-11 | 1.4342e-11
HGA | B | 3 | 1 | 16 | 1.4231e-11 | 1.4277e-11
HGA | B | 3 | 7 | 1 | 2.2198e-11 | 2.2626e-11
HGA | B | 3 | 7 | 2 | 2.2006e-11 | 2.242e-11
HGA | B | 3 | 7 | 3 | 2.201e-11 | 2.2419e-11
HGA | B | 3 | 7 | 4 | 2.1893e-11 | 2.2298e-11
HGA | B | 3 | 7 | 5 | 2.1901e-11 | 2.2308e-11
HGA | B | 3 | 7 | 6 | 2.1835e-11 | 2.2238e-11
HGA | B | 3 | 7 | 7 | 2.1855e-11 | 2.2254e-11
HGA | B | 3 | 7 | 8 | 2.1837e-11 | 2.2238e-11
HGA | B | 3 | 7 | 9 | 2.1878e-11 | 2.1938e-11
HGA | B | 3 | 7 | 10 | 2.184e-11 | 2.1894e-11
HGA | B | 3 | 7 | 11 | 2.1823e-11 | 2.1875e-11
HGA | B | 3 | 7 | 12 | 2.1872e-11 | 2.1917e-11
HGA | B | 3 | 7 | 13 | 2.1785e-11 | 2.1836e-11
HGA | B | 3 | 7 | 14 | 2.1882e-11 | 2.1929e-11
HGA | B | 3 | 7 | 15 | 2.1864e-11 | 2.1904e-11
HGA | B | 3 | 7 | 16 | 2.2054e-11 | 2.2099e-11
HGA | B | 3 | 11 | 1 | 3.2081e-11 | 3.2534e-11
HGA | B | 3 | 11 | 2 | 3.1763e-11 | 3.2216e-11
HGA | B | 3 | 11 | 3 | 3.1796e-11 | 3.2245e-11
HGA | B | 3 | 11 | 4 | 3.1841e-11 | 3.2303e-11
HGA | B | 3 | 11 | 5 | 3.1622e-11 | 3.2056e-11
HGA | B | 3 | 11 | 6 | 3.1587e-11 | 3.2032e-11
HGA | B | 3 | 11 | 7 | 3.1782e-11 | 3.2247e-11
HGA | B | 3 | 11 | 8 | 3.1579e-11 | 3.2038e-11
HGA | B | 3 | 11 | 9 | 3.166e-11 | 3.1625e-11
HGA | B | 3 | 11 | 10 | 3.1765e-11 | 3.1727e-11
HGA | B | 3 | 11 | 11 | 3.1581e-11 | 3.154e-11
HGA | B | 3 | 11 | 12 | 3.1595e-11 | 3.1556e-11
HGA | B | 3 | 11 | 13 | 3.174e-11 | 3.1705e-11
HGA | B | 3 | 11 | 14 | 3.1626e-11 | 3.1593e-11
HGA | B | 3 | 11 | 15 | 3.1591e-11 | 3.1547e-11
HGA | B | 3 | 11 | 16 | 3.1869e-11 | 3.1828e-11
HGA | B | 3 | 15 | 1 | 4.9511e-11 | 5.0108e-11
HGA | B | 3 | 15 | 2 | 4.9118e-11 | 4.9716e-11
HGA | B | 3 | 15 | 3 | 4.9047e-11 | 4.9637e-11
HGA | B | 3 | 15 | 4 | 4.9225e-11 | 4.9804e-11
HGA | B | 3 | 15 | 5 | 4.9186e-11 | 4.9776e-11
HGA | B | 3 | 15 | 6 | 4.9152e-11 | 4.9731e-11
HGA | B | 3 | 15 | 7 | 4.9207e-11 | 4.9783e-11
HGA | B | 3 | 15 | 8 | 4.8869e-11 | 4.9455e-11
HGA | B | 3 | 15 | 9 | 4.9003e-11 | 4.8818e-11
HGA | B | 3 | 15 | 10 | 4.9168e-11 | 4.8989e-11
HGA | B | 3 | 15 | 11 | 4.9156e-11 | 4.8961e-11
HGA | B | 3 | 15 | 12 | 4.9147e-11 | 4.8978e-11
HGA | B | 3 | 15 | 13 | 4.917e-11 | 4.8985e-11
HGA | B | 3 | 15 | 14 | 4.8867e-11 | 4.8692e-11
HGA | B | 3 | 15 | 15 | 4.8894e-11 | 4.8713e-11
HGA | B | 3 | 15 | 16 | 4.9254e-11 | 4.9073e-11
HGA | B | 3 | 20 | 1 | 7.3667e-11 | 7.4582e-11
HGA | B | 3 | 20 | 2 | 7.355e-11 | 7.4441e-11
HGA | B | 3 | 20 | 3 | 7.3434e-11 | 7.4322e-11
HGA | B | 3 | 20 | 4 | 7.3304e-11 | 7.4209e-11
HGA | B | 3 | 20 | 5 | 7.3238e-11 | 7.414e-11
HGA | B | 3 | 20 | 6 | 7.3093e-11 | 7.3986e-11
HGA | B | 3 | 20 | 7 | 7.3157e-11 | 7.4042e-11
HGA | B | 3 | 20 | 8 | 7.3096e-11 | 7.3978e-11
HGA | B | 3 | 20 | 9 | 7.3308e-11 | 7.3064e-11
HGA | B | 3 | 20 | 10 | 7.3181e-11 | 7.2927e-11
HGA | B | 3 | 20 | 11 | 7.3112e-11 | 7.2857e-11
HGA | B | 3 | 20 | 12 | 7.3265e-11 | 7.2999e-11
HGA | B | 3 | 20 | 13 | 7.3307e-11 | 7.3054e-11
HGA | B | 3 | 20 | 14 | 7.3273e-11 | 7.3035e-11
HGA | B | 3 | 20 | 15 | 7.3301e-11 | 7.3034e-11
HGA | B | 3 | 20 | 16 | 7.3369e-11 | 7.3116e-11
HGA | B | 3 | 26 | 1 | 1.1693e-10 | 1.1836e-10
HGA | B | 3 | 26 | 2 | 1.1664e-10 | 1.181e-10
HGA | B | 3 | 26 | 3 | 1.1646e-10 | 1.1791e-10
HGA | B | 3 | 26 | 4 | 1.1634e-10 | 1.1781e-10
HGA | B | 3 | 26 | 5 | 1.1629e-10 | 1.1774e-10
HGA | B | 3 | 26 | 6 | 1.1665e-10 | 1.181e-10
HGA | B | 3 | 26 | 7 | 1.1668e-10 | 1.1815e-10
HGA | B | 3 | 26 | 8 | 1.1653e-10 | 1.1797e-10
HGA | B | 3 | 26 | 9 | 1.1686e-10 | 1.1647e-10
HGA | B | 3 | 26 | 10 | 1.1681e-10 | 1.1643e-10
HGA | B | 3 | 26 | 11 | 1.1667e-10 | 1.163e-10
HGA | B | 3 | 26 | 12 | 1.1633e-10 | 1.1594e-10
HGA | B | 3 | 26 | 13 | 1.1637e-10 | 1.1599e-10
HGA | B | 3 | 26 | 14 | 1.1644e-10 | 1.1607e-10
HGA | B | 3 | 26 | 15 | 1.1637e-10 | 1.1599e-10
HGA | B | 3 | 26 | 16 | 1.1655e-10 | 1.162e-10
HGA | B | 3 | 32 | 1 | 1.8327e-10 | 1.8566e-10
HGA | B | 3 | 32 | 2 | 1.8281e-10 | 1.8517e-10
HGA | B | 3 | 32 | 3 | 1.8502e-10 | 1.8743e-10
HGA | B | 3 | 32 | 4 | 1.8259e-10 | 1.8499e-10
HGA | B | 3 | 32 | 5 | 1.8219e-10 | 1.8455e-10
HGA | B | 3 | 32 | 6 | 1.827e-10 | 1.8505e-10
HGA | B | 3 | 32 | 7 | 1.8257e-10 | 1.8497e-10
HGA | B | 3 | 32 | 8 | 1.8325e-10 | 1.8561e-10
HGA | B | 3 | 32 | 9 | 1.8375e-10 | 1.8329e-10
HGA | B | 3 | 32 | 10 | 1.8282e-10 | 1.8233e-10
HGA | B | 3 | 32 | 11 | 1.8276e-10 | 1.8229e-10
HGA | B | 3 | 32 | 12 | 1.8245e-10 | 1.8198e-10
HGA | B | 3 | 32 | 13 | 1.8251e-10 | 1.8202e-10
HGA | B | 3 | 32 | 14 | 1.8531e-10 | 1.8478e-10
HGA | B | 3 | 32 | 15 | 1.8294e-10 | 1.8242e-10
HGA | B | 3 | 32 | 16 | 1.8303e-10 | 1.8252e-10
HGA | B | 3 | 34 | 1 | 7.3868e-12 | 7.5324e-12
HGA | B | 3 | 34 | 2 | 7.4571e-12 | 7.5999e-12
HGA | B | 3 | 34 | 3 | 7.3959e-12 | 7.5334e-12
HGA | B | 3 | 34 | 4 | 7.3547e-12 | 7.4953e-12
HGA | B | 3 | 34 | 5 | 7.3792e-12 | 7.5196e-12
HGA | B | 3 | 34 | 6 | 7.3362e-12 | 7.4826e-12
HGA | B | 3 | 34 | 7 | 7.3171e-12 | 7.4543e-12
HGA | B | 3 | 34 | 8 | 7.3472e-12 | 7.4822e-12
HGA | B | 3 | 34 | 9 | 7.3509e-12 | 7.376e-12
HGA | B | 3 | 34 | 10 | 7.3047e-12 | 7.3323e-12
HGA | B | 3 | 34 | 11 | 7.3273e-12 | 7.3564e-12
HGA | B | 3 | 34 | 12 | 7.3399e-12 | 7.3716e-12
HGA | B | 3 | 34 | 13 | 7.3015e-12 | 7.3316e-12
HGA | B | 3 | 34 | 14 | 7.3264e-12 | 7.3508e-12
HGA | B | 3 | 34 | 15 | 7.3876e-12 | 7.4147e-12
HGA | B | 3 | 34 | 16 | 7.3244e-12 | 7.3541e-12
HGB | A | 3 | 1 | 1 | 1.4596e-11 | 1.4733e-11
HGB | A | 3 | 1 | 2 | 1.4781e-11 | 1.4911e-11
HGB | A | 3 | 1 | 3 | 1.4735e-11 | 1.487e-11
HGB | A | 3 | 1 | 4 | 1.4754e-11 | 1.4899e-11
HGB | A | 3 | 1 | 5 | 1.4807e-11 | 1.4935e-11
HGB | A | 3 | 1 | 6 | 1.4627e-11 | 1.476e-11
HGB | A | 3 | 1 | 7 | 1.4494e-11 | 1.4618e-11
HGB | A | 3 | 1 | 8 | 1.4529e-11 | 1.4652e-11
HGB | A | 3 | 1 | 9 | 1.4398e-11 | 1.4441e-11
HGB | A | 3 | 1 | 10 | 1.4358e-11 | 1.4402e-11
HGB | A | 3 | 1 | 11 | 1.4417e-11 | 1.4462e-11
HGB | A | 3 | 1 | 12 | 1.4472e-11 | 1.4518e-11
HGB | A | 3 | 1 | 13 | 1.4394e-11 | 1.4437e-11
HGB | A | 3 | 1 | 14 | 1.4369e-11 | 1.4413e-11
HGB | A | 3 | 1 | 15 | 1.4407e-11 | 1.4445e-11
HGB | A | 3 | 1 | 16 | 1.4311e-11 | 1.4354e-11
HGB | A | 3 | 7 | 1 | 2.2691e-11 | 2.2929e-11
HGB | A | 3 | 7 | 2 | 2.2609e-11 | 2.2848e-11
HGB | A | 3 | 7 | 3 | 2.2764e-11 | 2.2993e-11
HGB | A | 3 | 7 | 4 | 2.2708e-11 | 2.2928e-11
HGB | A | 3 | 7 | 5 | 2.2635e-11 | 2.2862e-11
HGB | A | 3 | 7 | 6 | 2.2415e-11 | 2.2641e-11
HGB | A | 3 | 7 | 7 | 2.2345e-11 | 2.256e-11
HGB | A | 3 | 7 | 8 | 2.2307e-11 | 2.2533e-11
HGB | A | 3 | 7 | 9 | 2.2123e-11 | 2.2203e-11
HGB | A | 3 | 7 | 10 | 2.2145e-11 | 2.2234e-11
HGB | A | 3 | 7 | 11 | 2.2174e-11 | 2.2251e-11
HGB | A | 3 | 7 | 12 | 2.2218e-11 | 2.2304e-11
HGB | A | 3 | 7 | 13 | 2.2173e-11 | 2.2257e-11
HGB | A | 3 | 7 | 14 | 2.2239e-11 | 2.232e-11
HGB | A | 3 | 7 | 15 | 2.2088e-11 | 2.2178e-11
HGB | A | 3 | 7 | 16 | 2.2172e-11 | 2.2269e-11
HGB | A | 3 | 11 | 1 | 3.2964e-11 | 3.3118e-11
HGB | A | 3 | 11 | 2 | 3.2812e-11 | 3.2938e-11
HGB | A | 3 | 11 | 3 | 3.2982e-11 | 3.3129e-11
HGB | A | 3 | 11 | 4 | 3.3023e-11 | 3.3154e-11
HGB | A | 3 | 11 | 5 | 3.2598e-11 | 3.2733e-11
HGB | A | 3 | 11 | 6 | 3.2395e-11 | 3.2532e-11
HGB | A | 3 | 11 | 7 | 3.2496e-11 | 3.263e-11
HGB | A | 3 | 11 | 8 | 3.229e-11 | 3.2422e-11
HGB | A | 3 | 11 | 9 | 3.1996e-11 | 3.1935e-11
HGB | A | 3 | 11 | 10 | 3.2195e-11 | 3.2144e-11
HGB | A | 3 | 11 | 11 | 3.2057e-11 | 3.2018e-11
HGB | A | 3 | 11 | 12 | 3.2089e-11 | 3.2026e-11
HGB | A | 3 | 11 | 13 | 3.2304e-11 | 3.2238e-11
HGB | A | 3 | 11 | 14 | 3.22e-11 | 3.2153e-11
HGB | A | 3 | 11 | 15 | 3.2047e-11 | 3.1995e-11
HGB | A | 3 | 11 | 16 | 3.2182e-11 | 3.2137e-11
HGB | A | 3 | 15 | 1 | 5.1044e-11 | 5.1165e-11
HGB | A | 3 | 15 | 2 | 5.0815e-11 | 5.0936e-11
HGB | A | 3 | 15 | 3 | 5.0784e-11 | 5.0904e-11
HGB | A | 3 | 15 | 4 | 5.0817e-11 | 5.093e-11
HGB | A | 3 | 15 | 5 | 5.0506e-11 | 5.0609e-11
HGB | A | 3 | 15 | 6 | 5.0289e-11 | 5.04e-11
HGB | A | 3 | 15 | 7 | 5.0227e-11 | 5.0335e-11
HGB | A | 3 | 15 | 8 | 4.9896e-11 | 4.9997e-11
HGB | A | 3 | 15 | 9 | 4.9443e-11 | 4.9255e-11
HGB | A | 3 | 15 | 10 | 4.9764e-11 | 4.9579e-11
HGB | A | 3 | 15 | 11 | 4.9828e-11 | 4.9645e-11
HGB | A | 3 | 15 | 12 | 4.9879e-11 | 4.9695e-11
HGB | A | 3 | 15 | 13 | 4.9893e-11 | 4.9711e-11
HGB | A | 3 | 15 | 14 | 4.9664e-11 | 4.9477e-11
HGB | A | 3 | 15 | 15 | 4.9671e-11 | 4.9506e-11
HGB | A | 3 | 15 | 16 | 4.9905e-11 | 4.9731e-11
HGB | A | 3 | 20 | 1 | 7.628e-11 | 7.6469e-11
HGB | A | 3 | 20 | 2 | 7.6227e-11 | 7.6427e-11
HGB | A | 3 | 20 | 3 | 7.5956e-11 | 7.615e-11
HGB | A | 3 | 20 | 4 | 7.5559e-11 | 7.5741e-11
HGB | A | 3 | 20 | 5 | 7.5137e-11 | 7.5318e-11
HGB | A | 3 | 20 | 6 | 7.4818e-11 | 7.5e-11
HGB | A | 3 | 20 | 7 | 7.4755e-11 | 7.4937e-11
HGB | A | 3 | 20 | 8 | 7.4725e-11 | 7.4892e-11
HGB | A | 3 | 20 | 9 | 7.4022e-11 | 7.376e-11
HGB | A | 3 | 20 | 10 | 7.4016e-11 | 7.377e-11
HGB | A | 3 | 20 | 11 | 7.4114e-11 | 7.3866e-11
HGB | A | 3 | 20 | 12 | 7.4386e-11 | 7.4146e-11
HGB | A | 3 | 20 | 13 | 7.4345e-11 | 7.4121e-11
HGB | A | 3 | 20 | 14 | 7.4421e-11 | 7.418e-11
HGB | A | 3 | 20 | 15 | 7.4489e-11 | 7.4238e-11
HGB | A | 3 | 20 | 16 | 7.4538e-11 | 7.4279e-11
HGB | A | 3 | 26 | 1 | 1.2113e-10 | 1.2142e-10
HGB | A | 3 | 26 | 2 | 1.2064e-10 | 1.2097e-10
HGB | A | 3 | 26 | 3 | 1.2011e-10 | 1.2042e-10
HGB | A | 3 | 26 | 4 | 1.1951e-10 | 1.1981e-10
HGB | A | 3 | 26 | 5 | 1.1906e-10 | 1.1936e-10
HGB | A | 3 | 26 | 6 | 1.1918e-10 | 1.1948e-10
HGB | A | 3 | 26 | 7 | 1.1913e-10 | 1.1944e-10
HGB | A | 3 | 26 | 8 | 1.1894e-10 | 1.1922e-10
HGB | A | 3 | 26 | 9 | 1.1789e-10 | 1.175e-10
HGB | A | 3 | 26 | 10 | 1.1794e-10 | 1.1754e-10
HGB | A | 3 | 26 | 11 | 1.1809e-10 | 1.1768e-10
HGB | A | 3 | 26 | 12 | 1.1793e-10 | 1.1754e-10
HGB | A | 3 | 26 | 13 | 1.1798e-10 | 1.176e-10
HGB | A | 3 | 26 | 14 | 1.1804e-10 | 1.1763e-10
HGB | A | 3 | 26 | 15 | 1.1814e-10 | 1.1775e-10
HGB | A | 3 | 26 | 16 | 1.1837e-10 | 1.18e-10
HGB | A | 3 | 32 | 1 | 1.894e-10 | 1.9006e-10
HGB | A | 3 | 32 | 2 | 1.8839e-10 | 1.8903e-10
HGB | A | 3 | 32 | 3 | 1.9002e-10 | 1.9064e-10
HGB | A | 3 | 32 | 4 | 1.8695e-10 | 1.8761e-10
HGB | A | 3 | 32 | 5 | 1.8626e-10 | 1.8683e-10
HGB | A | 3 | 32 | 6 | 1.8631e-10 | 1.8693e-10
HGB | A | 3 | 32 | 7 | 1.8618e-10 | 1.8679e-10
HGB | A | 3 | 32 | 8 | 1.8682e-10 | 1.8745e-10
HGB | A | 3 | 32 | 9 | 1.8524e-10 | 1.8479e-10
HGB | A | 3 | 32 | 10 | 1.8433e-10 | 1.8387e-10
HGB | A | 3 | 32 | 11 | 1.8475e-10 | 1.8432e-10
HGB | A | 3 | 32 | 12 | 1.8466e-10 | 1.8418e-10
HGB | A | 3 | 32 | 13 | 1.8507e-10 | 1.8464e-10
HGB | A | 3 | 32 | 14 | 1.875e-10 | 1.8703e-10
HGB | A | 3 | 32 | 15 | 1.8519e-10 | 1.8474e-10
HGB | A | 3 | 32 | 16 | 1.8563e-10 | 1.8516e-10
HGB | A | 3 | 34 | 1 | 7.5126e-12 | 7.5826e-12
HGB | A | 3 | 34 | 2 | 7.6292e-12 | 7.7019e-12
HGB | A | 3 | 34 | 3 | 7.6104e-12 | 7.6856e-12
HGB | A | 3 | 34 | 4 | 7.6251e-12 | 7.696e-12
HGB | A | 3 | 34 | 5 | 7.6481e-12 | 7.717e-12
HGB | A | 3 | 34 | 6 | 7.5594e-12 | 7.628e-12
HGB | A | 3 | 34 | 7 | 7.4894e-12 | 7.5549e-12
HGB | A | 3 | 34 | 8 | 7.5133e-12 | 7.5804e-12
HGB | A | 3 | 34 | 9 | 7.4416e-12 | 7.4687e-12
HGB | A | 3 | 34 | 10 | 7.4222e-12 | 7.4452e-12
HGB | A | 3 | 34 | 11 | 7.4528e-12 | 7.4811e-12
HGB | A | 3 | 34 | 12 | 7.4839e-12 | 7.5096e-12
HGB | A | 3 | 34 | 13 | 7.4341e-12 | 7.4628e-12
HGB | A | 3 | 34 | 14 | 7.4309e-12 | 7.4549e-12
HGB | A | 3 | 34 | 15 | 7.4453e-12 | 7.4667e-12
HGB | A | 3 | 34 | 16 | 7.3616e-12 | 7.3906e-12
HGB | B | 3 | 1 | 1 | 1.4583e-11 | 1.4718e-11
HGB | B | 3 | 1 | 2 | 1.4768e-11 | 1.4901e-11
HGB | B | 3 | 1 | 3 | 1.4729e-11 | 1.486e-11
HGB | B | 3 | 1 | 4 | 1.4744e-11 | 1.4884e-11
HGB | B | 3 | 1 | 5 | 1.479e-11 | 1.4928e-11
HGB | B | 3 | 1 | 6 | 1.4613e-11 | 1.4741e-11
HGB | B | 3 | 1 | 7 | 1.4481e-11 | 1.4607e-11
HGB | B | 3 | 1 | 8 | 1.4521e-11 | 1.4646e-11
HGB | B | 3 | 1 | 9 | 1.4392e-11 | 1.4429e-11
HGB | B | 3 | 1 | 10 | 1.4344e-11 | 1.4389e-11
HGB | B | 3 | 1 | 11 | 1.4411e-11 | 1.4448e-11
HGB | B | 3 | 1 | 12 | 1.4453e-11 | 1.4505e-11
HGB | B | 3 | 1 | 13 | 1.4372e-11 | 1.443e-11
HGB | B | 3 | 1 | 14 | 1.4362e-11 | 1.4407e-11
HGB | B | 3 | 1 | 15 | 1.4395e-11 | 1.4445e-11
HGB | B | 3 | 1 | 16 | 1.4296e-11 | 1.4343e-11
HGB | B | 3 | 7 | 1 | 2.2668e-11 | 2.2907e-11
HGB | B | 3 | 7 | 2 | 2.259e-11 | 2.2824e-11
HGB | B | 3 | 7 | 3 | 2.2741e-11 | 2.2981e-11
HGB | B | 3 | 7 | 4 | 2.2675e-11 | 2.2913e-11
HGB | B | 3 | 7 | 5 | 2.2616e-11 | 2.2844e-11
HGB | B | 3 | 7 | 6 | 2.2393e-11 | 2.2617e-11
HGB | B | 3 | 7 | 7 | 2.2314e-11 | 2.2541e-11
HGB | B | 3 | 7 | 8 | 2.2284e-11 | 2.2508e-11
HGB | B | 3 | 7 | 9 | 2.2091e-11 | 2.2178e-11
HGB | B | 3 | 7 | 10 | 2.2119e-11 | 2.221e-11
HGB | B | 3 | 7 | 11 | 2.2146e-11 | 2.224e-11
HGB | B | 3 | 7 | 12 | 2.2199e-11 | 2.2287e-11
HGB | B | 3 | 7 | 13 | 2.2155e-11 | 2.2244e-11
HGB | B | 3 | 7 | 14 | 2.2215e-11 | 2.2296e-11
HGB | B | 3 | 7 | 15 | 2.2069e-11 | 2.2153e-11
HGB | B | 3 | 7 | 16 | 2.2152e-11 | 2.2236e-11
HGB | B | 3 | 11 | 1 | 3.293e-11 | 3.3077e-11
HGB | B | 3 | 11 | 2 | 3.2783e-11 | 3.2925e-11
HGB | B | 3 | 11 | 3 | 3.2947e-11 | 3.309e-11
HGB | B | 3 | 11 | 4 | 3.2982e-11 | 3.3133e-11
HGB | B | 3 | 11 | 5 | 3.2566e-11 | 3.2706e-11
HGB | B | 3 | 11 | 6 | 3.2357e-11 | 3.2495e-11
HGB | B | 3 | 11 | 7 | 3.2467e-11 | 3.2597e-11
HGB | B | 3 | 11 | 8 | 3.2253e-11 | 3.2395e-11
HGB | B | 3 | 11 | 9 | 3.1957e-11 | 3.1898e-11
HGB | B | 3 | 11 | 10 | 3.2171e-11 | 3.2116e-11
HGB | B | 3 | 11 | 11 | 3.2039e-11 | 3.1991e-11
HGB | B | 3 | 11 | 12 | 3.2052e-11 | 3.1998e-11
HGB | B | 3 | 11 | 13 | 3.2259e-11 | 3.2216e-11
HGB | B | 3 | 11 | 14 | 3.2173e-11 | 3.2128e-11
HGB | B | 3 | 11 | 15 | 3.201e-11 | 3.1975e-11
HGB | B | 3 | 11 | 16 | 3.215e-11 | 3.2105e-11
HGB | B | 3 | 15 | 1 | 5.0985e-11 | 5.1105e-11
HGB | B | 3 | 15 | 2 | 5.0755e-11 | 5.0888e-11
HGB | B | 3 | 15 | 3 | 5.0712e-11 | 5.0841e-11
HGB | B | 3 | 15 | 4 | 5.0766e-11 | 5.0876e-11
HGB | B | 3 | 15 | 5 | 5.0447e-11 | 5.0552e-11
HGB | B | 3 | 15 | 6 | 5.023e-11 | 5.035e-11
HGB | B | 3 | 15 | 7 | 5.0166e-11 | 5.0284e-11
HGB | B | 3 | 15 | 8 | 4.9827e-11 | 4.9951e-11
HGB | B | 3 | 15 | 9 | 4.9381e-11 | 4.9205e-11
HGB | B | 3 | 15 | 10 | 4.9696e-11 | 4.9528e-11
HGB | B | 3 | 15 | 11 | 4.9779e-11 | 4.9589e-11
HGB | B | 3 | 15 | 12 | 4.982e-11 | 4.9639e-11
HGB | B | 3 | 15 | 13 | 4.9828e-11 | 4.9659e-11
HGB | B | 3 | 15 | 14 | 4.9596e-11 | 4.9418e-11
HGB | B | 3 | 15 | 15 | 4.9618e-11 | 4.9449e-11
HGB | B | 3 | 15 | 16 | 4.9853e-11 | 4.9673e-11
HGB | B | 3 | 20 | 1 | 7.6186e-11 | 7.6398e-11
HGB | B | 3 | 20 | 2 | 7.6139e-11 | 7.6347e-11
HGB | B | 3 | 20 | 3 | 7.5853e-11 | 7.6049e-11
HGB | B | 3 | 20 | 4 | 7.5462e-11 | 7.5655e-11
HGB | B | 3 | 20 | 5 | 7.5037e-11 | 7.5224e-11
HGB | B | 3 | 20 | 6 | 7.4728e-11 | 7.492e-11
HGB | B | 3 | 20 | 7 | 7.4665e-11 | 7.4847e-11
HGB | B | 3 | 20 | 8 | 7.4626e-11 | 7.48e-11
HGB | B | 3 | 20 | 9 | 7.3929e-11 | 7.3684e-11
HGB | B | 3 | 20 | 10 | 7.3919e-11 | 7.3671e-11
HGB | B | 3 | 20 | 11 | 7.4019e-11 | 7.3766e-11
HGB | B | 3 | 20 | 12 | 7.4283e-11 | 7.4069e-11
HGB | B | 3 | 20 | 13 | 7.4242e-11 | 7.4029e-11
HGB | B | 3 | 20 | 14 | 7.4332e-11 | 7.4093e-11
HGB | B | 3 | 20 | 15 | 7.4416e-11 | 7.4148e-11
HGB | B | 3 | 20 | 16 | 7.4454e-11 | 7.4199e-11
HGB | B | 3 | 26 | 1 | 1.2096e-10 | 1.2129e-10
HGB | B | 3 | 26 | 2 | 1.2049e-10 | 1.2081e-10
HGB | B | 3 | 26 | 3 | 1.1995e-10 | 1.2027e-10
HGB | B | 3 | 26 | 4 | 1.1935e-10 | 1.1966e-10
HGB | B | 3 | 26 | 5 | 1.1891e-10 | 1.192e-10
HGB | B | 3 | 26 | 6 | 1.1901e-10 | 1.1933e-10
HGB | B | 3 | 26 | 7 | 1.1898e-10 | 1.1927e-10
HGB | B | 3 | 26 | 8 | 1.1878e-10 | 1.1908e-10
HGB | B | 3 | 26 | 9 | 1.1772e-10 | 1.1734e-10
HGB | B | 3 | 26 | 10 | 1.1779e-10 | 1.1739e-10
HGB | B | 3 | 26 | 11 | 1.1792e-10 | 1.1752e-10
HGB | B | 3 | 26 | 12 | 1.1776e-10 | 1.1737e-10
HGB | B | 3 | 26 | 13 | 1.1783e-10 | 1.1743e-10
HGB | B | 3 | 26 | 14 | 1.1786e-10 | 1.1749e-10
HGB | B | 3 | 26 | 15 | 1.1797e-10 | 1.176e-10
HGB | B | 3 | 26 | 16 | 1.1822e-10 | 1.1784e-10
HGB | B | 3 | 32 | 1 | 1.8912e-10 | 1.8977e-10
HGB | B | 3 | 32 | 2 | 1.8811e-10 | 1.8874e-10
HGB | B | 3 | 32 | 3 | 1.8974e-10 | 1.9037e-10
HGB | B | 3 | 32 | 4 | 1.8667e-10 | 1.8733e-10
HGB | B | 3 | 32 | 5 | 1.8598e-10 | 1.866e-10
HGB | B | 3 | 32 | 6 | 1.8603e-10 | 1.8665e-10
HGB | B | 3 | 32 | 7 | 1.8591e-10 | 1.8652e-10
HGB | B | 3 | 32 | 8 | 1.8655e-10 | 1.8715e-10
HGB | B | 3 | 32 | 9 | 1.8497e-10 | 1.8451e-10
HGB | B | 3 | 32 | 10 | 1.8405e-10 | 1.8361e-10
HGB | B | 3 | 32 | 11 | 1.8447e-10 | 1.8406e-10
HGB | B | 3 | 32 | 12 | 1.8438e-10 | 1.8393e-10
HGB | B | 3 | 32 | 13 | 1.8481e-10 | 1.8437e-10
HGB | B | 3 | 32 | 14 | 1.872e-10 | 1.8677e-10
HGB | B | 3 | 32 | 15 | 1.8492e-10 | 1.8447e-10
HGB | B | 3 | 32 | 16 | 1.8536e-10 | 1.8489e-10
HGB | B | 3 | 34 | 1 | 7.5039e-12 | 7.5738e-12
HGB | B | 3 | 34 | 2 | 7.6163e-12 | 7.6926e-12
HGB | B | 3 | 34 | 3 | 7.6084e-12 | 7.679e-12
HGB | B | 3 | 34 | 4 | 7.6174e-12 | 7.6939e-12
HGB | B | 3 | 34 | 5 | 7.6458e-12 | 7.719e-12
HGB | B | 3 | 34 | 6 | 7.5546e-12 | 7.6214e-12
HGB | B | 3 | 34 | 7 | 7.4836e-12 | 7.5493e-12
HGB | B | 3 | 34 | 8 | 7.5059e-12 | 7.5685e-12
HGB | B | 3 | 34 | 9 | 7.4395e-12 | 7.4652e-12
HGB | B | 3 | 34 | 10 | 7.4155e-12 | 7.4431e-12
HGB | B | 3 | 34 | 11 | 7.4469e-12 | 7.4758e-12
HGB | B | 3 | 34 | 12 | 7.4723e-12 | 7.5058e-12
HGB | B | 3 | 34 | 13 | 7.4298e-12 | 7.4571e-12
HGB | B | 3 | 34 | 14 | 7.4205e-12 | 7.4444e-12
HGB | B | 3 | 34 | 15 | 7.4305e-12 | 7.4602e-12
HGB | B | 3 | 34 | 16 | 7.3527e-12 | 7.3815e-12

## TEB sensitivity
Band | HAM | Subsample | Gain State | Band Median |  | Band Average |  | NEdT at Ttyp Spec
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | NEdT at Ttyp | NEdT at Ttyp | NEdT at Ttyp | NEdT at Ttyp
I4 | A | 1 | H | 0.52356 | 0.61719 | 0.52059 | 0.62092 | 2.5
I4 | A | 2 | H | 0.52146 | 0.63031 | 0.52209 | 0.62357 | 2.5
I4 | B | 1 | H | 0.52837 | 0.62897 | 0.52745 | 0.62896 | 2.5
I4 | B | 2 | H | 0.52665 | 0.63096 | 0.52785 | 0.63233 | 2.5
I5 | A | 1 | H | 0.45462 | 0.46942 | 0.47226 | 0.46988 | 1.5
I5 | A | 2 | H | 0.4541 | 0.46592 | 0.47206 | 0.46758 | 1.5
I5 | B | 1 | H | 0.45423 | 0.46951 | 0.47312 | 0.46864 | 1.5
I5 | B | 2 | H | 0.45012 | 0.46549 | 0.47298 | 0.4683 | 1.5
M12 | A | 1 | H | 0.16159 | 0.18345 | 0.16162 | 0.18723 | 0.396
M12 | B | 1 | H | 0.16139 | 0.18041 | 0.16032 | 0.18146 | 0.396
M13 | A | 1 | H | 0.057504 | 0.05424 | 0.057579 | 0.054423 | 0.107
M13 | B | 1 | H | 0.057805 | 0.05455 | 0.057751 | 0.054405 | 0.107
M14 | A | 1 | H | 0.06973 | 0.072485 | 0.070249 | 0.073839 | 0.091
M14 | B | 1 | H | 0.070297 | 0.072425 | 0.071016 | 0.073918 | 0.091
M15 | A | 1 | H | 0.092764 | 0.054946 | 0.092531 | 0.055346 | 0.07
M15 | B | 1 | H | 0.090673 | 0.055037 | 0.092146 | 0.055646 | 0.07
M16A | A | 1 | H | 0.11186 | 0.1075 | 0.10966 | 0.11205 | 0.072
M16A | B | 1 | H | 0.1015 | 0.1024 | 0.10332 | 0.10554 | 0.072
M16B | A | 1 | H | 0.1506 | 0.11368 | 0.15525 | 0.11136 | 0.072
M16B | B | 1 | H | 0.15005 | 0.095004 | 0.14899 | 0.10116 | 0.072
M13 | A | 1 | L | 0.27018 | 0.28026 | 0.26305 | 0.28641 | 0.423
M13 | B | 1 | L | 0.26389 | 0.28071 | 0.26361 | 0.2813 | 0.423
Band | HAM | Detector | Subsample | Gain State | Detector Dependent |  | NEdT at Ttyp Spec |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 TEB radiometric sensitivity.
[5 empty] | NEdT at Ttyp | NEdT at Ttyp |  |  | 2. The results were derived from RC-1 using the LABB source (TMC BB for M13 low gain).
I4 | A | 1 | 1 | H | 0.49467 | 0.60029 | 2.5 |  | 3. Fits of the LABB path radiance difference vs SNR were used to determine the radiometric sensitivity.
I4 | A | 2 | 1 | H | 0.50849 | 0.68818 | 2.5 |  | 4. NEdT values are in K and were converted to beginning of scan.
I4 | A | 3 | 1 | H | 0.52503 | 0.59179 | 2.5 |  | 5. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I4 | A | 4 | 1 | H | 0.51913 | 0.61149 | 2.5 |  | 6. The test was set to operational day mode.
I4 | A | 5 | 1 | H | 0.53473 | 0.65409 | 2.5
I4 | A | 6 | 1 | H | 0.50364 | 0.60617 | 2.5
I4 | A | 7 | 1 | H | 0.50819 | 0.61552 | 2.5
I4 | A | 8 | 1 | H | 0.5196 | 0.64394 | 2.5
I4 | A | 9 | 1 | H | 0.50722 | 0.64691 | 2.5
I4 | A | 10 | 1 | H | 0.53476 | 0.64379 | 2.5
I4 | A | 11 | 1 | H | 0.51254 | 0.6461 | 2.5
I4 | A | 12 | 1 | H | 0.50199 | 0.66217 | 2.5
I4 | A | 13 | 1 | H | 0.52927 | 0.64552 | 2.5
I4 | A | 14 | 1 | H | 0.5252 | 0.61886 | 2.5
I4 | A | 15 | 1 | H | 0.5338 | 0.6042 | 2.5
I4 | A | 16 | 1 | H | 0.50661 | 0.60808 | 2.5
I4 | A | 17 | 1 | H | 0.50602 | 0.60546 | 2.5
I4 | A | 18 | 1 | H | 0.52804 | 0.64357 | 2.5
I4 | A | 19 | 1 | H | 0.5221 | 0.65592 | 2.5
I4 | A | 20 | 1 | H | 0.52883 | 0.62252 | 2.5
I4 | A | 21 | 1 | H | 0.54744 | 0.63691 | 2.5
I4 | A | 22 | 1 | H | 0.52645 | 0.63619 | 2.5
I4 | A | 23 | 1 | H | 0.51638 | 0.5893 | 2.5
I4 | A | 24 | 1 | H | 0.53099 | 0.61986 | 2.5
I4 | A | 25 | 1 | H | 0.53877 | 0.61108 | 2.5
I4 | A | 26 | 1 | H | 0.53358 | 0.583 | 2.5
I4 | A | 27 | 1 | H | 0.51762 | 0.65223 | 2.5
I4 | A | 28 | 1 | H | 0.53337 | 0.59116 | 2.5
I4 | A | 29 | 1 | H | 0.51404 | 0.57683 | 2.5
I4 | A | 30 | 1 | H | 0.52566 | 0.59563 | 2.5
I4 | A | 31 | 1 | H | 0.52928 | 0.58995 | 2.5
I4 | A | 32 | 1 | H | 0.49548 | 0.57272 | 2.5
I4 | A | 1 | 2 | H | 0.50857 | 0.60625 | 2.5
I4 | A | 2 | 2 | H | 0.53264 | 0.63289 | 2.5
I4 | A | 3 | 2 | H | 0.51756 | 0.65347 | 2.5
I4 | A | 4 | 2 | H | 0.53206 | 0.68308 | 2.5
I4 | A | 5 | 2 | H | 0.53591 | 0.63289 | 2.5
I4 | A | 6 | 2 | H | 0.51312 | 0.59555 | 2.5
I4 | A | 7 | 2 | H | 0.52398 | 0.62971 | 2.5
I4 | A | 8 | 2 | H | 0.52023 | 0.61328 | 2.5
I4 | A | 9 | 2 | H | 0.52065 | 0.63377 | 2.5
I4 | A | 10 | 2 | H | 0.52227 | 0.6351 | 2.5
I4 | A | 11 | 2 | H | 0.51271 | 0.61353 | 2.5
I4 | A | 12 | 2 | H | 0.51365 | 0.66902 | 2.5
I4 | A | 13 | 2 | H | 0.52478 | 0.63563 | 2.5
I4 | A | 14 | 2 | H | 0.54121 | 0.65393 | 2.5
I4 | A | 15 | 2 | H | 0.52874 | 0.63792 | 2.5
I4 | A | 16 | 2 | H | 0.52898 | 0.63737 | 2.5
I4 | A | 17 | 2 | H | 0.49762 | 0.64087 | 2.5
I4 | A | 18 | 2 | H | 0.53166 | 0.63091 | 2.5
I4 | A | 19 | 2 | H | 0.52331 | 0.63438 | 2.5
I4 | A | 20 | 2 | H | 0.51788 | 0.60639 | 2.5
I4 | A | 21 | 2 | H | 0.50892 | 0.63514 | 2.5
I4 | A | 22 | 2 | H | 0.53118 | 0.59778 | 2.5
I4 | A | 23 | 2 | H | 0.51749 | 0.62178 | 2.5
I4 | A | 24 | 2 | H | 0.54675 | 0.60773 | 2.5
I4 | A | 25 | 2 | H | 0.51371 | 0.63829 | 2.5
I4 | A | 26 | 2 | H | 0.51301 | 0.60838 | 2.5
I4 | A | 27 | 2 | H | 0.5325 | 0.59908 | 2.5
I4 | A | 28 | 2 | H | 0.53382 | 0.58437 | 2.5
I4 | A | 29 | 2 | H | 0.52327 | 0.59235 | 2.5
I4 | A | 30 | 2 | H | 0.51284 | 0.60618 | 2.5
I4 | A | 31 | 2 | H | 0.51583 | 0.58477 | 2.5
I4 | A | 32 | 2 | H | 0.5099 | 0.60236 | 2.5
I4 | B | 1 | 1 | H | 0.53568 | 0.65101 | 2.5
I4 | B | 2 | 1 | H | 0.54268 | 0.63963 | 2.5
I4 | B | 3 | 1 | H | 0.5129 | 0.65684 | 2.5
I4 | B | 4 | 1 | H | 0.50642 | 0.67438 | 2.5
I4 | B | 5 | 1 | H | 0.53401 | 0.67476 | 2.5
I4 | B | 6 | 1 | H | 0.53418 | 0.61562 | 2.5
I4 | B | 7 | 1 | H | 0.5654 | 0.64693 | 2.5
I4 | B | 8 | 1 | H | 0.51572 | 0.58869 | 2.5
I4 | B | 9 | 1 | H | 0.51111 | 0.64323 | 2.5
I4 | B | 10 | 1 | H | 0.51259 | 0.66921 | 2.5
I4 | B | 11 | 1 | H | 0.54218 | 0.67487 | 2.5
I4 | B | 12 | 1 | H | 0.51878 | 0.65516 | 2.5
I4 | B | 13 | 1 | H | 0.5307 | 0.65282 | 2.5
I4 | B | 14 | 1 | H | 0.52867 | 0.59568 | 2.5
I4 | B | 15 | 1 | H | 0.53104 | 0.61315 | 2.5
I4 | B | 16 | 1 | H | 0.52904 | 0.65939 | 2.5
I4 | B | 17 | 1 | H | 0.53892 | 0.64313 | 2.5
I4 | B | 18 | 1 | H | 0.55828 | 0.71961 | 2.5
I4 | B | 19 | 1 | H | 0.53926 | 0.63948 | 2.5
I4 | B | 20 | 1 | H | 0.53646 | 0.58992 | 2.5
I4 | B | 21 | 1 | H | 0.5084 | 0.62643 | 2.5
I4 | B | 22 | 1 | H | 0.51971 | 0.63151 | 2.5
I4 | B | 23 | 1 | H | 0.52119 | 0.61265 | 2.5
I4 | B | 24 | 1 | H | 0.53226 | 0.61837 | 2.5
I4 | B | 25 | 1 | H | 0.55996 | 0.58058 | 2.5
I4 | B | 26 | 1 | H | 0.51116 | 0.61755 | 2.5
I4 | B | 27 | 1 | H | 0.5153 | 0.62289 | 2.5
I4 | B | 28 | 1 | H | 0.52808 | 0.58797 | 2.5
I4 | B | 29 | 1 | H | 0.49516 | 0.58484 | 2.5
I4 | B | 30 | 1 | H | 0.51839 | 0.59485 | 2.5
I4 | B | 31 | 1 | H | 0.51802 | 0.57688 | 2.5
I4 | B | 32 | 1 | H | 0.5267 | 0.56861 | 2.5
I4 | B | 1 | 2 | H | 0.51733 | 0.69975 | 2.5
I4 | B | 2 | 2 | H | 0.54412 | 0.62069 | 2.5
I4 | B | 3 | 2 | H | 0.49895 | 0.60291 | 2.5
I4 | B | 4 | 2 | H | 0.51614 | 0.66019 | 2.5
I4 | B | 5 | 2 | H | 0.52625 | 0.61661 | 2.5
I4 | B | 6 | 2 | H | 0.53419 | 0.65118 | 2.5
I4 | B | 7 | 2 | H | 0.5268 | 0.65919 | 2.5
I4 | B | 8 | 2 | H | 0.52306 | 0.67987 | 2.5
I4 | B | 9 | 2 | H | 0.49959 | 0.65485 | 2.5
I4 | B | 10 | 2 | H | 0.52705 | 0.67557 | 2.5
I4 | B | 11 | 2 | H | 0.51937 | 0.62751 | 2.5
I4 | B | 12 | 2 | H | 0.52213 | 0.632 | 2.5
I4 | B | 13 | 2 | H | 0.53052 | 0.64719 | 2.5
I4 | B | 14 | 2 | H | 0.52604 | 0.62447 | 2.5
I4 | B | 15 | 2 | H | 0.52627 | 0.61547 | 2.5
I4 | B | 16 | 2 | H | 0.53825 | 0.6116 | 2.5
I4 | B | 17 | 2 | H | 0.53044 | 0.63639 | 2.5
I4 | B | 18 | 2 | H | 0.55563 | 0.63169 | 2.5
I4 | B | 19 | 2 | H | 0.52649 | 0.63388 | 2.5
I4 | B | 20 | 2 | H | 0.52725 | 0.65371 | 2.5
I4 | B | 21 | 2 | H | 0.52479 | 0.59887 | 2.5
I4 | B | 22 | 2 | H | 0.52528 | 0.61588 | 2.5
I4 | B | 23 | 2 | H | 0.5343 | 0.62973 | 2.5
I4 | B | 24 | 2 | H | 0.53406 | 0.61639 | 2.5
I4 | B | 25 | 2 | H | 0.54511 | 0.56871 | 2.5
I4 | B | 26 | 2 | H | 0.52974 | 0.60378 | 2.5
I4 | B | 27 | 2 | H | 0.51576 | 0.68323 | 2.5
I4 | B | 28 | 2 | H | 0.55572 | 0.63352 | 2.5
I4 | B | 29 | 2 | H | 0.52749 | 0.63529 | 2.5
I4 | B | 30 | 2 | H | 0.5168 | 0.57304 | 2.5
I4 | B | 31 | 2 | H | 0.54753 | 0.61121 | 2.5
I4 | B | 32 | 2 | H | 0.51889 | 0.63022 | 2.5
I5 | A | 1 | 1 | H | 0.45763 | 0.45874 | 1.5
I5 | A | 2 | 1 | H | 0.43401 | 0.41574 | 1.5
I5 | A | 3 | 1 | H | 0.40529 | 0.43829 | 1.5
I5 | A | 4 | 1 | H | 0.42914 | 0.47049 | 1.5
I5 | A | 5 | 1 | H | 0.44249 | 0.42053 | 1.5
I5 | A | 6 | 1 | H | 0.39351 | 0.40975 | 1.5
I5 | A | 7 | 1 | H | 0.4406 | 0.4405 | 1.5
I5 | A | 8 | 1 | H | 0.47688 | 0.48902 | 1.5
I5 | A | 9 | 1 | H | 0.51346 | 0.54983 | 1.5
I5 | A | 10 | 1 | H | 0.48083 | 0.48081 | 1.5
I5 | A | 11 | 1 | H | 0.41042 | 0.4155 | 1.5
I5 | A | 12 | 1 | H | 0.42759 | 0.43512 | 1.5
I5 | A | 13 | 1 | H | 0.49029 | 0.47194 | 1.5
I5 | A | 14 | 1 | H | 0.84466 | 0.48676 | 1.5
I5 | A | 15 | 1 | H | 0.3898 | 0.421 | 1.5
I5 | A | 16 | 1 | H | 0.48584 | 0.48195 | 1.5
I5 | A | 17 | 1 | H | 0.44684 | 0.46491 | 1.5
I5 | A | 18 | 1 | H | 0.44297 | 0.49015 | 1.5
I5 | A | 19 | 1 | H | 0.43348 | 0.46484 | 1.5
I5 | A | 20 | 1 | H | 0.4557 | 0.47226 | 1.5
I5 | A | 21 | 1 | H | 0.42661 | 0.44525 | 1.5
I5 | A | 22 | 1 | H | 0.5076 | 0.50534 | 1.5
I5 | A | 23 | 1 | H | 0.43653 | 0.45444 | 1.5
I5 | A | 24 | 1 | H | 0.51432 | 0.50625 | 1.5
I5 | A | 25 | 1 | H | 0.44614 | 0.44943 | 1.5
I5 | A | 26 | 1 | H | 0.54654 | 0.50527 | 1.5
I5 | A | 27 | 1 | H | 0.47983 | 0.48718 | 1.5
I5 | A | 28 | 1 | H | 0.55741 | 0.56658 | 1.5
I5 | A | 29 | 1 | H | 0.45354 | 0.46834 | 1.5
I5 | A | 30 | 1 | H | 0.461 | 0.46202 | 1.5
I5 | A | 31 | 1 | H | 0.4841 | 0.49207 | 1.5
I5 | A | 32 | 1 | H | 0.49727 | 0.51598 | 1.5
I5 | A | 1 | 2 | H | 0.45905 | 0.46695 | 1.5
I5 | A | 2 | 2 | H | 0.43167 | 0.42017 | 1.5
I5 | A | 3 | 2 | H | 0.40836 | 0.43403 | 1.5
I5 | A | 4 | 2 | H | 0.42953 | 0.46042 | 1.5
I5 | A | 5 | 2 | H | 0.43304 | 0.40939 | 1.5
I5 | A | 6 | 2 | H | 0.4139 | 0.40259 | 1.5
I5 | A | 7 | 2 | H | 0.43551 | 0.44071 | 1.5
I5 | A | 8 | 2 | H | 0.47544 | 0.48313 | 1.5
I5 | A | 9 | 2 | H | 0.52524 | 0.56605 | 1.5
I5 | A | 10 | 2 | H | 0.48879 | 0.47513 | 1.5
I5 | A | 11 | 2 | H | 0.41347 | 0.41097 | 1.5
I5 | A | 12 | 2 | H | 0.43287 | 0.43094 | 1.5
I5 | A | 13 | 2 | H | 0.50183 | 0.45494 | 1.5
I5 | A | 14 | 2 | H | 0.85452 | 0.49079 | 1.5
I5 | A | 15 | 2 | H | 0.38325 | 0.41706 | 1.5
I5 | A | 16 | 2 | H | 0.49443 | 0.48775 | 1.5
I5 | A | 17 | 2 | H | 0.45288 | 0.47584 | 1.5
I5 | A | 18 | 2 | H | 0.42404 | 0.47778 | 1.5
I5 | A | 19 | 2 | H | 0.43392 | 0.45412 | 1.5
I5 | A | 20 | 2 | H | 0.45531 | 0.46489 | 1.5
I5 | A | 21 | 2 | H | 0.41843 | 0.45337 | 1.5
I5 | A | 22 | 2 | H | 0.512 | 0.49917 | 1.5
I5 | A | 23 | 2 | H | 0.44064 | 0.45416 | 1.5
I5 | A | 24 | 2 | H | 0.5239 | 0.51634 | 1.5
I5 | A | 25 | 2 | H | 0.42948 | 0.45469 | 1.5
I5 | A | 26 | 2 | H | 0.53112 | 0.49595 | 1.5
I5 | A | 27 | 2 | H | 0.46657 | 0.47011 | 1.5
I5 | A | 28 | 2 | H | 0.55508 | 0.55678 | 1.5
I5 | A | 29 | 2 | H | 0.44667 | 0.4801 | 1.5
I5 | A | 30 | 2 | H | 0.45834 | 0.45784 | 1.5
I5 | A | 31 | 2 | H | 0.48812 | 0.47373 | 1.5
I5 | A | 32 | 2 | H | 0.4884 | 0.52665 | 1.5
I5 | B | 1 | 1 | H | 0.47702 | 0.45538 | 1.5
I5 | B | 2 | 1 | H | 0.45609 | 0.43304 | 1.5
I5 | B | 3 | 1 | H | 0.41921 | 0.43471 | 1.5
I5 | B | 4 | 1 | H | 0.44433 | 0.46544 | 1.5
I5 | B | 5 | 1 | H | 0.43762 | 0.4343 | 1.5
I5 | B | 6 | 1 | H | 0.39483 | 0.40848 | 1.5
I5 | B | 7 | 1 | H | 0.42404 | 0.42725 | 1.5
I5 | B | 8 | 1 | H | 0.4793 | 0.47678 | 1.5
I5 | B | 9 | 1 | H | 0.55668 | 0.54297 | 1.5
I5 | B | 10 | 1 | H | 0.48276 | 0.48263 | 1.5
I5 | B | 11 | 1 | H | 0.40826 | 0.43679 | 1.5
I5 | B | 12 | 1 | H | 0.42612 | 0.43038 | 1.5
I5 | B | 13 | 1 | H | 0.47823 | 0.46242 | 1.5
I5 | B | 14 | 1 | H | 0.83766 | 0.49375 | 1.5
I5 | B | 15 | 1 | H | 0.39875 | 0.41991 | 1.5
I5 | B | 16 | 1 | H | 0.48817 | 0.46981 | 1.5
I5 | B | 17 | 1 | H | 0.43547 | 0.47219 | 1.5
I5 | B | 18 | 1 | H | 0.42528 | 0.47834 | 1.5
I5 | B | 19 | 1 | H | 0.45303 | 0.45313 | 1.5
I5 | B | 20 | 1 | H | 0.43641 | 0.46933 | 1.5
I5 | B | 21 | 1 | H | 0.41923 | 0.44027 | 1.5
I5 | B | 22 | 1 | H | 0.50711 | 0.50536 | 1.5
I5 | B | 23 | 1 | H | 0.43086 | 0.45795 | 1.5
I5 | B | 24 | 1 | H | 0.52538 | 0.53501 | 1.5
I5 | B | 25 | 1 | H | 0.452 | 0.4527 | 1.5
I5 | B | 26 | 1 | H | 0.54696 | 0.49001 | 1.5
I5 | B | 27 | 1 | H | 0.45264 | 0.46968 | 1.5
I5 | B | 28 | 1 | H | 0.56245 | 0.55926 | 1.5
I5 | B | 29 | 1 | H | 0.45773 | 0.47341 | 1.5
I5 | B | 30 | 1 | H | 0.45543 | 0.47042 | 1.5
I5 | B | 31 | 1 | H | 0.4782 | 0.47418 | 1.5
I5 | B | 32 | 1 | H | 0.49272 | 0.52107 | 1.5
I5 | B | 1 | 2 | H | 0.45308 | 0.44375 | 1.5
I5 | B | 2 | 2 | H | 0.43378 | 0.42926 | 1.5
I5 | B | 3 | 2 | H | 0.40881 | 0.43141 | 1.5
I5 | B | 4 | 2 | H | 0.43348 | 0.46697 | 1.5
I5 | B | 5 | 2 | H | 0.44681 | 0.43575 | 1.5
I5 | B | 6 | 2 | H | 0.39867 | 0.4137 | 1.5
I5 | B | 7 | 2 | H | 0.42911 | 0.4312 | 1.5
I5 | B | 8 | 2 | H | 0.47979 | 0.46821 | 1.5
I5 | B | 9 | 2 | H | 0.54332 | 0.54897 | 1.5
I5 | B | 10 | 2 | H | 0.4939 | 0.4815 | 1.5
I5 | B | 11 | 2 | H | 0.41426 | 0.42508 | 1.5
I5 | B | 12 | 2 | H | 0.4382 | 0.42253 | 1.5
I5 | B | 13 | 2 | H | 0.4829 | 0.46379 | 1.5
I5 | B | 14 | 2 | H | 0.85 | 0.48369 | 1.5
I5 | B | 15 | 2 | H | 0.38394 | 0.42079 | 1.5
I5 | B | 16 | 2 | H | 0.49778 | 0.48793 | 1.5
I5 | B | 17 | 2 | H | 0.44758 | 0.47939 | 1.5
I5 | B | 18 | 2 | H | 0.42815 | 0.47439 | 1.5
I5 | B | 19 | 2 | H | 0.44875 | 0.45464 | 1.5
I5 | B | 20 | 2 | H | 0.44626 | 0.46265 | 1.5
I5 | B | 21 | 2 | H | 0.42659 | 0.44018 | 1.5
I5 | B | 22 | 2 | H | 0.49597 | 0.50245 | 1.5
I5 | B | 23 | 2 | H | 0.43371 | 0.44744 | 1.5
I5 | B | 24 | 2 | H | 0.52269 | 0.51871 | 1.5
I5 | B | 25 | 2 | H | 0.44159 | 0.46117 | 1.5
I5 | B | 26 | 2 | H | 0.55092 | 0.50617 | 1.5
I5 | B | 27 | 2 | H | 0.47314 | 0.48027 | 1.5
I5 | B | 28 | 2 | H | 0.54246 | 0.56147 | 1.5
I5 | B | 29 | 2 | H | 0.45149 | 0.48774 | 1.5
I5 | B | 30 | 2 | H | 0.45806 | 0.46401 | 1.5
I5 | B | 31 | 2 | H | 0.48644 | 0.47139 | 1.5
I5 | B | 32 | 2 | H | 0.49366 | 0.51912 | 1.5
M12 | A | 1 | 1 | H | 0.16405 | 0.17932 | 0.396
M12 | A | 2 | 1 | H | 0.15885 | 0.1842 | 0.396
M12 | A | 3 | 1 | H | 0.16154 | 0.17557 | 0.396
M12 | A | 4 | 1 | H | 0.15939 | 0.18137 | 0.396
M12 | A | 5 | 1 | H | 0.16305 | 0.19394 | 0.396
M12 | A | 6 | 1 | H | 0.16196 | 0.17812 | 0.396
M12 | A | 7 | 1 | H | 0.16527 | 0.18202 | 0.396
M12 | A | 8 | 1 | H | 0.16165 | 0.17456 | 0.396
M12 | A | 9 | 1 | H | 0.16089 | 0.17967 | 0.396
M12 | A | 10 | 1 | H | 0.16059 | 0.1827 | 0.396
M12 | A | 11 | 1 | H | 0.16422 | 0.19652 | 0.396
M12 | A | 12 | 1 | H | 0.15781 | 0.19123 | 0.396
M12 | A | 13 | 1 | H | 0.16555 | 0.19602 | 0.396
M12 | A | 14 | 1 | H | 0.15635 | 0.19251 | 0.396
M12 | A | 15 | 1 | H | 0.16755 | 0.20021 | 0.396
M12 | A | 16 | 1 | H | 0.15714 | 0.20773 | 0.396
M12 | B | 1 | 1 | H | 0.16357 | 0.17672 | 0.396
M12 | B | 2 | 1 | H | 0.15775 | 0.16824 | 0.396
M12 | B | 3 | 1 | H | 0.1656 | 0.17222 | 0.396
M12 | B | 4 | 1 | H | 0.1528 | 0.17327 | 0.396
M12 | B | 5 | 1 | H | 0.16406 | 0.17316 | 0.396
M12 | B | 6 | 1 | H | 0.15751 | 0.18038 | 0.396
M12 | B | 7 | 1 | H | 0.16319 | 0.18045 | 0.396
M12 | B | 8 | 1 | H | 0.15557 | 0.17197 | 0.396
M12 | B | 9 | 1 | H | 0.16145 | 0.19035 | 0.396
M12 | B | 10 | 1 | H | 0.16216 | 0.18997 | 0.396
M12 | B | 11 | 1 | H | 0.15945 | 0.18303 | 0.396
M12 | B | 12 | 1 | H | 0.16091 | 0.1901 | 0.396
M12 | B | 13 | 1 | H | 0.16132 | 0.1823 | 0.396
M12 | B | 14 | 1 | H | 0.15379 | 0.17809 | 0.396
M12 | B | 15 | 1 | H | 0.16175 | 0.1898 | 0.396
M12 | B | 16 | 1 | H | 0.16427 | 0.20333 | 0.396
M13 | A | 1 | 1 | H | 0.056433 | 0.05526 | 0.107
M13 | A | 2 | 1 | H | 0.058812 | 0.054587 | 0.107
M13 | A | 3 | 1 | H | 0.057138 | 0.053157 | 0.107
M13 | A | 4 | 1 | H | 0.058966 | 0.054112 | 0.107
M13 | A | 5 | 1 | H | 0.055753 | 0.053281 | 0.107
M13 | A | 6 | 1 | H | 0.057418 | 0.05395 | 0.107
M13 | A | 7 | 1 | H | 0.056223 | 0.053133 | 0.107
M13 | A | 8 | 1 | H | 0.057462 | 0.053677 | 0.107
M13 | A | 9 | 1 | H | 0.056589 | 0.055292 | 0.107
M13 | A | 10 | 1 | H | 0.058402 | 0.054368 | 0.107
M13 | A | 11 | 1 | H | 0.057867 | 0.053186 | 0.107
M13 | A | 12 | 1 | H | 0.058411 | 0.05496 | 0.107
M13 | A | 13 | 1 | H | 0.056342 | 0.0538 | 0.107
M13 | A | 14 | 1 | H | 0.058574 | 0.056675 | 0.107
M13 | A | 15 | 1 | H | 0.057546 | 0.054716 | 0.107
M13 | A | 16 | 1 | H | 0.05933 | 0.056613 | 0.107
M13 | B | 1 | 1 | H | 0.057144 | 0.054962 | 0.107
M13 | B | 2 | 1 | H | 0.05942 | 0.055557 | 0.107
M13 | B | 3 | 1 | H | 0.056182 | 0.053426 | 0.107
M13 | B | 4 | 1 | H | 0.058287 | 0.054578 | 0.107
M13 | B | 5 | 1 | H | 0.056176 | 0.052648 | 0.107
M13 | B | 6 | 1 | H | 0.057557 | 0.054914 | 0.107
M13 | B | 7 | 1 | H | 0.056373 | 0.052616 | 0.107
M13 | B | 8 | 1 | H | 0.057715 | 0.054096 | 0.107
M13 | B | 9 | 1 | H | 0.057896 | 0.05556 | 0.107
M13 | B | 10 | 1 | H | 0.058278 | 0.054221 | 0.107
M13 | B | 11 | 1 | H | 0.057587 | 0.05355 | 0.107
M13 | B | 12 | 1 | H | 0.058751 | 0.054793 | 0.107
M13 | B | 13 | 1 | H | 0.056341 | 0.053481 | 0.107
M13 | B | 14 | 1 | H | 0.058516 | 0.056323 | 0.107
M13 | B | 15 | 1 | H | 0.058514 | 0.054522 | 0.107
M13 | B | 16 | 1 | H | 0.059282 | 0.055235 | 0.107
M14 | A | 1 | 1 | H | 0.068539 | 0.073107 | 0.091
M14 | A | 2 | 1 | H | 0.068169 | 0.070994 | 0.091
M14 | A | 3 | 1 | H | 0.068305 | 0.071863 | 0.091
M14 | A | 4 | 1 | H | 0.077025 | 0.079409 | 0.091
M14 | A | 5 | 1 | H | 0.068182 | 0.070964 | 0.091
M14 | A | 6 | 1 | H | 0.067702 | 0.071424 | 0.091
M14 | A | 7 | 1 | H | 0.06697 | 0.070554 | 0.091
M14 | A | 8 | 1 | H | 0.069444 | 0.068913 | 0.091
M14 | A | 9 | 1 | H | 0.073814 | 0.074129 | 0.091
M14 | A | 10 | 1 | H | 0.068584 | 0.070218 | 0.091
M14 | A | 11 | 1 | H | 0.07194 | 0.074886 | 0.091
M14 | A | 12 | 1 | H | 0.070036 | 0.073697 | 0.091
M14 | A | 13 | 1 | H | 0.070016 | 0.073183 | 0.091
M14 | A | 14 | 1 | H | 0.071642 | 0.071506 | 0.091
M14 | A | 15 | 1 | H | 0.071084 | 0.074875 | 0.091
M14 | A | 16 | 1 | H | 0.07253 | 0.091697 | 0.091
M14 | B | 1 | 1 | H | 0.073246 | 0.072553 | 0.091
M14 | B | 2 | 1 | H | 0.06979 | 0.070617 | 0.091
M14 | B | 3 | 1 | H | 0.070178 | 0.072772 | 0.091
M14 | B | 4 | 1 | H | 0.079972 | 0.078101 | 0.091
M14 | B | 5 | 1 | H | 0.070416 | 0.069996 | 0.091
M14 | B | 6 | 1 | H | 0.067851 | 0.071996 | 0.091
M14 | B | 7 | 1 | H | 0.068111 | 0.070806 | 0.091
M14 | B | 8 | 1 | H | 0.071159 | 0.071705 | 0.091
M14 | B | 9 | 1 | H | 0.075281 | 0.075675 | 0.091
M14 | B | 10 | 1 | H | 0.068387 | 0.07091 | 0.091
M14 | B | 11 | 1 | H | 0.071087 | 0.074916 | 0.091
M14 | B | 12 | 1 | H | 0.069593 | 0.07241 | 0.091
M14 | B | 13 | 1 | H | 0.068224 | 0.071676 | 0.091
M14 | B | 14 | 1 | H | 0.071702 | 0.07244 | 0.091
M14 | B | 15 | 1 | H | 0.067568 | 0.072607 | 0.091
M14 | B | 16 | 1 | H | 0.073693 | 0.093509 | 0.091
M15 | A | 1 | 1 | H | 0.090738 | 0.05554 | 0.07
M15 | A | 2 | 1 | H | 0.084362 | 0.057886 | 0.07
M15 | A | 3 | 1 | H | 0.10547 | 0.055189 | 0.07
M15 | A | 4 | 1 | H | 0.096774 | 0.057422 | 0.07
M15 | A | 5 | 1 | H | 0.089373 | 0.054727 | 0.07
M15 | A | 6 | 1 | H | 0.095379 | 0.053819 | 0.07
M15 | A | 7 | 1 | H | 0.095589 | 0.054137 | 0.07
M15 | A | 8 | 1 | H | 0.09438 | 0.054275 | 0.07
M15 | A | 9 | 1 | H | 0.096677 | 0.054617 | 0.07
M15 | A | 10 | 1 | H | 0.092402 | 0.057179 | 0.07
M15 | A | 11 | 1 | H | 0.087038 | 0.053873 | 0.07
M15 | A | 12 | 1 | H | 0.081414 | 0.053164 | 0.07
M15 | A | 13 | 1 | H | 0.096318 | 0.054526 | 0.07
M15 | A | 14 | 1 | H | 0.089054 | 0.058236 | 0.07
M15 | A | 15 | 1 | H | 0.09297 | 0.055165 | 0.07
M15 | A | 16 | 1 | H | 0.092558 | 0.055784 | 0.07
M15 | B | 1 | 1 | H | 0.088935 | 0.057581 | 0.07
M15 | B | 2 | 1 | H | 0.089263 | 0.056713 | 0.07
M15 | B | 3 | 1 | H | 0.090302 | 0.053684 | 0.07
M15 | B | 4 | 1 | H | 0.090652 | 0.054067 | 0.07
M15 | B | 5 | 1 | H | 0.088334 | 0.053348 | 0.07
M15 | B | 6 | 1 | H | 0.090693 | 0.054728 | 0.07
M15 | B | 7 | 1 | H | 0.093151 | 0.055975 | 0.07
M15 | B | 8 | 1 | H | 0.092356 | 0.052939 | 0.07
M15 | B | 9 | 1 | H | 0.093484 | 0.053623 | 0.07
M15 | B | 10 | 1 | H | 0.097381 | 0.054487 | 0.07
M15 | B | 11 | 1 | H | 0.10279 | 0.05183 | 0.07
M15 | B | 12 | 1 | H | 0.089332 | 0.055346 | 0.07
M15 | B | 13 | 1 | H | 0.089142 | 0.056378 | 0.07
M15 | B | 14 | 1 | H | 0.094796 | 0.060708 | 0.07
M15 | B | 15 | 1 | H | 0.0895 | 0.060088 | 0.07
M15 | B | 16 | 1 | H | 0.094218 | 0.058842 | 0.07
M16A | A | 1 | 1 | H | 0.094973 | 0.14434 | 0.072
M16A | A | 2 | 1 | H | 0.099026 | 0.14093 | 0.072
M16A | A | 3 | 1 | H | 0.12632 | 0.14466 | 0.072
M16A | A | 4 | 1 | H | 0.1097 | 0.11986 | 0.072
M16A | A | 5 | 1 | H | 0.11403 | 0.12158 | 0.072
M16A | A | 6 | 1 | H | 0.081606 | 0.12595 | 0.072
M16A | A | 7 | 1 | H | 0.12072 | 0.12218 | 0.072
M16A | A | 8 | 1 | H | 0.082165 | 0.10687 | 0.072
M16A | A | 9 | 1 | H | 0.11849 | 0.10066 | 0.072
M16A | A | 10 | 1 | H | 0.10008 | 0.096204 | 0.072
M16A | A | 11 | 1 | H | 0.11918 | 0.10813 | 0.072
M16A | A | 12 | 1 | H | 0.12581 | 0.096514 | 0.072
M16A | A | 13 | 1 | H | 0.12623 | 0.094084 | 0.072
M16A | A | 14 | 1 | H | 0.12319 | 0.091272 | 0.072
M16A | A | 15 | 1 | H | 0.10499 | 0.09556 | 0.072
M16A | A | 16 | 1 | H | 0.10809 | 0.084084 | 0.072
M16A | B | 1 | 1 | H | 0.11688 | 0.15543 | 0.072
M16A | B | 2 | 1 | H | 0.10133 | 0.12231 | 0.072
M16A | B | 3 | 1 | H | 0.090472 | 0.11863 | 0.072
M16A | B | 4 | 1 | H | 0.10975 | 0.11174 | 0.072
M16A | B | 5 | 1 | H | 0.10831 | 0.11538 | 0.072
M16A | B | 6 | 1 | H | 0.086882 | 0.11678 | 0.072
M16A | B | 7 | 1 | H | 0.11824 | 0.1057 | 0.072
M16A | B | 8 | 1 | H | 0.087126 | 0.099103 | 0.072
M16A | B | 9 | 1 | H | 0.087393 | 0.1086 | 0.072
M16A | B | 10 | 1 | H | 0.094439 | 0.095194 | 0.072
M16A | B | 11 | 1 | H | 0.089135 | 0.09383 | 0.072
M16A | B | 12 | 1 | H | 0.14147 | 0.095761 | 0.072
M16A | B | 13 | 1 | H | 0.10166 | 0.091936 | 0.072
M16A | B | 14 | 1 | H | 0.10953 | 0.086338 | 0.072
M16A | B | 15 | 1 | H | 0.092931 | 0.082616 | 0.072
M16A | B | 16 | 1 | H | 0.11754 | 0.089285 | 0.072
M16B | A | 1 | 1 | H | 0.13348 | 0.13334 | 0.072
M16B | A | 2 | 1 | H | 0.16724 | 0.13098 | 0.072
M16B | A | 3 | 1 | H | 0.19243 | 0.11509 | 0.072
M16B | A | 4 | 1 | H | 0.1416 | 0.13536 | 0.072
M16B | A | 5 | 1 | H | 0.12055 | 0.1272 | 0.072
M16B | A | 6 | 1 | H | 0.13089 | 0.10907 | 0.072
M16B | A | 7 | 1 | H | 0.17114 | 0.11792 | 0.072
M16B | A | 8 | 1 | H | 0.1728 | 0.11374 | 0.072
M16B | A | 9 | 1 | H | 0.12489 | 0.11363 | 0.072
M16B | A | 10 | 1 | H | 0.15933 | 0.11712 | 0.072
M16B | A | 11 | 1 | H | 0.13665 | 0.09631 | 0.072
M16B | A | 12 | 1 | H | 0.12985 | 0.10236 | 0.072
M16B | A | 13 | 1 | H | 0.15864 | 0.095695 | 0.072
M16B | A | 14 | 1 | H | 0.16678 | 0.093102 | 0.072
M16B | A | 15 | 1 | H | 0.23512 | 0.094661 | 0.072
M16B | A | 16 | 1 | H | 0.14257 | 0.086237 | 0.072
M16B | B | 1 | 1 | H | 0.13392 | 0.12632 | 0.072
M16B | B | 2 | 1 | H | 0.14778 | 0.10845 | 0.072
M16B | B | 3 | 1 | H | 0.12969 | 0.11411 | 0.072
M16B | B | 4 | 1 | H | 0.17575 | 0.11076 | 0.072
M16B | B | 5 | 1 | H | 0.1286 | 0.14596 | 0.072
M16B | B | 6 | 1 | H | 0.15745 | 0.10952 | 0.072
M16B | B | 7 | 1 | H | 0.11764 | 0.11824 | 0.072
M16B | B | 8 | 1 | H | 0.149 | 0.097886 | 0.072
M16B | B | 9 | 1 | H | 0.1521 | 0.089453 | 0.072
M16B | B | 10 | 1 | H | 0.14563 | 0.092121 | 0.072
M16B | B | 11 | 1 | H | 0.1627 | 0.091262 | 0.072
M16B | B | 12 | 1 | H | 0.1511 | 0.082015 | 0.072
M16B | B | 13 | 1 | H | 0.16966 | 0.088731 | 0.072
M16B | B | 14 | 1 | H | 0.14314 | 0.077643 | 0.072
M16B | B | 15 | 1 | H | 0.15885 | 0.086735 | 0.072
M16B | B | 16 | 1 | H | 0.16088 | 0.079405 | 0.072
M13 | A | 1 | 1 | L | 0.27688 | 0.27618 | 0.423
M13 | A | 2 | 1 | L | 0.22197 | 0.32846 | 0.423
M13 | A | 3 | 1 | L | 0.27108 | 0.2722 | 0.423
M13 | A | 4 | 1 | L | 0.25927 | 0.2133 | 0.423
M13 | A | 5 | 1 | L | 0.25042 | 0.2527 | 0.423
M13 | A | 6 | 1 | L | 0.23799 | 0.27309 | 0.423
M13 | A | 7 | 1 | L | 0.23892 | 0.35041 | 0.423
M13 | A | 8 | 1 | L | 0.27168 | 0.28434 | 0.423
M13 | A | 9 | 1 | L | 0.29894 | 0.33428 | 0.423
M13 | A | 10 | 1 | L | 0.20472 | 0.27199 | 0.423
M13 | A | 11 | 1 | L | 0.3231 | 0.27043 | 0.423
M13 | A | 12 | 1 | L | 0.31241 | 0.31543 | 0.423
M13 | A | 13 | 1 | L | 0.26929 | 0.25462 | 0.423
M13 | A | 14 | 1 | L | 0.28234 | 0.29161 | 0.423
M13 | A | 15 | 1 | L | 0.21806 | 0.29762 | 0.423
M13 | A | 16 | 1 | L | 0.27176 | 0.29594 | 0.423
M13 | B | 1 | 1 | L | 0.26923 | 0.28093 | 0.423
M13 | B | 2 | 1 | L | 0.21106 | 0.32108 | 0.423
M13 | B | 3 | 1 | L | 0.2697 | 0.24381 | 0.423
M13 | B | 4 | 1 | L | 0.25627 | 0.21268 | 0.423
M13 | B | 5 | 1 | L | 0.24912 | 0.25336 | 0.423
M13 | B | 6 | 1 | L | 0.25191 | 0.27358 | 0.423
M13 | B | 7 | 1 | L | 0.23839 | 0.34467 | 0.423
M13 | B | 8 | 1 | L | 0.28544 | 0.29099 | 0.423
M13 | B | 9 | 1 | L | 0.29279 | 0.33567 | 0.423
M13 | B | 10 | 1 | L | 0.21918 | 0.28049 | 0.423
M13 | B | 11 | 1 | L | 0.31952 | 0.25865 | 0.423
M13 | B | 12 | 1 | L | 0.30903 | 0.29992 | 0.423
M13 | B | 13 | 1 | L | 0.27483 | 0.24175 | 0.423
M13 | B | 14 | 1 | L | 0.25856 | 0.25966 | 0.423
M13 | B | 15 | 1 | L | 0.22458 | 0.28452 | 0.423
M13 | B | 16 | 1 | L | 0.28809 | 0.31908 | 0.423

## RSB sensitivity
Band | HAM | Subsample | Gain State | Band Median |  | Band Average |  | SNR at Ltyp Spec
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | SNR at Ltyp | SNR at Ltyp | SNR at Ltyp | SNR at Ltyp
I1 | A | 1 | H | 161.4 | 158.905 | 173.6694 | 164.5559 | 119
I1 | A | 2 | H | 166.425 | 176.735 | 178.0337 | 176.1191 | 119
I1 | B | 1 | H | 172.155 | 164.8 | 193.8725 | 172.2275 | 119
I1 | B | 2 | H | 175.44 | 176.465 | 193.0191 | 186.1253 | 119
I2 | A | 1 | H | 241.07 | 237.285 | 239.9019 | 237.0191 | 150
I2 | A | 2 | H | 240.275 | 235.245 | 241.5934 | 236.5622 | 150
I2 | B | 1 | H | 235.96 | 236.34 | 239.1919 | 237.1269 | 150
I2 | B | 2 | H | 240.12 | 235.11 | 238.9844 | 236.1147 | 150
I3 | A | 1 | H | 174.815 | 175.515 | 170.0081 | 171.1716 | 6
I3 | B | 2 | H | 176.35 | 174.32 | 170.6037 | 170.0047 | 6
I3 | A | 1 | H | 174.655 | 176.005 | 169.7881 | 170.7266 | 6
I3 | B | 2 | H | 176.205 | 176.55 | 170.0953 | 171.7547 | 6
M1 | A | 1 | H | 599.07 | 590.515 | 604.5894 | 589.0838 | 352
M1 | B | 1 | H | 603.39 | 578.2 | 603.5319 | 578.8369 | 352
M2 | A | 1 | H | 538.115 | 525.87 | 539.6269 | 525.0013 | 380
M2 | B | 1 | H | 533.045 | 510.305 | 534.2269 | 510.3531 | 380
M3 | A | 1 | H | 672.065 | 641.615 | 667.9644 | 644.9244 | 416
M3 | B | 1 | H | 668.395 | 646.3 | 666.5913 | 647.345 | 416
M4 | A | 1 | H | 521.9 | 505.525 | 516.0675 | 504.83 | 362
M4 | B | 1 | H | 519.175 | 506.455 | 517.2412 | 509.9363 | 362
M5 | A | 1 | H | 360.255 | 349.575 | 360.0613 | 350.5056 | 242
M5 | B | 1 | H | 362.045 | 347.07 | 361.1812 | 342.0206 | 242
M6 | A | 1 | H | 376.745 | 376.025 | 375.8231 | 375.7244 | 199
M6 | B | 1 | H | 380.65 | 377.645 | 380.9375 | 376.8413 | 199
M7 | A | 1 | H | 513.1 | 522.245 | 513.3763 | 522.2713 | 215
M7 | B | 1 | H | 509.265 | 512.56 | 510.2906 | 518.7531 | 215
M8 | A | 1 | H | 291.225 | 261.965 | 300.8781 | 277.095 | 74
M8 | B | 1 | H | 276.75 | 305.16 | 285.7513 | 298.8375 | 74
M9 | A | 1 | H | 248.25 | 230.15 | 247.0444 | 232.1406 | 83
M9 | B | 1 | H | 249.305 | 233.58 | 248.7556 | 233.3263 | 83
M10 | A | 1 | H | 629.21 | 639.76 | 645.5725 | 643.6175 | 342
M10 | B | 1 | H | 636.055 | 637.81 | 646.2581 | 671.0912 | 342
M11 | A | 1 | H | 58.54 | 44.84 | 54.7794 | 62.5481 | 10
M11 | B | 1 | H | 51.615 | 56.61 | 53.2475 | 55.6675 | 10
M1 | A | 1 | L | 1020.54 | 1002.31 | 1026.4519 | 990.5281 | 316
M1 | B | 1 | L | 1017.735 | 968.565 | 1016.5675 | 961.64 | 316
M2 | A | 1 | L | 927.61 | 905.065 | 930.6319 | 914.7256 | 409
M2 | B | 1 | L | 934.045 | 909.32 | 942.0212 | 907.0212 | 409
M3 | A | 1 | L | 1029.485 | 1065.66 | 1026.4425 | 1056.115 | 414
M3 | B | 1 | L | 1005.38 | 1045.785 | 986.0875 | 1058.3756 | 414
M4 | A | 1 | L | 782.83 | 770 | 786.8575 | 776.1294 | 315
M4 | B | 1 | L | 783.15 | 767.215 | 780.2525 | 768.2894 | 315
M5 | A | 1 | L | 698.085 | 677.49 | 719.3594 | 685.0113 | 360
M5 | B | 1 | L | 700.725 | 677.245 | 702.0119 | 714.5531 | 360
M7 | A | 1 | L | 684.88 | 655.34 | 698.1425 | 683.0169 | 340
M7 | B | 1 | L | 691.075 | 661.885 | 685.6788 | 658.77 | 340
Band | HAM | Detector | Subsample | Gain State | Detector Dependent |  | SNR at Ltyp Spec |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 RSB radiometric sensitivity.
[5 empty] | SNR at Ltyp | SNR at Ltyp |  |  | 2. The results were derived from RC-1 using the SIS-100 source (ScMA for M1, M2, and M3 low gain).
I1 | A | 1 | 1 | H | 166.15 | 213.57 | 119 |  | 3. Fits of the SIS-100 radiance vs SNR were used to determine the radiometric sensitivity.
I1 | A | 2 | 1 | H | 135.24 | 164.61 | 119 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | A | 3 | 1 | H | 206.54 | 139.78 | 119 |  | 5. The test was set to operational day mode.
I1 | A | 4 | 1 | H | 153.66 | 156.37 | 119
I1 | A | 5 | 1 | H | 192.32 | 194.68 | 119
I1 | A | 6 | 1 | H | 242.17 | 167.12 | 119
I1 | A | 7 | 1 | H | 161.46 | 198.92 | 119
I1 | A | 8 | 1 | H | 189.11 | 153.62 | 119
I1 | A | 9 | 1 | H | 304.03 | 166.05 | 119
I1 | A | 10 | 1 | H | 154.92 | 193.24 | 119
I1 | A | 11 | 1 | H | 143.7 | 162.15 | 119
I1 | A | 12 | 1 | H | 182.03 | 209.91 | 119
I1 | A | 13 | 1 | H | 186.45 | 125.81 | 119
I1 | A | 14 | 1 | H | 152.25 | 127.97 | 119
I1 | A | 15 | 1 | H | 142.66 | 144 | 119
I1 | A | 16 | 1 | H | 159.97 | 135.67 | 119
I1 | A | 17 | 1 | H | 157.66 | 135.19 | 119
I1 | A | 18 | 1 | H | 168.78 | 171.62 | 119
I1 | A | 19 | 1 | H | 171.68 | 172.06 | 119
I1 | A | 20 | 1 | H | 222.27 | 161.44 | 119
I1 | A | 21 | 1 | H | 156.7 | 151.7 | 119
I1 | A | 22 | 1 | H | 197.04 | 266.29 | 119
I1 | A | 23 | 1 | H | 155.01 | 341.48 | 119
I1 | A | 24 | 1 | H | 156.51 | 123.9 | 119
I1 | A | 25 | 1 | H | 133.14 | 168.02 | 119
I1 | A | 26 | 1 | H | 181.84 | 127.76 | 119
I1 | A | 27 | 1 | H | 133.4 | 131.93 | 119
I1 | A | 28 | 1 | H | 161.34 | 130.02 | 119
I1 | A | 29 | 1 | H | 156.26 | 135.11 | 119
I1 | A | 30 | 1 | H | 207.68 | 163.87 | 119
I1 | A | 31 | 1 | H | 187.61 | 84.68 | 119
I1 | A | 32 | 1 | H | 137.84 | 147.25 | 119
I1 | A | 1 | 2 | H | 169.6 | 173.81 | 119
I1 | A | 2 | 2 | H | 170.24 | 137.18 | 119
I1 | A | 3 | 2 | H | 140.45 | 180.72 | 119
I1 | A | 4 | 2 | H | 159.47 | 174.12 | 119
I1 | A | 5 | 2 | H | 153.77 | 176.18 | 119
I1 | A | 6 | 2 | H | 130.86 | 186.82 | 119
I1 | A | 7 | 2 | H | 126.1 | 120.84 | 119
I1 | A | 8 | 2 | H | 200.86 | 194.57 | 119
I1 | A | 9 | 2 | H | 158.26 | 137.09 | 119
I1 | A | 10 | 2 | H | 169.41 | 141.71 | 119
I1 | A | 11 | 2 | H | 180.9 | 138.8 | 119
I1 | A | 12 | 2 | H | 201.95 | 169.57 | 119
I1 | A | 13 | 2 | H | 232.61 | 148.02 | 119
I1 | A | 14 | 2 | H | 161.29 | 308.35 | 119
I1 | A | 15 | 2 | H | 160.57 | 177.29 | 119
I1 | A | 16 | 2 | H | 294.86 | 155.39 | 119
I1 | A | 17 | 2 | H | 296.18 | 130.42 | 119
I1 | A | 18 | 2 | H | 171.5 | 194.32 | 119
I1 | A | 19 | 2 | H | 157.97 | 236.16 | 119
I1 | A | 20 | 2 | H | 128.78 | 131.92 | 119
I1 | A | 21 | 2 | H | 149.37 | 180.9 | 119
I1 | A | 22 | 2 | H | 247.08 | 285.73 | 119
I1 | A | 23 | 2 | H | 144.3 | 182.77 | 119
I1 | A | 24 | 2 | H | 157.47 | 221.85 | 119
I1 | A | 25 | 2 | H | 221.7 | 181.67 | 119
I1 | A | 26 | 2 | H | 174.51 | 143.1 | 119
I1 | A | 27 | 2 | H | 149.54 | 184.05 | 119
I1 | A | 28 | 2 | H | 196.97 | 183.14 | 119
I1 | A | 29 | 2 | H | 182.54 | 131.74 | 119
I1 | A | 30 | 2 | H | 163.44 | 203.2 | 119
I1 | A | 31 | 2 | H | 151.33 | 199.92 | 119
I1 | A | 32 | 2 | H | 193.2 | 124.46 | 119
I1 | B | 1 | 1 | H | 318.26 | 224.65 | 119
I1 | B | 2 | 1 | H | 188.92 | 231.46 | 119
I1 | B | 3 | 1 | H | 161.71 | 146.77 | 119
I1 | B | 4 | 1 | H | 156.03 | 144.7 | 119
I1 | B | 5 | 1 | H | 155.31 | 164.97 | 119
I1 | B | 6 | 1 | H | 166.84 | 140.24 | 119
I1 | B | 7 | 1 | H | 136.46 | 184.76 | 119
I1 | B | 8 | 1 | H | 129.98 | 154.59 | 119
I1 | B | 9 | 1 | H | 207.62 | 152.59 | 119
I1 | B | 10 | 1 | H | 157.65 | 187.6 | 119
I1 | B | 11 | 1 | H | 279.62 | 213.45 | 119
I1 | B | 12 | 1 | H | 274.78 | 220.05 | 119
I1 | B | 13 | 1 | H | 157.71 | 155.7 | 119
I1 | B | 14 | 1 | H | 153.39 | 127.47 | 119
I1 | B | 15 | 1 | H | 196.19 | 135.04 | 119
I1 | B | 16 | 1 | H | 159.76 | 161.95 | 119
I1 | B | 17 | 1 | H | 345.87 | 210.91 | 119
I1 | B | 18 | 1 | H | 200.53 | 188.41 | 119
I1 | B | 19 | 1 | H | 177.47 | 205.46 | 119
I1 | B | 20 | 1 | H | 157.48 | 197.01 | 119
I1 | B | 21 | 1 | H | 195.64 | 176.49 | 119
I1 | B | 22 | 1 | H | 161.56 | 152.26 | 119
I1 | B | 23 | 1 | H | 153.46 | 131.27 | 119
I1 | B | 24 | 1 | H | 157.64 | 182.55 | 119
I1 | B | 25 | 1 | H | 205.51 | 146.7 | 119
I1 | B | 26 | 1 | H | 198.23 | 170.42 | 119
I1 | B | 27 | 1 | H | 208.47 | 192.69 | 119
I1 | B | 28 | 1 | H | 157.55 | 112.88 | 119
I1 | B | 29 | 1 | H | 193.21 | 157.96 | 119
I1 | B | 30 | 1 | H | 301.88 | 132.27 | 119
I1 | B | 31 | 1 | H | 244.17 | 164.63 | 119
I1 | B | 32 | 1 | H | 145.02 | 243.38 | 119
I1 | B | 1 | 2 | H | 174.77 | 161.63 | 119
I1 | B | 2 | 2 | H | 234.81 | 218.81 | 119
I1 | B | 3 | 2 | H | 234.23 | 204.18 | 119
I1 | B | 4 | 2 | H | 139.25 | 357.05 | 119
I1 | B | 5 | 2 | H | 153.54 | 217.55 | 119
I1 | B | 6 | 2 | H | 178.18 | 124.03 | 119
I1 | B | 7 | 2 | H | 198.12 | 149.48 | 119
I1 | B | 8 | 2 | H | 187.74 | 122.55 | 119
I1 | B | 9 | 2 | H | 195.82 | 245.88 | 119
I1 | B | 10 | 2 | H | 152.73 | 206.68 | 119
I1 | B | 11 | 2 | H | 359.32 | 288.93 | 119
I1 | B | 12 | 2 | H | 247.93 | 209.2 | 119
I1 | B | 13 | 2 | H | 314.64 | 239.74 | 119
I1 | B | 14 | 2 | H | 165.71 | 144.97 | 119
I1 | B | 15 | 2 | H | 125.54 | 119.06 | 119
I1 | B | 16 | 2 | H | 118.21 | 174.4 | 119
I1 | B | 17 | 2 | H | 158.91 | 207.74 | 119
I1 | B | 18 | 2 | H | 158.84 | 152.06 | 119
I1 | B | 19 | 2 | H | 230.17 | 158.81 | 119
I1 | B | 20 | 2 | H | 157.19 | 154.29 | 119
I1 | B | 21 | 2 | H | 132.33 | 162.39 | 119
I1 | B | 22 | 2 | H | 120.6 | 204.58 | 119
I1 | B | 23 | 2 | H | 220.65 | 151.2 | 119
I1 | B | 24 | 2 | H | 138.97 | 158.35 | 119
I1 | B | 25 | 2 | H | 121.6 | 287.96 | 119
I1 | B | 26 | 2 | H | 176.11 | 180.12 | 119
I1 | B | 27 | 2 | H | 445.33 | 178.53 | 119
I1 | B | 28 | 2 | H | 136.72 | 180.24 | 119
I1 | B | 29 | 2 | H | 177.53 | 136.73 | 119
I1 | B | 30 | 2 | H | 172.94 | 192.22 | 119
I1 | B | 31 | 2 | H | 197.86 | 130.41 | 119
I1 | B | 32 | 2 | H | 250.32 | 136.24 | 119
I2 | A | 1 | 1 | H | 246.55 | 246.09 | 150
I2 | A | 2 | 1 | H | 244.94 | 231.69 | 150
I2 | A | 3 | 1 | H | 239.16 | 235.15 | 150
I2 | A | 4 | 1 | H | 236.19 | 232.05 | 150
I2 | A | 5 | 1 | H | 248.39 | 234.3 | 150
I2 | A | 6 | 1 | H | 246.54 | 246.08 | 150
I2 | A | 7 | 1 | H | 248.55 | 238.85 | 150
I2 | A | 8 | 1 | H | 247.02 | 246.2 | 150
I2 | A | 9 | 1 | H | 242.26 | 238.14 | 150
I2 | A | 10 | 1 | H | 253.01 | 240.41 | 150
I2 | A | 11 | 1 | H | 235.9 | 238.65 | 150
I2 | A | 12 | 1 | H | 238.38 | 241.49 | 150
I2 | A | 13 | 1 | H | 244.22 | 232.47 | 150
I2 | A | 14 | 1 | H | 247.05 | 232.4 | 150
I2 | A | 15 | 1 | H | 234.09 | 240.29 | 150
I2 | A | 16 | 1 | H | 234.84 | 227.65 | 150
I2 | A | 17 | 1 | H | 227.23 | 244.98 | 150
I2 | A | 18 | 1 | H | 231.28 | 228.59 | 150
I2 | A | 19 | 1 | H | 244.96 | 251.05 | 150
I2 | A | 20 | 1 | H | 247.78 | 230.42 | 150
I2 | A | 21 | 1 | H | 232.67 | 243.54 | 150
I2 | A | 22 | 1 | H | 224.58 | 247.75 | 150
I2 | A | 23 | 1 | H | 245.86 | 232.45 | 150
I2 | A | 24 | 1 | H | 244.16 | 235.01 | 150
I2 | A | 25 | 1 | H | 228.91 | 228.13 | 150
I2 | A | 26 | 1 | H | 239.88 | 242.59 | 150
I2 | A | 27 | 1 | H | 242.32 | 240.13 | 150
I2 | A | 28 | 1 | H | 246.44 | 242.63 | 150
I2 | A | 29 | 1 | H | 234.21 | 228.67 | 150
I2 | A | 30 | 1 | H | 234.86 | 236.43 | 150
I2 | A | 31 | 1 | H | 237.55 | 221.7 | 150
I2 | A | 32 | 1 | H | 227.08 | 228.63 | 150
I2 | A | 1 | 2 | H | 232.62 | 235.55 | 150
I2 | A | 2 | 2 | H | 258.24 | 249.36 | 150
I2 | A | 3 | 2 | H | 246.9 | 233.13 | 150
I2 | A | 4 | 2 | H | 250.24 | 230.38 | 150
I2 | A | 5 | 2 | H | 239.32 | 234.94 | 150
I2 | A | 6 | 2 | H | 238.12 | 252.77 | 150
I2 | A | 7 | 2 | H | 247.39 | 241.21 | 150
I2 | A | 8 | 2 | H | 243.54 | 252.97 | 150
I2 | A | 9 | 2 | H | 239.97 | 236.42 | 150
I2 | A | 10 | 2 | H | 237.33 | 250.89 | 150
I2 | A | 11 | 2 | H | 239.87 | 246.34 | 150
I2 | A | 12 | 2 | H | 251.34 | 236.5 | 150
I2 | A | 13 | 2 | H | 242.14 | 233.01 | 150
I2 | A | 14 | 2 | H | 240.58 | 245 | 150
I2 | A | 15 | 2 | H | 237.62 | 238.39 | 150
I2 | A | 16 | 2 | H | 241.49 | 243.65 | 150
I2 | A | 17 | 2 | H | 257.88 | 225.12 | 150
I2 | A | 18 | 2 | H | 245.21 | 233.3 | 150
I2 | A | 19 | 2 | H | 234.48 | 233.22 | 150
I2 | A | 20 | 2 | H | 239.02 | 240.41 | 150
I2 | A | 21 | 2 | H | 241.86 | 227.75 | 150
I2 | A | 22 | 2 | H | 238.42 | 232.69 | 150
I2 | A | 23 | 2 | H | 233.11 | 235.96 | 150
I2 | A | 24 | 2 | H | 250.48 | 238.31 | 150
I2 | A | 25 | 2 | H | 242.07 | 234.36 | 150
I2 | A | 26 | 2 | H | 235.54 | 229.46 | 150
I2 | A | 27 | 2 | H | 235.11 | 236.38 | 150
I2 | A | 28 | 2 | H | 233.17 | 231.98 | 150
I2 | A | 29 | 2 | H | 250.81 | 224.16 | 150
I2 | A | 30 | 2 | H | 245.25 | 231.74 | 150
I2 | A | 31 | 2 | H | 227.61 | 225.11 | 150
I2 | A | 32 | 2 | H | 234.26 | 229.53 | 150
I2 | B | 1 | 1 | H | 251.25 | 250.27 | 150
I2 | B | 2 | 1 | H | 239.39 | 232.26 | 150
I2 | B | 3 | 1 | H | 236.02 | 236.62 | 150
I2 | B | 4 | 1 | H | 236.92 | 235.97 | 150
I2 | B | 5 | 1 | H | 268.76 | 237.58 | 150
I2 | B | 6 | 1 | H | 241.59 | 246.19 | 150
I2 | B | 7 | 1 | H | 256.52 | 239.05 | 150
I2 | B | 8 | 1 | H | 239.83 | 241.44 | 150
I2 | B | 9 | 1 | H | 258.61 | 240.64 | 150
I2 | B | 10 | 1 | H | 255.9 | 247.97 | 150
I2 | B | 11 | 1 | H | 233.35 | 233.59 | 150
I2 | B | 12 | 1 | H | 235.86 | 238.91 | 150
I2 | B | 13 | 1 | H | 251.39 | 233.05 | 150
I2 | B | 14 | 1 | H | 239.05 | 227.06 | 150
I2 | B | 15 | 1 | H | 225.83 | 246.64 | 150
I2 | B | 16 | 1 | H | 232.06 | 236.06 | 150
I2 | B | 17 | 1 | H | 234.71 | 239.51 | 150
I2 | B | 18 | 1 | H | 230.01 | 231.66 | 150
I2 | B | 19 | 1 | H | 241.47 | 238.13 | 150
I2 | B | 20 | 1 | H | 230.59 | 231.57 | 150
I2 | B | 21 | 1 | H | 235.9 | 246.07 | 150
I2 | B | 22 | 1 | H | 230.05 | 248.16 | 150
I2 | B | 23 | 1 | H | 237.47 | 227.26 | 150
I2 | B | 24 | 1 | H | 230.75 | 230.11 | 150
I2 | B | 25 | 1 | H | 223.18 | 237.97 | 150
I2 | B | 26 | 1 | H | 226.1 | 228.17 | 150
I2 | B | 27 | 1 | H | 233.45 | 251.75 | 150
I2 | B | 28 | 1 | H | 250.51 | 232.14 | 150
I2 | B | 29 | 1 | H | 233.51 | 232.36 | 150
I2 | B | 30 | 1 | H | 232.16 | 233.21 | 150
I2 | B | 31 | 1 | H | 257.82 | 226.54 | 150
I2 | B | 32 | 1 | H | 224.13 | 230.15 | 150
I2 | B | 1 | 2 | H | 239.11 | 232.78 | 150
I2 | B | 2 | 2 | H | 244.08 | 242.78 | 150
I2 | B | 3 | 2 | H | 236.21 | 233.97 | 150
I2 | B | 4 | 2 | H | 247.89 | 236.07 | 150
I2 | B | 5 | 2 | H | 241.57 | 235.79 | 150
I2 | B | 6 | 2 | H | 237.16 | 244.55 | 150
I2 | B | 7 | 2 | H | 234.46 | 244.54 | 150
I2 | B | 8 | 2 | H | 242.43 | 252.33 | 150
I2 | B | 9 | 2 | H | 235.52 | 233.32 | 150
I2 | B | 10 | 2 | H | 244.65 | 245.12 | 150
I2 | B | 11 | 2 | H | 236.18 | 240.17 | 150
I2 | B | 12 | 2 | H | 242.12 | 236.12 | 150
I2 | B | 13 | 2 | H | 241.87 | 231.41 | 150
I2 | B | 14 | 2 | H | 244.47 | 257.61 | 150
I2 | B | 15 | 2 | H | 246.99 | 236.13 | 150
I2 | B | 16 | 2 | H | 244.41 | 237.34 | 150
I2 | B | 17 | 2 | H | 252.54 | 233.41 | 150
I2 | B | 18 | 2 | H | 230.36 | 232.51 | 150
I2 | B | 19 | 2 | H | 241.13 | 240.62 | 150
I2 | B | 20 | 2 | H | 237.02 | 234.43 | 150
I2 | B | 21 | 2 | H | 235.76 | 233.34 | 150
I2 | B | 22 | 2 | H | 244.49 | 229.75 | 150
I2 | B | 23 | 2 | H | 242.24 | 225.29 | 150
I2 | B | 24 | 2 | H | 245.88 | 232.84 | 150
I2 | B | 25 | 2 | H | 232.9 | 229.33 | 150
I2 | B | 26 | 2 | H | 225.4 | 239.33 | 150
I2 | B | 27 | 2 | H | 227.26 | 242.01 | 150
I2 | B | 28 | 2 | H | 234.07 | 227.67 | 150
I2 | B | 29 | 2 | H | 241.91 | 226.01 | 150
I2 | B | 30 | 2 | H | 236.89 | 231.55 | 150
I2 | B | 31 | 2 | H | 233.59 | 221.48 | 150
I2 | B | 32 | 2 | H | 226.94 | 236.07 | 150
I3 | A | 1 | 1 | H | 167.79 | 171.36 | 6
I3 | A | 2 | 1 | H | 176.7 | 173.66 | 6
I3 | A | 3 | 1 | H | 174.64 | 179 | 6
I3 | A | 4 | 1 | H | 7 | 4.92 | 6
I3 | A | 5 | 1 | H | 183.14 | 175.91 | 6
I3 | A | 6 | 1 | H | 176.68 | 180.83 | 6
I3 | A | 7 | 1 | H | 175.32 | 176.53 | 6
I3 | A | 8 | 1 | H | 175.38 | 177.41 | 6
I3 | A | 9 | 1 | H | 187.47 | 183.63 | 6
I3 | A | 10 | 1 | H | 178.67 | 182.78 | 6
I3 | A | 11 | 1 | H | 178.51 | 173.16 | 6
I3 | A | 12 | 1 | H | 176.77 | 185.07 | 6
I3 | A | 13 | 1 | H | 184.21 | 183.86 | 6
I3 | A | 14 | 1 | H | 174.02 | 174.7 | 6
I3 | A | 15 | 1 | H | 170.34 | 177.41 | 6
I3 | A | 16 | 1 | H | 173.94 | 178.63 | 6
I3 | A | 17 | 1 | H | 172.9 | 174.14 | 6
I3 | A | 18 | 1 | H | 167.23 | 172.22 | 6
I3 | A | 19 | 1 | H | 174.99 | 171.7 | 6
I3 | A | 20 | 1 | H | 180.51 | 175.68 | 6
I3 | A | 21 | 1 | H | 181.9 | 179.39 | 6
I3 | A | 22 | 1 | H | 178.81 | 169.88 | 6
I3 | A | 23 | 1 | H | 171.51 | 175.35 | 6
I3 | A | 24 | 1 | H | 165.35 | 174.63 | 6
I3 | A | 25 | 1 | H | 179.31 | 179.69 | 6
I3 | A | 26 | 1 | H | 165.02 | 174.08 | 6
I3 | A | 27 | 1 | H | 180 | 181.55 | 6
I3 | A | 28 | 1 | H | 168.19 | 175.03 | 6
I3 | A | 29 | 1 | H | 174.41 | 172.9 | 6
I3 | A | 30 | 1 | H | 173.49 | 171.39 | 6
I3 | A | 31 | 1 | H | 172.29 | 174.64 | 6
I3 | A | 32 | 1 | H | 173.77 | 176.36 | 6
I3 | A | 1 | 2 | H | 169.35 | 172.33 | 6
I3 | A | 2 | 2 | H | 176.97 | 173.9 | 6
I3 | A | 3 | 2 | H | 168.66 | 177.72 | 6
I3 | A | 4 | 2 | H | 6.4 | 4.66 | 6
I3 | A | 5 | 2 | H | 177.46 | 175.83 | 6
I3 | A | 6 | 2 | H | 177.69 | 182.73 | 6
I3 | A | 7 | 2 | H | 171.43 | 177.64 | 6
I3 | A | 8 | 2 | H | 178.64 | 178.61 | 6
I3 | A | 9 | 2 | H | 190.38 | 181.76 | 6
I3 | A | 10 | 2 | H | 181.7 | 182 | 6
I3 | A | 11 | 2 | H | 178.57 | 180.61 | 6
I3 | A | 12 | 2 | H | 176.31 | 180.98 | 6
I3 | A | 13 | 2 | H | 180.64 | 174.61 | 6
I3 | A | 14 | 2 | H | 172.96 | 179.81 | 6
I3 | A | 15 | 2 | H | 181.58 | 181.62 | 6
I3 | A | 16 | 2 | H | 175.37 | 171.96 | 6
I3 | A | 17 | 2 | H | 178.92 | 172.28 | 6
I3 | A | 18 | 2 | H | 172.08 | 172 | 6
I3 | A | 19 | 2 | H | 181.49 | 170.05 | 6
I3 | A | 20 | 2 | H | 176.78 | 174.27 | 6
I3 | A | 21 | 2 | H | 178.22 | 173.37 | 6
I3 | A | 22 | 2 | H | 176.11 | 171.94 | 6
I3 | A | 23 | 2 | H | 173.42 | 177.74 | 6
I3 | A | 24 | 2 | H | 175.46 | 173.25 | 6
I3 | A | 25 | 2 | H | 176.43 | 178.67 | 6
I3 | A | 26 | 2 | H | 171.55 | 174.37 | 6
I3 | A | 27 | 2 | H | 180.23 | 177.55 | 6
I3 | A | 28 | 2 | H | 169.7 | 170.73 | 6
I3 | A | 29 | 2 | H | 176.39 | 172.57 | 6
I3 | A | 30 | 2 | H | 171.73 | 164.77 | 6
I3 | A | 31 | 2 | H | 168.92 | 168.42 | 6
I3 | A | 32 | 2 | H | 167.78 | 171.4 | 6
I3 | A | 1 | 1 | H | 175.53 | 172.16 | 6
I3 | A | 2 | 1 | H | 176.63 | 179.74 | 6
I3 | B | 3 | 1 | H | 171.01 | 179.78 | 6
I3 | B | 4 | 1 | H | 7.08 | 5.12 | 6
I3 | B | 5 | 1 | H | 177.05 | 176.35 | 6
I3 | B | 6 | 1 | H | 184.17 | 177.71 | 6
I3 | B | 7 | 1 | H | 174.66 | 181.52 | 6
I3 | B | 8 | 1 | H | 185.26 | 178.19 | 6
I3 | B | 9 | 1 | H | 182.67 | 185.74 | 6
I3 | B | 10 | 1 | H | 172.72 | 176.15 | 6
I3 | B | 11 | 1 | H | 176.21 | 172.53 | 6
I3 | B | 12 | 1 | H | 177.44 | 178.07 | 6
I3 | B | 13 | 1 | H | 187.12 | 178.08 | 6
I3 | B | 14 | 1 | H | 177.76 | 175.55 | 6
I3 | B | 15 | 1 | H | 186.87 | 186.12 | 6
I3 | B | 16 | 1 | H | 172.79 | 174.87 | 6
I3 | B | 17 | 1 | H | 177.06 | 172.41 | 6
I3 | B | 18 | 1 | H | 166.99 | 180.65 | 6
I3 | B | 19 | 1 | H | 168.67 | 169.77 | 6
I3 | B | 20 | 1 | H | 177.79 | 179.2 | 6
I3 | B | 21 | 1 | H | 178.8 | 181.07 | 6
I3 | B | 22 | 1 | H | 173.51 | 174.79 | 6
I3 | B | 23 | 1 | H | 174.11 | 175.86 | 6
I3 | B | 24 | 1 | H | 171.05 | 172.76 | 6
I3 | B | 25 | 1 | H | 174.65 | 176.27 | 6
I3 | B | 26 | 1 | H | 164.58 | 173.81 | 6
I3 | B | 27 | 1 | H | 177.33 | 177.77 | 6
I3 | B | 28 | 1 | H | 167.88 | 167.89 | 6
I3 | B | 29 | 1 | H | 171.27 | 169.35 | 6
I3 | B | 30 | 1 | H | 166.66 | 167.7 | 6
I3 | B | 31 | 1 | H | 170.4 | 174.14 | 6
I3 | B | 32 | 1 | H | 167.5 | 172.13 | 6
I3 | B | 1 | 2 | H | 176.86 | 173.69 | 6
I3 | B | 2 | 2 | H | 172.34 | 174.56 | 6
I3 | B | 3 | 2 | H | 173.39 | 175.01 | 6
I3 | B | 4 | 2 | H | 6.45 | 4.69 | 6
I3 | B | 5 | 2 | H | 174.24 | 183.09 | 6
I3 | B | 6 | 2 | H | 178.48 | 184.81 | 6
I3 | B | 7 | 2 | H | 176.29 | 181.23 | 6
I3 | B | 8 | 2 | H | 177.35 | 185.25 | 6
I3 | B | 9 | 2 | H | 184 | 184.74 | 6
I3 | B | 10 | 2 | H | 174.54 | 175.63 | 6
I3 | B | 11 | 2 | H | 182.63 | 178.06 | 6
I3 | B | 12 | 2 | H | 179.22 | 182.72 | 6
I3 | B | 13 | 2 | H | 181.79 | 180 | 6
I3 | B | 14 | 2 | H | 185.16 | 178.98 | 6
I3 | B | 15 | 2 | H | 184.76 | 178.43 | 6
I3 | B | 16 | 2 | H | 177.34 | 181.15 | 6
I3 | B | 17 | 2 | H | 177.76 | 173.35 | 6
I3 | B | 18 | 2 | H | 169.83 | 173.08 | 6
I3 | B | 19 | 2 | H | 176.18 | 178.51 | 6
I3 | B | 20 | 2 | H | 177.84 | 176.04 | 6
I3 | B | 21 | 2 | H | 177.67 | 175.68 | 6
I3 | B | 22 | 2 | H | 173.14 | 167.17 | 6
I3 | B | 23 | 2 | H | 171.4 | 176.99 | 6
I3 | B | 24 | 2 | H | 164.31 | 175.62 | 6
I3 | B | 25 | 2 | H | 170.96 | 180.42 | 6
I3 | B | 26 | 2 | H | 171.36 | 176.2 | 6
I3 | B | 27 | 2 | H | 177.21 | 177.28 | 6
I3 | B | 28 | 2 | H | 169.19 | 176.9 | 6
I3 | B | 29 | 2 | H | 176.23 | 172.76 | 6
I3 | B | 30 | 2 | H | 167 | 165.78 | 6
I3 | B | 31 | 2 | H | 172.63 | 173.38 | 6
I3 | B | 32 | 2 | H | 165.5 | 174.95 | 6
M1 | A | 1 | 1 | H | 590.45 | 589 | 352
M1 | A | 2 | 1 | H | 592.08 | 568.98 | 352
M1 | A | 3 | 1 | H | 596.67 | 586.59 | 352
M1 | A | 4 | 1 | H | 613.89 | 573.68 | 352
M1 | A | 5 | 1 | H | 595.96 | 596.51 | 352
M1 | A | 6 | 1 | H | 595.39 | 584.6 | 352
M1 | A | 7 | 1 | H | 596.59 | 597.9 | 352
M1 | A | 8 | 1 | H | 617.77 | 601.91 | 352
M1 | A | 9 | 1 | H | 608.16 | 596.21 | 352
M1 | A | 10 | 1 | H | 630.13 | 607.12 | 352
M1 | A | 11 | 1 | H | 600.5 | 592.42 | 352
M1 | A | 12 | 1 | H | 621.06 | 604.35 | 352
M1 | A | 13 | 1 | H | 597.64 | 592.03 | 352
M1 | A | 14 | 1 | H | 596.29 | 586.65 | 352
M1 | A | 15 | 1 | H | 618.51 | 572.6 | 352
M1 | A | 16 | 1 | H | 602.34 | 574.79 | 352
M1 | B | 1 | 1 | H | 594.34 | 583.5 | 352
M1 | B | 2 | 1 | H | 585.4 | 563.59 | 352
M1 | B | 3 | 1 | H | 591.61 | 577.42 | 352
M1 | B | 4 | 1 | H | 602.32 | 539.08 | 352
M1 | B | 5 | 1 | H | 586.68 | 573.81 | 352
M1 | B | 6 | 1 | H | 595.97 | 567.35 | 352
M1 | B | 7 | 1 | H | 611.48 | 598.26 | 352
M1 | B | 8 | 1 | H | 614.54 | 596.76 | 352
M1 | B | 9 | 1 | H | 612.48 | 598.25 | 352
M1 | B | 10 | 1 | H | 622.75 | 599.13 | 352
M1 | B | 11 | 1 | H | 614.76 | 589.79 | 352
M1 | B | 12 | 1 | H | 625.31 | 587.05 | 352
M1 | B | 13 | 1 | H | 605.2 | 561.59 | 352
M1 | B | 14 | 1 | H | 598.34 | 571.1 | 352
M1 | B | 15 | 1 | H | 590.87 | 575.73 | 352
M1 | B | 16 | 1 | H | 604.46 | 578.98 | 352
M2 | A | 1 | 1 | H | 518.32 | 515.21 | 380
M2 | A | 2 | 1 | H | 539.99 | 524.42 | 380
M2 | A | 3 | 1 | H | 533.56 | 519.96 | 380
M2 | A | 4 | 1 | H | 533.27 | 510.67 | 380
M2 | A | 5 | 1 | H | 548.85 | 511.13 | 380
M2 | A | 6 | 1 | H | 533.6 | 528.07 | 380
M2 | A | 7 | 1 | H | 552.3 | 518.24 | 380
M2 | A | 8 | 1 | H | 513.73 | 534.88 | 380
M2 | A | 9 | 1 | H | 560.25 | 531.13 | 380
M2 | A | 10 | 1 | H | 549.4 | 527.32 | 380
M2 | A | 11 | 1 | H | 536.24 | 519.61 | 380
M2 | A | 12 | 1 | H | 530.15 | 530.57 | 380
M2 | A | 13 | 1 | H | 569.79 | 539 | 380
M2 | A | 14 | 1 | H | 541.04 | 539.56 | 380
M2 | A | 15 | 1 | H | 531.07 | 512.42 | 380
M2 | A | 16 | 1 | H | 542.47 | 537.83 | 380
M2 | B | 1 | 1 | H | 529.02 | 517 | 380
M2 | B | 2 | 1 | H | 529.85 | 494.14 | 380
M2 | B | 3 | 1 | H | 530.35 | 525.52 | 380
M2 | B | 4 | 1 | H | 549.48 | 514.5 | 380
M2 | B | 5 | 1 | H | 530.88 | 506.31 | 380
M2 | B | 6 | 1 | H | 534.74 | 497.94 | 380
M2 | B | 7 | 1 | H | 539.73 | 507.15 | 380
M2 | B | 8 | 1 | H | 541.11 | 513.46 | 380
M2 | B | 9 | 1 | H | 520.36 | 519.1 | 380
M2 | B | 10 | 1 | H | 549.36 | 505.02 | 380
M2 | B | 11 | 1 | H | 533.27 | 515.21 | 380
M2 | B | 12 | 1 | H | 525.93 | 488.96 | 380
M2 | B | 13 | 1 | H | 541.72 | 504.73 | 380
M2 | B | 14 | 1 | H | 521.19 | 505.1 | 380
M2 | B | 15 | 1 | H | 537.82 | 529.53 | 380
M2 | B | 16 | 1 | H | 532.82 | 521.98 | 380
M3 | A | 1 | 1 | H | 683.5 | 640.64 | 416
M3 | A | 2 | 1 | H | 673.9 | 644.8 | 416
M3 | A | 3 | 1 | H | 684.63 | 658.42 | 416
M3 | A | 4 | 1 | H | 672.54 | 638.14 | 416
M3 | A | 5 | 1 | H | 674.35 | 640.01 | 416
M3 | A | 6 | 1 | H | 671.83 | 640.24 | 416
M3 | A | 7 | 1 | H | 655.61 | 635.97 | 416
M3 | A | 8 | 1 | H | 673.12 | 641.55 | 416
M3 | A | 9 | 1 | H | 656.37 | 660.84 | 416
M3 | A | 10 | 1 | H | 674.97 | 662.22 | 416
M3 | A | 11 | 1 | H | 668.27 | 639.63 | 416
M3 | A | 12 | 1 | H | 652.61 | 641.68 | 416
M3 | A | 13 | 1 | H | 665.09 | 623.41 | 416
M3 | A | 14 | 1 | H | 653.76 | 649.99 | 416
M3 | A | 15 | 1 | H | 672.3 | 656.46 | 416
M3 | A | 16 | 1 | H | 654.58 | 644.79 | 416
M3 | B | 1 | 1 | H | 663.17 | 656.5 | 416
M3 | B | 2 | 1 | H | 672.99 | 640.59 | 416
M3 | B | 3 | 1 | H | 660.9 | 646.02 | 416
M3 | B | 4 | 1 | H | 653.98 | 656.02 | 416
M3 | B | 5 | 1 | H | 656.3 | 675.96 | 416
M3 | B | 6 | 1 | H | 669.78 | 633.69 | 416
M3 | B | 7 | 1 | H | 674.96 | 657.11 | 416
M3 | B | 8 | 1 | H | 661.4 | 646.58 | 416
M3 | B | 9 | 1 | H | 680.72 | 664.57 | 416
M3 | B | 10 | 1 | H | 670.01 | 634.33 | 416
M3 | B | 11 | 1 | H | 667.01 | 625.65 | 416
M3 | B | 12 | 1 | H | 676.14 | 661.32 | 416
M3 | B | 13 | 1 | H | 676.02 | 629.12 | 416
M3 | B | 14 | 1 | H | 645.61 | 654.17 | 416
M3 | B | 15 | 1 | H | 657.03 | 641.82 | 416
M3 | B | 16 | 1 | H | 679.44 | 634.07 | 416
M4 | A | 1 | 1 | H | 534.94 | 524.69 | 362
M4 | A | 2 | 1 | H | 525.87 | 500.95 | 362
M4 | A | 3 | 1 | H | 527.56 | 521.24 | 362
M4 | A | 4 | 1 | H | 482.44 | 507.99 | 362
M4 | A | 5 | 1 | H | 494.71 | 488.16 | 362
M4 | A | 6 | 1 | H | 496.14 | 489.4 | 362
M4 | A | 7 | 1 | H | 524.36 | 520.51 | 362
M4 | A | 8 | 1 | H | 530.01 | 520.78 | 362
M4 | A | 9 | 1 | H | 513.52 | 496.67 | 362
M4 | A | 10 | 1 | H | 518.74 | 505.89 | 362
M4 | A | 11 | 1 | H | 523.5 | 525.56 | 362
M4 | A | 12 | 1 | H | 528.94 | 481.79 | 362
M4 | A | 13 | 1 | H | 507.83 | 492.29 | 362
M4 | A | 14 | 1 | H | 520.3 | 485.78 | 362
M4 | A | 15 | 1 | H | 533.39 | 505.16 | 362
M4 | A | 16 | 1 | H | 494.83 | 510.42 | 362
M4 | B | 1 | 1 | H | 506.64 | 521.31 | 362
M4 | B | 2 | 1 | H | 543.56 | 491.24 | 362
M4 | B | 3 | 1 | H | 525.76 | 519.95 | 362
M4 | B | 4 | 1 | H | 483.42 | 500.08 | 362
M4 | B | 5 | 1 | H | 499.62 | 508.01 | 362
M4 | B | 6 | 1 | H | 487.44 | 499.51 | 362
M4 | B | 7 | 1 | H | 512.44 | 507.28 | 362
M4 | B | 8 | 1 | H | 524.28 | 532.74 | 362
M4 | B | 9 | 1 | H | 518.9 | 500.06 | 362
M4 | B | 10 | 1 | H | 507.8 | 530.16 | 362
M4 | B | 11 | 1 | H | 531.46 | 526.14 | 362
M4 | B | 12 | 1 | H | 547.04 | 491.72 | 362
M4 | B | 13 | 1 | H | 519.45 | 503.56 | 362
M4 | B | 14 | 1 | H | 523.79 | 486.94 | 362
M4 | B | 15 | 1 | H | 518.09 | 534.65 | 362
M4 | B | 16 | 1 | H | 526.17 | 505.63 | 362
M5 | A | 1 | 1 | H | 355.79 | 346.83 | 242
M5 | A | 2 | 1 | H | 350.67 | 333.45 | 242
M5 | A | 3 | 1 | H | 363.39 | 354.56 | 242
M5 | A | 4 | 1 | H | 371.21 | 347.86 | 242
M5 | A | 5 | 1 | H | 351.33 | 351.47 | 242
M5 | A | 6 | 1 | H | 357.61 | 345.44 | 242
M5 | A | 7 | 1 | H | 353.23 | 358.78 | 242
M5 | A | 8 | 1 | H | 354.58 | 375.06 | 242
M5 | A | 9 | 1 | H | 368.29 | 332.31 | 242
M5 | A | 10 | 1 | H | 358.19 | 346.49 | 242
M5 | A | 11 | 1 | H | 363.72 | 382.35 | 242
M5 | A | 12 | 1 | H | 362.41 | 336.69 | 242
M5 | A | 13 | 1 | H | 365.75 | 353.67 | 242
M5 | A | 14 | 1 | H | 360.09 | 359.2 | 242
M5 | A | 15 | 1 | H | 360.42 | 332.64 | 242
M5 | A | 16 | 1 | H | 364.3 | 351.29 | 242
M5 | B | 1 | 1 | H | 370.56 | 351.09 | 242
M5 | B | 2 | 1 | H | 363.23 | 348.44 | 242
M5 | B | 3 | 1 | H | 353.29 | 328.68 | 242
M5 | B | 4 | 1 | H | 367.03 | 348.49 | 242
M5 | B | 5 | 1 | H | 351.53 | 333.86 | 242
M5 | B | 6 | 1 | H | 377.01 | 325.95 | 242
M5 | B | 7 | 1 | H | 360.86 | 347.56 | 242
M5 | B | 8 | 1 | H | 354.41 | 348.41 | 242
M5 | B | 9 | 1 | H | 356.43 | 346.58 | 242
M5 | B | 10 | 1 | H | 369.81 | 356.67 | 242
M5 | B | 11 | 1 | H | 374.43 | 351.01 | 242
M5 | B | 12 | 1 | H | 350.06 | 317.96 | 242
M5 | B | 13 | 1 | H | 373.71 | 345.07 | 242
M5 | B | 14 | 1 | H | 366.21 | 333.21 | 242
M5 | B | 15 | 1 | H | 343.17 | 348.69 | 242
M5 | B | 16 | 1 | H | 347.16 | 340.66 | 242
M6 | A | 1 | 1 | H | 387.56 | 366.26 | 199
M6 | A | 2 | 1 | H | 378.95 | 361.18 | 199
M6 | A | 3 | 1 | H | 377.81 | 359.46 | 199
M6 | A | 4 | 1 | H | 376.31 | 393.15 | 199
M6 | A | 5 | 1 | H | 386.24 | 378.48 | 199
M6 | A | 6 | 1 | H | 382.42 | 379.74 | 199
M6 | A | 7 | 1 | H | 364.99 | 374.42 | 199
M6 | A | 8 | 1 | H | 377.18 | 370.14 | 199
M6 | A | 9 | 1 | H | 372.5 | 377.59 | 199
M6 | A | 10 | 1 | H | 368.45 | 379.51 | 199
M6 | A | 11 | 1 | H | 365.21 | 382.89 | 199
M6 | A | 12 | 1 | H | 368.67 | 379.87 | 199
M6 | A | 13 | 1 | H | 383.16 | 369.05 | 199
M6 | A | 14 | 1 | H | 364.24 | 372.55 | 199
M6 | A | 15 | 1 | H | 384.93 | 392.84 | 199
M6 | A | 16 | 1 | H | 374.55 | 374.46 | 199
M6 | B | 1 | 1 | H | 382.67 | 377.23 | 199
M6 | B | 2 | 1 | H | 373.61 | 365.92 | 199
M6 | B | 3 | 1 | H | 390.46 | 368.38 | 199
M6 | B | 4 | 1 | H | 388.49 | 363.29 | 199
M6 | B | 5 | 1 | H | 374.26 | 381.83 | 199
M6 | B | 6 | 1 | H | 368.21 | 384.63 | 199
M6 | B | 7 | 1 | H | 368.33 | 386.77 | 199
M6 | B | 8 | 1 | H | 404.94 | 386.34 | 199
M6 | B | 9 | 1 | H | 382.26 | 386.66 | 199
M6 | B | 10 | 1 | H | 392.8 | 394.76 | 199
M6 | B | 11 | 1 | H | 379.33 | 383.11 | 199
M6 | B | 12 | 1 | H | 377.73 | 370.21 | 199
M6 | B | 13 | 1 | H | 381.97 | 378.06 | 199
M6 | B | 14 | 1 | H | 382.1 | 371.23 | 199
M6 | B | 15 | 1 | H | 377.78 | 367.27 | 199
M6 | B | 16 | 1 | H | 370.06 | 363.77 | 199
M7 | A | 1 | 1 | H | 487.69 | 554.37 | 215
M7 | A | 2 | 1 | H | 518.54 | 464.91 | 215
M7 | A | 3 | 1 | H | 513.25 | 540.82 | 215
M7 | A | 4 | 1 | H | 501.06 | 560 | 215
M7 | A | 5 | 1 | H | 525.69 | 537.4 | 215
M7 | A | 6 | 1 | H | 544.3 | 522.42 | 215
M7 | A | 7 | 1 | H | 480.04 | 542.99 | 215
M7 | A | 8 | 1 | H | 512.95 | 522.07 | 215
M7 | A | 9 | 1 | H | 530.97 | 510.59 | 215
M7 | A | 10 | 1 | H | 478.01 | 528.12 | 215
M7 | A | 11 | 1 | H | 503.54 | 532.67 | 215
M7 | A | 12 | 1 | H | 512.93 | 491.55 | 215
M7 | A | 13 | 1 | H | 531.63 | 512.87 | 215
M7 | A | 14 | 1 | H | 530.79 | 510.04 | 215
M7 | A | 15 | 1 | H | 532.54 | 509.63 | 215
M7 | A | 16 | 1 | H | 510.09 | 515.89 | 215
M7 | B | 1 | 1 | H | 481.14 | 552.41 | 215
M7 | B | 2 | 1 | H | 522.15 | 494.3 | 215
M7 | B | 3 | 1 | H | 531.68 | 504.74 | 215
M7 | B | 4 | 1 | H | 528.54 | 541.22 | 215
M7 | B | 5 | 1 | H | 529.28 | 535.1 | 215
M7 | B | 6 | 1 | H | 534.83 | 516.78 | 215
M7 | B | 7 | 1 | H | 479.57 | 546.84 | 215
M7 | B | 8 | 1 | H | 503.81 | 521.49 | 215
M7 | B | 9 | 1 | H | 509.69 | 503.34 | 215
M7 | B | 10 | 1 | H | 497.91 | 551.69 | 215
M7 | B | 11 | 1 | H | 508.84 | 502.21 | 215
M7 | B | 12 | 1 | H | 521.07 | 508.34 | 215
M7 | B | 13 | 1 | H | 514.31 | 503.69 | 215
M7 | B | 14 | 1 | H | 490.18 | 497.85 | 215
M7 | B | 15 | 1 | H | 506.6 | 519.78 | 215
M7 | B | 16 | 1 | H | 505.05 | 500.27 | 215
M8 | A | 1 | 1 | H | 266.11 | 372.9 | 74
M8 | A | 2 | 1 | H | 349.08 | 296.8 | 74
M8 | A | 3 | 1 | H | 317.94 | 267.64 | 74
M8 | A | 4 | 1 | H | 285.7 | 316.01 | 74
M8 | A | 5 | 1 | H | 353.44 | 322.71 | 74
M8 | A | 6 | 1 | H | 342.25 | 302.58 | 74
M8 | A | 7 | 1 | H | 222.51 | 232.63 | 74
M8 | A | 8 | 1 | H | 284.32 | 256.29 | 74
M8 | A | 9 | 1 | H | 389.11 | 249.23 | 74
M8 | A | 10 | 1 | H | 245.32 | 223.49 | 74
M8 | A | 11 | 1 | H | 306.36 | 320.41 | 74
M8 | A | 12 | 1 | H | 278.12 | 213.2 | 74
M8 | A | 13 | 1 | H | 249.18 | 252.95 | 74
M8 | A | 14 | 1 | H | 194.23 | 246.76 | 74
M8 | A | 15 | 1 | H | 296.75 | 237 | 74
M8 | A | 16 | 1 | H | 433.63 | 322.92 | 74
M8 | B | 1 | 1 | H | 272.59 | 312.23 | 74
M8 | B | 2 | 1 | H | 255.98 | 305.57 | 74
M8 | B | 3 | 1 | H | 250.89 | 336.11 | 74
M8 | B | 4 | 1 | H | 413.79 | 306.73 | 74
M8 | B | 5 | 1 | H | 307.95 | 253.45 | 74
M8 | B | 6 | 1 | H | 249.4 | 324.16 | 74
M8 | B | 7 | 1 | H | 339.5 | 367.29 | 74
M8 | B | 8 | 1 | H | 238.72 | 259.69 | 74
M8 | B | 9 | 1 | H | 281.66 | 304.75 | 74
M8 | B | 10 | 1 | H | 269.34 | 287.76 | 74
M8 | B | 11 | 1 | H | 229.96 | 269.43 | 74
M8 | B | 12 | 1 | H | 342.63 | 266.06 | 74
M8 | B | 13 | 1 | H | 280.91 | 279.76 | 74
M8 | B | 14 | 1 | H | 308.57 | 314.19 | 74
M8 | B | 15 | 1 | H | 314.71 | 320.11 | 74
M8 | B | 16 | 1 | H | 215.42 | 274.11 | 74
M9 | A | 1 | 1 | H | 270.86 | 249.23 | 83
M9 | A | 2 | 1 | H | 266.09 | 245.64 | 83
M9 | A | 3 | 1 | H | 254.3 | 247.17 | 83
M9 | A | 4 | 1 | H | 258.01 | 247.83 | 83
M9 | A | 5 | 1 | H | 255.85 | 243.61 | 83
M9 | A | 6 | 1 | H | 257.46 | 232.5 | 83
M9 | A | 7 | 1 | H | 246.34 | 236.59 | 83
M9 | A | 8 | 1 | H | 250.16 | 223.8 | 83
M9 | A | 9 | 1 | H | 251.72 | 225.48 | 83
M9 | A | 10 | 1 | H | 243.36 | 229.53 | 83
M9 | A | 11 | 1 | H | 244.13 | 230.77 | 83
M9 | A | 12 | 1 | H | 242.92 | 222.8 | 83
M9 | A | 13 | 1 | H | 227.21 | 225.88 | 83
M9 | A | 14 | 1 | H | 232 | 223.26 | 83
M9 | A | 15 | 1 | H | 232.4 | 212.24 | 83
M9 | A | 16 | 1 | H | 219.9 | 217.92 | 83
M9 | B | 1 | 1 | H | 269.09 | 261.81 | 83
M9 | B | 2 | 1 | H | 259.67 | 250.33 | 83
M9 | B | 3 | 1 | H | 245.47 | 246.9 | 83
M9 | B | 4 | 1 | H | 255.43 | 241.25 | 83
M9 | B | 5 | 1 | H | 271.16 | 237.75 | 83
M9 | B | 6 | 1 | H | 253.86 | 242.8 | 83
M9 | B | 7 | 1 | H | 250.98 | 232.79 | 83
M9 | B | 8 | 1 | H | 248.83 | 234.37 | 83
M9 | B | 9 | 1 | H | 250.61 | 232.3 | 83
M9 | B | 10 | 1 | H | 244.15 | 219.84 | 83
M9 | B | 11 | 1 | H | 245.87 | 226.56 | 83
M9 | B | 12 | 1 | H | 249.78 | 235.29 | 83
M9 | B | 13 | 1 | H | 227.78 | 227.19 | 83
M9 | B | 14 | 1 | H | 235.24 | 223.26 | 83
M9 | B | 15 | 1 | H | 239.15 | 211.23 | 83
M9 | B | 16 | 1 | H | 233.02 | 209.55 | 83
M10 | A | 1 | 1 | H | 635.04 | 628.69 | 342
M10 | A | 2 | 1 | H | 712.83 | 686.11 | 342
M10 | A | 3 | 1 | H | 674.96 | 648.44 | 342
M10 | A | 4 | 1 | H | 659.02 | 643.98 | 342
M10 | A | 5 | 1 | H | 553.23 | 599.83 | 342
M10 | A | 6 | 1 | H | 607.64 | 573.53 | 342
M10 | A | 7 | 1 | H | 748.63 | 679.95 | 342
M10 | A | 8 | 1 | H | 603.37 | 635.54 | 342
M10 | A | 9 | 1 | H | 599.92 | 649.94 | 342
M10 | A | 10 | 1 | H | 766.19 | 634.42 | 342
M10 | A | 11 | 1 | H | 616.04 | 565.29 | 342
M10 | A | 12 | 1 | H | 614.94 | 734.07 | 342
M10 | A | 13 | 1 | H | 642.36 | 722.12 | 342
M10 | A | 14 | 1 | H | 705.27 | 673.53 | 342
M10 | A | 15 | 1 | H | 623.38 | 629.81 | 342
M10 | A | 16 | 1 | H | 566.34 | 592.63 | 342
M10 | B | 1 | 1 | H | 770.01 | 620.83 | 342
M10 | B | 2 | 1 | H | 673.19 | 873.81 | 342
M10 | B | 3 | 1 | H | 651.69 | 695.68 | 342
M10 | B | 4 | 1 | H | 804.67 | 626.3 | 342
M10 | B | 5 | 1 | H | 678.21 | 796.87 | 342
M10 | B | 6 | 1 | H | 560.01 | 591.49 | 342
M10 | B | 7 | 1 | H | 667.13 | 722.08 | 342
M10 | B | 8 | 1 | H | 639.38 | 611.17 | 342
M10 | B | 9 | 1 | H | 632.73 | 645.32 | 342
M10 | B | 10 | 1 | H | 559.13 | 629.91 | 342
M10 | B | 11 | 1 | H | 848.94 | 699.37 | 342
M10 | B | 12 | 1 | H | 617.33 | 686.67 | 342
M10 | B | 13 | 1 | H | 623.87 | 630.3 | 342
M10 | B | 14 | 1 | H | 581.57 | 716.8 | 342
M10 | B | 15 | 1 | H | 524.36 | 606.24 | 342
M10 | B | 16 | 1 | H | 507.91 | 584.62 | 342
M11 | A | 1 | 1 | H | 159.68 | 45.88 | 10
M11 | A | 2 | 1 | H | 63.53 | 39.87 | 10
M11 | A | 3 | 1 | H | 1.39 | 7.61 | 10
M11 | A | 4 | 1 | H | 0.65 | 149.54 | 10
M11 | A | 5 | 1 | H | 60.35 | 28.79 | 10
M11 | A | 6 | 1 | H | 68.42 | 94.24 | 10
M11 | A | 7 | 1 | H | 73.49 | 39.94 | 10
M11 | A | 8 | 1 | H | 80.34 | 43.8 | 10
M11 | A | 9 | 1 | H | 37.37 | 84.18 | 10
M11 | A | 10 | 1 | H | 56.73 | 59.65 | 10
M11 | A | 11 | 1 | H | 64.62 | 34.29 | 10
M11 | A | 12 | 1 | H | 49.55 | 70.99 | 10
M11 | A | 13 | 1 | H | 68.51 | 38.66 | 10
M11 | A | 14 | 1 | H | 50.44 | 38.48 | 10
M11 | A | 15 | 1 | H | 26.44 | 112.97 | 10
M11 | A | 16 | 1 | H | 14.96 | 111.88 | 10
M11 | B | 1 | 1 | H | 96.96 | 57.83 | 10
M11 | B | 2 | 1 | H | 52.23 | 55.39 | 10
M11 | B | 3 | 1 | H | 19.12 | 31.64 | 10
M11 | B | 4 | 1 | H | 19.35 | 86.51 | 10
M11 | B | 5 | 1 | H | 56.96 | 10.81 | 10
M11 | B | 6 | 1 | H | 51 | 85.91 | 10
M11 | B | 7 | 1 | H | 121.08 | 13.38 | 10
M11 | B | 8 | 1 | H | 110.21 | 67.83 | 10
M11 | B | 9 | 1 | H | 21.9 | 73.31 | 10
M11 | B | 10 | 1 | H | 70.55 | 25.93 | 10
M11 | B | 11 | 1 | H | 71.8 | 36.57 | 10
M11 | B | 12 | 1 | H | 76.13 | 77.64 | 10
M11 | B | 13 | 1 | H | 44.82 | 100.36 | 10
M11 | B | 14 | 1 | H | 24.54 | 15.13 | 10
M11 | B | 15 | 1 | H | 1.14 | 103.52 | 10
M11 | B | 16 | 1 | H | 14.17 | 48.92 | 10
M1 | A | 1 | 1 | L | 978.46 | 961.73 | 316
M1 | A | 2 | 1 | L | 1024.11 | 997.49 | 316
M1 | A | 3 | 1 | L | 1021.64 | 1009.37 | 316
M1 | A | 4 | 1 | L | 1049.12 | 921.08 | 316
M1 | A | 5 | 1 | L | 978.69 | 1007.13 | 316
M1 | A | 6 | 1 | L | 1051.78 | 974.29 | 316
M1 | A | 7 | 1 | L | 1011.14 | 1008.2 | 316
M1 | A | 8 | 1 | L | 1012.72 | 1010.77 | 316
M1 | A | 9 | 1 | L | 1019.44 | 1014.51 | 316
M1 | A | 10 | 1 | L | 1019.38 | 1101.86 | 316
M1 | A | 11 | 1 | L | 1155.86 | 1038.64 | 316
M1 | A | 12 | 1 | L | 1025.53 | 903.92 | 316
M1 | A | 13 | 1 | L | 1053.5 | 952.46 | 316
M1 | A | 14 | 1 | L | 1035.17 | 1030.9 | 316
M1 | A | 15 | 1 | L | 989 | 986.48 | 316
M1 | A | 16 | 1 | L | 997.69 | 929.62 | 316
M1 | B | 1 | 1 | L | 1011.2 | 970.39 | 316
M1 | B | 2 | 1 | L | 1012.08 | 960.74 | 316
M1 | B | 3 | 1 | L | 1051.2 | 960.63 | 316
M1 | B | 4 | 1 | L | 998.29 | 952.69 | 316
M1 | B | 5 | 1 | L | 1026.9 | 941.63 | 316
M1 | B | 6 | 1 | L | 1027.06 | 970.01 | 316
M1 | B | 7 | 1 | L | 1044.23 | 968.37 | 316
M1 | B | 8 | 1 | L | 1023.39 | 968.76 | 316
M1 | B | 9 | 1 | L | 1008.77 | 969.91 | 316
M1 | B | 10 | 1 | L | 1042.9 | 974.41 | 316
M1 | B | 11 | 1 | L | 1038.87 | 978.72 | 316
M1 | B | 12 | 1 | L | 1029.7 | 994.35 | 316
M1 | B | 13 | 1 | L | 1010.06 | 918.96 | 316
M1 | B | 14 | 1 | L | 1009.58 | 979.01 | 316
M1 | B | 15 | 1 | L | 1000.85 | 953.17 | 316
M1 | B | 16 | 1 | L | 930 | 924.49 | 316
M2 | A | 1 | 1 | L | 920.04 | 941.21 | 409
M2 | A | 2 | 1 | L | 911.96 | 904.48 | 409
M2 | A | 3 | 1 | L | 933.03 | 902.75 | 409
M2 | A | 4 | 1 | L | 1004.8 | 934.93 | 409
M2 | A | 5 | 1 | L | 876.17 | 931.14 | 409
M2 | A | 6 | 1 | L | 960.13 | 898.07 | 409
M2 | A | 7 | 1 | L | 911.53 | 888.35 | 409
M2 | A | 8 | 1 | L | 955.18 | 905.09 | 409
M2 | A | 9 | 1 | L | 940.83 | 905.04 | 409
M2 | A | 10 | 1 | L | 913.02 | 954.68 | 409
M2 | A | 11 | 1 | L | 951.78 | 898.88 | 409
M2 | A | 12 | 1 | L | 936.2 | 871.89 | 409
M2 | A | 13 | 1 | L | 949.73 | 968.51 | 409
M2 | A | 14 | 1 | L | 900.73 | 916.29 | 409
M2 | A | 15 | 1 | L | 922.19 | 911.69 | 409
M2 | A | 16 | 1 | L | 902.79 | 902.61 | 409
M2 | B | 1 | 1 | L | 935.46 | 914.71 | 409
M2 | B | 2 | 1 | L | 936.05 | 866.31 | 409
M2 | B | 3 | 1 | L | 959.4 | 898.99 | 409
M2 | B | 4 | 1 | L | 896.86 | 911.61 | 409
M2 | B | 5 | 1 | L | 932.63 | 891.13 | 409
M2 | B | 6 | 1 | L | 1009.11 | 952.72 | 409
M2 | B | 7 | 1 | L | 961.69 | 886.55 | 409
M2 | B | 8 | 1 | L | 973.17 | 917.11 | 409
M2 | B | 9 | 1 | L | 963.39 | 935.66 | 409
M2 | B | 10 | 1 | L | 922.76 | 902.89 | 409
M2 | B | 11 | 1 | L | 1005.03 | 921.92 | 409
M2 | B | 12 | 1 | L | 923.51 | 910.47 | 409
M2 | B | 13 | 1 | L | 879.63 | 939.3 | 409
M2 | B | 14 | 1 | L | 929.98 | 896.26 | 409
M2 | B | 15 | 1 | L | 929.12 | 858.54 | 409
M2 | B | 16 | 1 | L | 914.55 | 908.17 | 409
M3 | A | 1 | 1 | L | 1034.95 | 1070.81 | 414
M3 | A | 2 | 1 | L | 1190.25 | 1036.26 | 414
M3 | A | 3 | 1 | L | 937.88 | 958.74 | 414
M3 | A | 4 | 1 | L | 975.65 | 1050.58 | 414
M3 | A | 5 | 1 | L | 1188.28 | 1069.89 | 414
M3 | A | 6 | 1 | L | 1072.32 | 1187.14 | 414
M3 | A | 7 | 1 | L | 1026.86 | 1142.92 | 414
M3 | A | 8 | 1 | L | 1024.58 | 1091.67 | 414
M3 | A | 9 | 1 | L | 900.63 | 1128.33 | 414
M3 | A | 10 | 1 | L | 1052.6 | 950.12 | 414
M3 | A | 11 | 1 | L | 1045.3 | 1142.25 | 414
M3 | A | 12 | 1 | L | 968.56 | 1044.08 | 414
M3 | A | 13 | 1 | L | 1075.66 | 1074.66 | 414
M3 | A | 14 | 1 | L | 1009.86 | 960.73 | 414
M3 | A | 15 | 1 | L | 1032.11 | 1061.43 | 414
M3 | A | 16 | 1 | L | 887.59 | 928.23 | 414
M3 | B | 1 | 1 | L | 959.35 | 1121.82 | 414
M3 | B | 2 | 1 | L | 1091.9 | 977.69 | 414
M3 | B | 3 | 1 | L | 1037.66 | 1142.7 | 414
M3 | B | 4 | 1 | L | 883.74 | 1013.28 | 414
M3 | B | 5 | 1 | L | 1041.51 | 1049.82 | 414
M3 | B | 6 | 1 | L | 1016.01 | 915.68 | 414
M3 | B | 7 | 1 | L | 883.3 | 1095.46 | 414
M3 | B | 8 | 1 | L | 945.03 | 1213.64 | 414
M3 | B | 9 | 1 | L | 1018.13 | 1101.71 | 414
M3 | B | 10 | 1 | L | 1074.48 | 985.5 | 414
M3 | B | 11 | 1 | L | 1006.89 | 904.49 | 414
M3 | B | 12 | 1 | L | 919.78 | 1083.89 | 414
M3 | B | 13 | 1 | L | 995.26 | 1041.75 | 414
M3 | B | 14 | 1 | L | 1003.87 | 1028.5 | 414
M3 | B | 15 | 1 | L | 887.72 | 912.88 | 414
M3 | B | 16 | 1 | L | 1012.77 | 1345.2 | 414
M4 | A | 1 | 1 | L | 753.42 | 758.01 | 315
M4 | A | 2 | 1 | L | 755.77 | 930.75 | 315
M4 | A | 3 | 1 | L | 857.5 | 735.31 | 315
M4 | A | 4 | 1 | L | 745 | 771.19 | 315
M4 | A | 5 | 1 | L | 784.11 | 732.39 | 315
M4 | A | 6 | 1 | L | 827.69 | 771.82 | 315
M4 | A | 7 | 1 | L | 781.55 | 750.71 | 315
M4 | A | 8 | 1 | L | 800.19 | 768.81 | 315
M4 | A | 9 | 1 | L | 820.2 | 830.07 | 315
M4 | A | 10 | 1 | L | 769.23 | 720.4 | 315
M4 | A | 11 | 1 | L | 739.81 | 810.52 | 315
M4 | A | 12 | 1 | L | 811.94 | 730.47 | 315
M4 | A | 13 | 1 | L | 753.29 | 775.76 | 315
M4 | A | 14 | 1 | L | 828.14 | 793.05 | 315
M4 | A | 15 | 1 | L | 770.25 | 790.01 | 315
M4 | A | 16 | 1 | L | 791.63 | 748.8 | 315
M4 | B | 1 | 1 | L | 779.06 | 765.53 | 315
M4 | B | 2 | 1 | L | 765.36 | 821.61 | 315
M4 | B | 3 | 1 | L | 787.82 | 722.69 | 315
M4 | B | 4 | 1 | L | 801.22 | 780.33 | 315
M4 | B | 5 | 1 | L | 804.78 | 767.52 | 315
M4 | B | 6 | 1 | L | 781.45 | 763.75 | 315
M4 | B | 7 | 1 | L | 788.09 | 823.27 | 315
M4 | B | 8 | 1 | L | 788.26 | 775.78 | 315
M4 | B | 9 | 1 | L | 761.13 | 777.27 | 315
M4 | B | 10 | 1 | L | 754.59 | 790.59 | 315
M4 | B | 11 | 1 | L | 749.98 | 762.13 | 315
M4 | B | 12 | 1 | L | 790.02 | 694.4 | 315
M4 | B | 13 | 1 | L | 729.06 | 766.91 | 315
M4 | B | 14 | 1 | L | 780.37 | 777.77 | 315
M4 | B | 15 | 1 | L | 838 | 759.06 | 315
M4 | B | 16 | 1 | L | 784.85 | 744.02 | 315
M5 | A | 1 | 1 | L | 829.6 | 631.68 | 360
M5 | A | 2 | 1 | L | 865.28 | 735.88 | 360
M5 | A | 3 | 1 | L | 682.36 | 635.57 | 360
M5 | A | 4 | 1 | L | 627.01 | 801.12 | 360
M5 | A | 5 | 1 | L | 707.46 | 634.21 | 360
M5 | A | 6 | 1 | L | 761.94 | 694.54 | 360
M5 | A | 7 | 1 | L | 654.53 | 691.59 | 360
M5 | A | 8 | 1 | L | 650.63 | 691.04 | 360
M5 | A | 9 | 1 | L | 714.8 | 635.84 | 360
M5 | A | 10 | 1 | L | 827.77 | 645.81 | 360
M5 | A | 11 | 1 | L | 778.2 | 623.15 | 360
M5 | A | 12 | 1 | L | 759.55 | 663.94 | 360
M5 | A | 13 | 1 | L | 641.8 | 795.32 | 360
M5 | A | 14 | 1 | L | 688.71 | 610.82 | 360
M5 | A | 15 | 1 | L | 635.24 | 774.77 | 360
M5 | A | 16 | 1 | L | 684.87 | 694.9 | 360
M5 | B | 1 | 1 | L | 664.56 | 687.78 | 360
M5 | B | 2 | 1 | L | 639.08 | 631.44 | 360
M5 | B | 3 | 1 | L | 699.83 | 651.71 | 360
M5 | B | 4 | 1 | L | 718.77 | 676.81 | 360
M5 | B | 5 | 1 | L | 771.83 | 631.52 | 360
M5 | B | 6 | 1 | L | 701.62 | 658.51 | 360
M5 | B | 7 | 1 | L | 748.77 | 666.12 | 360
M5 | B | 8 | 1 | L | 770.73 | 747.26 | 360
M5 | B | 9 | 1 | L | 698.53 | 764.51 | 360
M5 | B | 10 | 1 | L | 659.79 | 793.4 | 360
M5 | B | 11 | 1 | L | 717.6 | 677.68 | 360
M5 | B | 12 | 1 | L | 637.7 | 751.55 | 360
M5 | B | 13 | 1 | L | 737.82 | 652.94 | 360
M5 | B | 14 | 1 | L | 740.96 | 927.62 | 360
M5 | B | 15 | 1 | L | 664 | 648.57 | 360
M5 | B | 16 | 1 | L | 660.6 | 865.43 | 360
M7 | A | 1 | 1 | L | 722.26 | 622.58 | 340
M7 | A | 2 | 1 | L | 751.73 | 605.5 | 340
M7 | A | 3 | 1 | L | 682.45 | 611.16 | 340
M7 | A | 4 | 1 | L | 711.04 | 758.27 | 340
M7 | A | 5 | 1 | L | 636.73 | 627.15 | 340
M7 | A | 6 | 1 | L | 790.57 | 646.99 | 340
M7 | A | 7 | 1 | L | 667.49 | 696.8 | 340
M7 | A | 8 | 1 | L | 629.49 | 650.94 | 340
M7 | A | 9 | 1 | L | 773.79 | 846.66 | 340
M7 | A | 10 | 1 | L | 627.68 | 602.5 | 340
M7 | A | 11 | 1 | L | 808.09 | 757.27 | 340
M7 | A | 12 | 1 | L | 644.7 | 671.12 | 340
M7 | A | 13 | 1 | L | 665.78 | 788.66 | 340
M7 | A | 14 | 1 | L | 656.92 | 659.49 | 340
M7 | A | 15 | 1 | L | 687.31 | 731.99 | 340
M7 | A | 16 | 1 | L | 714.25 | 651.19 | 340
M7 | B | 1 | 1 | L | 796.39 | 572.22 | 340
M7 | B | 2 | 1 | L | 695.45 | 610.91 | 340
M7 | B | 3 | 1 | L | 596.42 | 673.32 | 340
M7 | B | 4 | 1 | L | 715.33 | 676.4 | 340
M7 | B | 5 | 1 | L | 690.52 | 675.02 | 340
M7 | B | 6 | 1 | L | 681.5 | 581.96 | 340
M7 | B | 7 | 1 | L | 629.75 | 602.97 | 340
M7 | B | 8 | 1 | L | 658.6 | 702.25 | 340
M7 | B | 9 | 1 | L | 691.63 | 652.52 | 340
M7 | B | 10 | 1 | L | 581.61 | 624.83 | 340
M7 | B | 11 | 1 | L | 701.44 | 671.25 | 340
M7 | B | 12 | 1 | L | 675.28 | 777.82 | 340
M7 | B | 13 | 1 | L | 727.48 | 711.81 | 340
M7 | B | 14 | 1 | L | 660.86 | 649.64 | 340
M7 | B | 15 | 1 | L | 769.64 | 750.31 | 340
M7 | B | 16 | 1 | L | 698.96 | 607.09 | 340

## DNB sensitivity
Band | HAM | Gain State | Agg Zone | Band Median |  | Band Average |  | SNR at Lmin Spec
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | SNR at Lmin | SNR at Lmin | SNR at Lmin | SNR at Lmin
HGA | A | 3 | 1 | 7.2015 | 10.383 | 7.2021 | 10.4064 | 6
HGA | A | 3 | 7 | 6.4249 | 8.1528 | 6.4253 | 8.2196 | 6
HGA | A | 3 | 11 | 5.4463 | 7.0118 | 5.4778 | 6.9595 | 6
HGA | A | 3 | 15 | 4.5516 | 5.4693 | 4.5421 | 5.5021 | 6
HGA | A | 3 | 20 | 3.6979 | 4.6084 | 3.6877 | 4.6232 | 6
HGA | A | 3 | 26 | 2.972 | 3.5048 | 2.9453 | 3.5046 | 6
HGA | A | 3 | 32 | 1.9435 | 2.2369 | 2.0593 | 2.3435 | 5
HGA | A | 3 | 34 | 10.5726 | 14.5284 | 13.2616 | 14.7167 | ~
HGA | B | 3 | 1 | 6.0363 | 8.1528 | 7.2266 | 10.2708 | 6
HGA | B | 3 | 7 | 6.4846 | 8.2047 | 6.4867 | 8.2754 | 6
HGA | B | 3 | 11 | 5.3787 | 6.9205 | 5.391 | 6.8941 | 6
HGA | B | 3 | 15 | 4.474 | 5.4693 | 4.4721 | 5.5191 | 6
HGA | B | 3 | 20 | 3.7108 | 4.6292 | 3.7035 | 4.6447 | 6
HGA | B | 3 | 26 | 2.9385 | 3.5079 | 2.9501 | 3.5228 | 6
HGA | B | 3 | 32 | 2.0088 | 2.3605 | 2.1323 | 2.4756 | 5
HGA | B | 3 | 34 | 10.4743 | 15.0377 | 10.9963 | 15.0207 | ~
HGB | A | 3 | 1 | 6.4149 | 8.0407 | 6.248 | 7.9815 | 6
HGB | A | 3 | 7 | 5.4378 | 7.0118 | 5.307 | 6.678 | 6
HGB | A | 3 | 11 | 4.5516 | 5.4342 | 4.3983 | 5.3823 | 6
HGB | A | 3 | 15 | 3.6979 | 4.5931 | 3.5487 | 4.4117 | 6
HGB | A | 3 | 20 | 2.972 | 3.4762 | 2.7654 | 3.2574 | 6
HGB | A | 3 | 26 | 2.1755 | 2.4208 | 6.3694 | 4.7176 | 6
HGB | A | 3 | 32 | 10.3759 | 14.2917 | 9.8661 | 13.7725 | 5
HGB | A | 3 | 34 | 7.2154 | 10.1243 | 7.1346 | 9.9014 | ~
HGB | B | 3 | 1 | 5.529 | #NUM! | 6.3218 | 8.0248 | 6
HGB | B | 3 | 7 | 5.3534 | 6.9205 | 5.2203 | 6.675 | 6
HGB | B | 3 | 11 | 4.4361 | 5.4693 | 4.304 | 5.3578 | 6
HGB | B | 3 | 15 | 3.7108 | 4.5666 | 3.5689 | 4.4524 | 6
HGB | B | 3 | 20 | 2.9385 | 3.5079 | 2.7698 | 3.287 | 6
HGB | B | 3 | 26 | 2.1933 | 2.4581 | 4.0833 | 4.6707 | 6
HGB | B | 3 | 32 | 10.332 | 15.0021 | 9.8849 | 14.0611 | 5
HGB | B | 3 | 34 | 6.9316 | 9.9433 | 6.8163 | 9.7403 | ~
Band | HAM | Gain State | Agg Zone | Detector | Detector Dependent |  | SNR at Lmin Spec |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 DNB radiometric sensitivity.
[5 empty] | SNR at Lmin | SNR at Lmin |  |  | 2. The results were derived from RC-1 using the SIS-100 source.
HGA | A | 3 | 1 | 1 | 7.1815 | 10.0641 | 6 |  | 3. Fits of the SIS-100 radiance vs SNR were used to determine the radiometric sensitivity.
HGA | A | 3 | 1 | 2 | 7.2215 | 9.7737 | 6 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
HGA | A | 3 | 1 | 3 | 6.6081 | 10.3743 | 6 |  | 5. The test was set to diagnostic mode.
HGA | A | 3 | 1 | 4 | 6.9066 | 9.8919 | 6
HGA | A | 3 | 1 | 5 | 7.2406 | 10.1237 | 6
HGA | A | 3 | 1 | 6 | 7.0022 | 11.1226 | 6
HGA | A | 3 | 1 | 7 | 7.0456 | 10.4566 | 6
HGA | A | 3 | 1 | 8 | 7.4762 | 10.9117 | 6
HGA | A | 3 | 1 | 9 | 7.1012 | 10.1577 | 6
HGA | A | 3 | 1 | 10 | 7.1724 | 10.2362 | 6
HGA | A | 3 | 1 | 11 | 7.3686 | 10.3916 | 6
HGA | A | 3 | 1 | 12 | 7.6152 | 9.6703 | 6
HGA | A | 3 | 1 | 13 | 7.1418 | 10.8798 | 6
HGA | A | 3 | 1 | 14 | 7.2982 | 10.5808 | 6
HGA | A | 3 | 1 | 15 | 7.3654 | 10.6138 | 6
HGA | A | 3 | 1 | 16 | 7.4883 | 11.254 | 6
HGA | A | 3 | 7 | 1 | 5.9623 | 7.6495 | 6
HGA | A | 3 | 7 | 2 | 6.4277 | 8.1718 | 6
HGA | A | 3 | 7 | 3 | 6.2299 | 8.1339 | 6
HGA | A | 3 | 7 | 4 | 6.4076 | 7.8696 | 6
HGA | A | 3 | 7 | 5 | 6.557 | 8.1791 | 6
HGA | A | 3 | 7 | 6 | 6.1915 | 7.7482 | 6
HGA | A | 3 | 7 | 7 | 6.2058 | 7.8569 | 6
HGA | A | 3 | 7 | 8 | 6.5001 | 8.0934 | 6
HGA | A | 3 | 7 | 9 | 6.2732 | 7.9724 | 6
HGA | A | 3 | 7 | 10 | 6.7153 | 8.2434 | 6
HGA | A | 3 | 7 | 11 | 6.7994 | 8.7314 | 6
HGA | A | 3 | 7 | 12 | 6.6151 | 8.6504 | 6
HGA | A | 3 | 7 | 13 | 6.1869 | 8.4153 | 6
HGA | A | 3 | 7 | 14 | 6.6567 | 8.4539 | 6
HGA | A | 3 | 7 | 15 | 6.4221 | 9.3568 | 6
HGA | A | 3 | 7 | 16 | 6.6548 | 7.9881 | 6
HGA | A | 3 | 11 | 1 | 4.9455 | 6.4222 | 6
HGA | A | 3 | 11 | 2 | 5.3959 | 6.7225 | 6
HGA | A | 3 | 11 | 3 | 5.4419 | 7.0002 | 6
HGA | A | 3 | 11 | 4 | 5.2509 | 7.3402 | 6
HGA | A | 3 | 11 | 5 | 5.2371 | 7.1095 | 6
HGA | A | 3 | 11 | 6 | 5.2945 | 7.1842 | 6
HGA | A | 3 | 11 | 7 | 5.4249 | 6.5233 | 6
HGA | A | 3 | 11 | 8 | 5.8998 | 7.0121 | 6
HGA | A | 3 | 11 | 9 | 5.7912 | 7.1451 | 6
HGA | A | 3 | 11 | 10 | 5.4507 | 7.0941 | 6
HGA | A | 3 | 11 | 11 | 5.7057 | 6.8631 | 6
HGA | A | 3 | 11 | 12 | 5.8366 | 7.0995 | 6
HGA | A | 3 | 11 | 13 | 5.5906 | 7.0115 | 6
HGA | A | 3 | 11 | 14 | 5.4519 | 6.5828 | 6
HGA | A | 3 | 11 | 15 | 5.6956 | 6.9648 | 6
HGA | A | 3 | 11 | 16 | 5.2322 | 7.277 | 6
HGA | A | 3 | 15 | 1 | 4.1185 | 5.0056 | 6
HGA | A | 3 | 15 | 2 | 4.4008 | 5.518 | 6
HGA | A | 3 | 15 | 3 | 4.5311 | 5.117 | 6
HGA | A | 3 | 15 | 4 | 4.7914 | 5.4244 | 6
HGA | A | 3 | 15 | 5 | 4.3352 | 5.5568 | 6
HGA | A | 3 | 15 | 6 | 4.5621 | 5.276 | 6
HGA | A | 3 | 15 | 7 | 4.7631 | 5.5846 | 6
HGA | A | 3 | 15 | 8 | 4.2934 | 5.4141 | 6
HGA | A | 3 | 15 | 9 | 4.6718 | 6.0429 | 6
HGA | A | 3 | 15 | 10 | 4.6543 | 5.4946 | 6
HGA | A | 3 | 15 | 11 | 4.5411 | 5.584 | 6
HGA | A | 3 | 15 | 12 | 4.7845 | 6.0228 | 6
HGA | A | 3 | 15 | 13 | 4.6426 | 5.3635 | 6
HGA | A | 3 | 15 | 14 | 4.4484 | 5.391 | 6
HGA | A | 3 | 15 | 15 | 4.6379 | 5.4439 | 6
HGA | A | 3 | 15 | 16 | 4.4968 | 5.7942 | 6
HGA | A | 3 | 20 | 1 | 3.5894 | 4.5268 | 6
HGA | A | 3 | 20 | 2 | 3.6795 | 4.3857 | 6
HGA | A | 3 | 20 | 3 | 3.4813 | 4.812 | 6
HGA | A | 3 | 20 | 4 | 3.6169 | 4.5935 | 6
HGA | A | 3 | 20 | 5 | 3.735 | 4.2432 | 6
HGA | A | 3 | 20 | 6 | 3.4787 | 4.8214 | 6
HGA | A | 3 | 20 | 7 | 3.8342 | 4.8314 | 6
HGA | A | 3 | 20 | 8 | 3.619 | 4.5019 | 6
HGA | A | 3 | 20 | 9 | 3.5346 | 4.5926 | 6
HGA | A | 3 | 20 | 10 | 3.8616 | 4.3867 | 6
HGA | A | 3 | 20 | 11 | 3.7007 | 4.9996 | 6
HGA | A | 3 | 20 | 12 | 3.7837 | 4.6233 | 6
HGA | A | 3 | 20 | 13 | 3.811 | 4.7572 | 6
HGA | A | 3 | 20 | 14 | 3.8419 | 4.7198 | 6
HGA | A | 3 | 20 | 15 | 3.6951 | 4.5357 | 6
HGA | A | 3 | 20 | 16 | 3.7405 | 4.6413 | 6
HGA | A | 3 | 26 | 1 | 2.8186 | 3.3956 | 6
HGA | A | 3 | 26 | 2 | 2.8661 | 3.542 | 6
HGA | A | 3 | 26 | 3 | 2.8419 | 3.4024 | 6
HGA | A | 3 | 26 | 4 | 2.968 | 3.3871 | 6
HGA | A | 3 | 26 | 5 | 2.9643 | 3.4522 | 6
HGA | A | 3 | 26 | 6 | 2.8575 | 3.3834 | 6
HGA | A | 3 | 26 | 7 | 2.976 | 3.3591 | 6
HGA | A | 3 | 26 | 8 | 2.979 | 3.5401 | 6
HGA | A | 3 | 26 | 9 | 3.0672 | 3.5155 | 6
HGA | A | 3 | 26 | 10 | 3.0183 | 3.8069 | 6
HGA | A | 3 | 26 | 11 | 2.9808 | 3.6903 | 6
HGA | A | 3 | 26 | 12 | 2.9812 | 3.4113 | 6
HGA | A | 3 | 26 | 13 | 2.7816 | 3.5815 | 6
HGA | A | 3 | 26 | 14 | 2.9872 | 3.5968 | 6
HGA | A | 3 | 26 | 15 | 2.9519 | 3.5093 | 6
HGA | A | 3 | 26 | 16 | 3.0855 | 3.5002 | 6
HGA | A | 3 | 32 | 1 | 1.8233 | 2.1192 | 6
HGA | A | 3 | 32 | 2 | 1.9657 | 2.1603 | 6
HGA | A | 3 | 32 | 3 | 1.8594 | 2.106 | 6
HGA | A | 3 | 32 | 4 | 2.3598 | 2.8612 | 6
HGA | A | 3 | 32 | 5 | 1.9154 | 2.0801 | 6
HGA | A | 3 | 32 | 6 | 2.0847 | 2.2691 | 6
HGA | A | 3 | 32 | 7 | 2.1376 | 2.5086 | 6
HGA | A | 3 | 32 | 8 | 2.2134 | 2.4813 | 6
HGA | A | 3 | 32 | 9 | 2.3206 | 2.723 | 6
HGA | A | 3 | 32 | 10 | 2.2552 | 2.3602 | 6
HGA | A | 3 | 32 | 11 | 1.9213 | 2.2878 | 6
HGA | A | 3 | 32 | 12 | 1.8992 | 2.1437 | 6
HGA | A | 3 | 32 | 13 | 2.5552 | 3.0411 | 6
HGA | A | 3 | 32 | 14 | 1.8574 | 2.0014 | 6
HGA | A | 3 | 32 | 15 | 1.8869 | 2.2048 | 6
HGA | A | 3 | 32 | 16 | 1.8934 | 2.1486 | 6
HGA | A | 3 | 34 | 1 | 13.9427 | 14.3236 | 6
HGA | A | 3 | 34 | 2 | 30.8745 | 16.1192 | 6
HGA | A | 3 | 34 | 3 | 29.7929 | 13.9285 | 6
HGA | A | 3 | 34 | 4 | 10.357 | 13.8227 | 6
HGA | A | 3 | 34 | 5 | 9.4169 | 13.4212 | 6
HGA | A | 3 | 34 | 6 | 10.6884 | 14.6274 | 6
HGA | A | 3 | 34 | 7 | 10.4735 | 16.0593 | 6
HGA | A | 3 | 34 | 8 | 10.5665 | 13.6398 | 6
HGA | A | 3 | 34 | 9 | 9.6145 | 15.0838 | 6
HGA | A | 3 | 34 | 10 | 10.3949 | 14.8643 | 6
HGA | A | 3 | 34 | 11 | 9.3888 | 14.1542 | 6
HGA | A | 3 | 34 | 12 | 14.106 | 15.0397 | 6
HGA | A | 3 | 34 | 13 | 10.5788 | 15.998 | 6
HGA | A | 3 | 34 | 14 | 10.6276 | 13.872 | 6
HGA | A | 3 | 34 | 15 | 10.2432 | 16.085 | 6
HGA | A | 3 | 34 | 16 | 11.1197 | 14.4293 | 6
HGA | B | 3 | 1 | 1 | 7.0472 | 9.8935 | 6
HGA | B | 3 | 1 | 2 | 6.7746 | 9.5851 | 6
HGA | B | 3 | 1 | 3 | 6.46 | 9.7853 | 6
HGA | B | 3 | 1 | 4 | 7.0224 | 10.2142 | 6
HGA | B | 3 | 1 | 5 | 7.0866 | 9.3223 | 6
HGA | B | 3 | 1 | 6 | 7.7398 | 9.9544 | 6
HGA | B | 3 | 1 | 7 | 7.2443 | 10.3532 | 6
HGA | B | 3 | 1 | 8 | 7.0797 | 9.7314 | 6
HGA | B | 3 | 1 | 9 | 7.2132 | 10.0344 | 6
HGA | B | 3 | 1 | 10 | 7.0511 | 11.0275 | 6
HGA | B | 3 | 1 | 11 | 7.2176 | 9.8864 | 6
HGA | B | 3 | 1 | 12 | 7.3605 | 10.9626 | 6
HGA | B | 3 | 1 | 13 | 7.3334 | 10.9722 | 6
HGA | B | 3 | 1 | 14 | 7.6677 | 10.6165 | 6
HGA | B | 3 | 1 | 15 | 7.5982 | 11.135 | 6
HGA | B | 3 | 1 | 16 | 7.7289 | 10.8596 | 6
HGA | B | 3 | 7 | 1 | 6.3646 | 7.3488 | 6
HGA | B | 3 | 7 | 2 | 6.1904 | 8.0069 | 6
HGA | B | 3 | 7 | 3 | 6.2548 | 7.9976 | 6
HGA | B | 3 | 7 | 4 | 6.3489 | 7.8018 | 6
HGA | B | 3 | 7 | 5 | 6.2527 | 8.2898 | 6
HGA | B | 3 | 7 | 6 | 6.461 | 8.2583 | 6
HGA | B | 3 | 7 | 7 | 6.5282 | 8.5084 | 6
HGA | B | 3 | 7 | 8 | 6.6494 | 8.0784 | 6
HGA | B | 3 | 7 | 9 | 6.7032 | 8.2941 | 6
HGA | B | 3 | 7 | 10 | 6.5081 | 8.0037 | 6
HGA | B | 3 | 7 | 11 | 6.457 | 8.6058 | 6
HGA | B | 3 | 7 | 12 | 6.7289 | 8.7032 | 6
HGA | B | 3 | 7 | 13 | 6.4065 | 8.0122 | 6
HGA | B | 3 | 7 | 14 | 6.5095 | 8.1511 | 6
HGA | B | 3 | 7 | 15 | 6.874 | 9.5085 | 6
HGA | B | 3 | 7 | 16 | 6.55 | 8.8372 | 6
HGA | B | 3 | 11 | 1 | 5.308 | 6.4844 | 6
HGA | B | 3 | 11 | 2 | 5.4147 | 6.2963 | 6
HGA | B | 3 | 11 | 3 | 5.4483 | 6.563 | 6
HGA | B | 3 | 11 | 4 | 5.3694 | 6.3795 | 6
HGA | B | 3 | 11 | 5 | 5.3981 | 6.801 | 6
HGA | B | 3 | 11 | 6 | 5.6783 | 7.0882 | 6
HGA | B | 3 | 11 | 7 | 5.3172 | 6.6508 | 6
HGA | B | 3 | 11 | 8 | 5.3727 | 7.0401 | 6
HGA | B | 3 | 11 | 9 | 5.3146 | 7.3224 | 6
HGA | B | 3 | 11 | 10 | 5.4598 | 7.3557 | 6
HGA | B | 3 | 11 | 11 | 5.4905 | 7.061 | 6
HGA | B | 3 | 11 | 12 | 5.3306 | 7.1851 | 6
HGA | B | 3 | 11 | 13 | 5.2417 | 6.7829 | 6
HGA | B | 3 | 11 | 14 | 5.3374 | 7.3547 | 6
HGA | B | 3 | 11 | 15 | 5.3902 | 7.1839 | 6
HGA | B | 3 | 11 | 16 | 5.3848 | 6.7564 | 6
HGA | B | 3 | 15 | 1 | 4.3986 | 5.2152 | 6
HGA | B | 3 | 15 | 2 | 4.4755 | 5.1941 | 6
HGA | B | 3 | 15 | 3 | 4.5657 | 5.4291 | 6
HGA | B | 3 | 15 | 4 | 4.4725 | 5.4874 | 6
HGA | B | 3 | 15 | 5 | 4.2828 | 5.3773 | 6
HGA | B | 3 | 15 | 6 | 4.4074 | 5.2756 | 6
HGA | B | 3 | 15 | 7 | 4.5019 | 5.5923 | 6
HGA | B | 3 | 15 | 8 | 4.5391 | 5.4614 | 6
HGA | B | 3 | 15 | 9 | 4.4283 | 6.2693 | 6
HGA | B | 3 | 15 | 10 | 4.4439 | 5.4088 | 6
HGA | B | 3 | 15 | 11 | 4.5575 | 5.3381 | 6
HGA | B | 3 | 15 | 12 | 4.5932 | 5.8398 | 6
HGA | B | 3 | 15 | 13 | 4.3859 | 5.4772 | 6
HGA | B | 3 | 15 | 14 | 4.5432 | 5.6482 | 6
HGA | B | 3 | 15 | 15 | 4.3975 | 5.6222 | 6
HGA | B | 3 | 15 | 16 | 4.5605 | 5.6697 | 6
HGA | B | 3 | 20 | 1 | 3.5559 | 4.2698 | 6
HGA | B | 3 | 20 | 2 | 3.6742 | 4.3637 | 6
HGA | B | 3 | 20 | 3 | 3.5201 | 4.6245 | 6
HGA | B | 3 | 20 | 4 | 3.5887 | 4.3967 | 6
HGA | B | 3 | 20 | 5 | 3.7887 | 4.9137 | 6
HGA | B | 3 | 20 | 6 | 3.7245 | 4.4666 | 6
HGA | B | 3 | 20 | 7 | 3.714 | 4.4993 | 6
HGA | B | 3 | 20 | 8 | 3.7076 | 4.3958 | 6
HGA | B | 3 | 20 | 9 | 3.8605 | 4.9442 | 6
HGA | B | 3 | 20 | 10 | 3.7794 | 4.4405 | 6
HGA | B | 3 | 20 | 11 | 3.9416 | 5.1595 | 6
HGA | B | 3 | 20 | 12 | 3.5558 | 4.6339 | 6
HGA | B | 3 | 20 | 13 | 3.8272 | 4.6493 | 6
HGA | B | 3 | 20 | 14 | 3.4433 | 4.7586 | 6
HGA | B | 3 | 20 | 15 | 3.6679 | 4.8558 | 6
HGA | B | 3 | 20 | 16 | 3.9063 | 4.9437 | 6
HGA | B | 3 | 26 | 1 | 2.8127 | 3.4137 | 6
HGA | B | 3 | 26 | 2 | 2.8941 | 3.3537 | 6
HGA | B | 3 | 26 | 3 | 2.8894 | 3.4125 | 6
HGA | B | 3 | 26 | 4 | 2.9587 | 3.7879 | 6
HGA | B | 3 | 26 | 5 | 3.0518 | 3.4687 | 6
HGA | B | 3 | 26 | 6 | 2.9309 | 3.6235 | 6
HGA | B | 3 | 26 | 7 | 2.9694 | 3.6337 | 6
HGA | B | 3 | 26 | 8 | 2.9236 | 3.4963 | 6
HGA | B | 3 | 26 | 9 | 2.9271 | 3.377 | 6
HGA | B | 3 | 26 | 10 | 3.0085 | 3.6988 | 6
HGA | B | 3 | 26 | 11 | 3.0304 | 3.564 | 6
HGA | B | 3 | 26 | 12 | 2.9421 | 3.4016 | 6
HGA | B | 3 | 26 | 13 | 2.9603 | 3.4118 | 6
HGA | B | 3 | 26 | 14 | 3.0757 | 3.5768 | 6
HGA | B | 3 | 26 | 15 | 2.9349 | 3.5195 | 6
HGA | B | 3 | 26 | 16 | 2.8915 | 3.6252 | 6
HGA | B | 3 | 32 | 1 | 1.8654 | 2.0652 | 5
HGA | B | 3 | 32 | 2 | 1.9701 | 2.1821 | 5
HGA | B | 3 | 32 | 3 | 1.8771 | 2.1594 | 5
HGA | B | 3 | 32 | 4 | 2.4726 | 3.0295 | 5
HGA | B | 3 | 32 | 5 | 1.8384 | 2.3002 | 5
HGA | B | 3 | 32 | 6 | 2.1341 | 2.4408 | 5
HGA | B | 3 | 32 | 7 | 2.1754 | 2.4487 | 5
HGA | B | 3 | 32 | 8 | 2.3305 | 2.529 | 5
HGA | B | 3 | 32 | 9 | 2.2113 | 2.7745 | 5
HGA | B | 3 | 32 | 10 | 2.2315 | 2.4208 | 5
HGA | B | 3 | 32 | 11 | 2.0351 | 2.4675 | 5
HGA | B | 3 | 32 | 12 | 1.9826 | 2.0675 | 5
HGA | B | 3 | 32 | 13 | 2.5757 | 3.244 | 5
HGA | B | 3 | 32 | 14 | 1.8212 | 2.0174 | 5
HGA | B | 3 | 32 | 15 | 1.9353 | 2.164 | 5
HGA | B | 3 | 32 | 16 | 1.9016 | 2.1494 | 5
HGA | B | 3 | 34 | 1 | 14.1264 | 14.7783 | ~
HGA | B | 3 | 34 | 2 | 14.015 | 15.4995 | ~
HGA | B | 3 | 34 | 3 | 9.5465 | 12.4005 | ~
HGA | B | 3 | 34 | 4 | 9.5937 | 15.1105 | ~
HGA | B | 3 | 34 | 5 | 10.7405 | 14.7502 | ~
HGA | B | 3 | 34 | 6 | 14.001 | 13.9596 | ~
HGA | B | 3 | 34 | 7 | 11.0853 | 14.9854 | ~
HGA | B | 3 | 34 | 8 | 10.2317 | 14.0727 | ~
HGA | B | 3 | 34 | 9 | 10.657 | 17.3338 | ~
HGA | B | 3 | 34 | 10 | 10.4383 | 14.9015 | ~
HGA | B | 3 | 34 | 11 | 11.228 | 15.0567 | ~
HGA | B | 3 | 34 | 12 | 9.8715 | 15.0188 | ~
HGA | B | 3 | 34 | 13 | 10.5104 | 16.1498 | ~
HGA | B | 3 | 34 | 14 | 9.5021 | 15.3681 | ~
HGA | B | 3 | 34 | 15 | 9.9614 | 15.2329 | ~
HGA | B | 3 | 34 | 16 | 10.4324 | 15.7135 | ~
HGB | A | 3 | 1 | 1 | 6.7743 | 10.0873 | 6
HGB | A | 3 | 1 | 2 | 6.6757 | 9.5374 | 6
HGB | A | 3 | 1 | 3 | 6.7936 | 9.3599 | 6
HGB | A | 3 | 1 | 4 | 6.9964 | 9.8448 | 6
HGB | A | 3 | 1 | 5 | 6.8728 | 9.6262 | 6
HGB | A | 3 | 1 | 6 | 6.9471 | 10.8944 | 6
HGB | A | 3 | 1 | 7 | 6.8722 | 10.5973 | 6
HGB | A | 3 | 1 | 8 | 7.1181 | 10.2315 | 6
HGB | A | 3 | 1 | 9 | 6.6534 | 9.7562 | 6
HGB | A | 3 | 1 | 10 | 7.0531 | 9.3393 | 6
HGB | A | 3 | 1 | 11 | 7.0125 | 9.7309 | 6
HGB | A | 3 | 1 | 12 | 6.683 | 10.3518 | 6
HGB | A | 3 | 1 | 13 | 6.9523 | 10.0419 | 6
HGB | A | 3 | 1 | 14 | 7.2521 | 10.8436 | 6
HGB | A | 3 | 1 | 15 | 6.9162 | 10.3569 | 6
HGB | A | 3 | 1 | 16 | 7.3612 | 10.4824 | 6
HGB | A | 3 | 7 | 1 | 6.2681 | 7.4065 | 6
HGB | A | 3 | 7 | 2 | 6.0213 | 8.3063 | 6
HGB | A | 3 | 7 | 3 | 6.0807 | 8.0348 | 6
HGB | A | 3 | 7 | 4 | 6.2335 | 9.1161 | 6
HGB | A | 3 | 7 | 5 | 6.4157 | 7.9298 | 6
HGB | A | 3 | 7 | 6 | 6.4744 | 7.8976 | 6
HGB | A | 3 | 7 | 7 | 6.3803 | 7.903 | 6
HGB | A | 3 | 7 | 8 | 6.504 | 8.2595 | 6
HGB | A | 3 | 7 | 9 | 6.2407 | 8.4252 | 6
HGB | A | 3 | 7 | 10 | 6.2498 | 8.2512 | 6
HGB | A | 3 | 7 | 11 | 6.125 | 8.3255 | 6
HGB | A | 3 | 7 | 12 | 6.3371 | 7.484 | 6
HGB | A | 3 | 7 | 13 | 6.3875 | 8.0745 | 6
HGB | A | 3 | 7 | 14 | 6.454 | 8.2474 | 6
HGB | A | 3 | 7 | 15 | 6.5553 | 8.0255 | 6
HGB | A | 3 | 7 | 16 | 6.7107 | 8.3366 | 6
HGB | A | 3 | 11 | 1 | 5.4116 | 6.5233 | 6
HGB | A | 3 | 11 | 2 | 5.4697 | 6.7208 | 6
HGB | A | 3 | 11 | 3 | 5.0639 | 6.5898 | 6
HGB | A | 3 | 11 | 4 | 5.0903 | 6.8681 | 6
HGB | A | 3 | 11 | 5 | 5.3136 | 7.2353 | 6
HGB | A | 3 | 11 | 6 | 5.3879 | 6.9035 | 6
HGB | A | 3 | 11 | 7 | 5.217 | 7.3793 | 6
HGB | A | 3 | 11 | 8 | 5.679 | 7.0986 | 6
HGB | A | 3 | 11 | 9 | 5.1632 | 6.9363 | 6
HGB | A | 3 | 11 | 10 | 5.4967 | 6.6924 | 6
HGB | A | 3 | 11 | 11 | 5.2499 | 6.5038 | 6
HGB | A | 3 | 11 | 12 | 5.4054 | 6.691 | 6
HGB | A | 3 | 11 | 13 | 5.1341 | 6.748 | 6
HGB | A | 3 | 11 | 14 | 5.2266 | 6.6185 | 6
HGB | A | 3 | 11 | 15 | 5.3936 | 6.9889 | 6
HGB | A | 3 | 11 | 16 | 5.4432 | 6.8865 | 6
HGB | A | 3 | 15 | 1 | 4.383 | 5.2879 | 6
HGB | A | 3 | 15 | 2 | 4.1478 | 5.174 | 6
HGB | A | 3 | 15 | 3 | 4.3262 | 5.263 | 6
HGB | A | 3 | 15 | 4 | 4.3509 | 5.2916 | 6
HGB | A | 3 | 15 | 5 | 4.5084 | 5.3409 | 6
HGB | A | 3 | 15 | 6 | 4.5535 | 5.5246 | 6
HGB | A | 3 | 15 | 7 | 4.4972 | 5.5931 | 6
HGB | A | 3 | 15 | 8 | 4.7163 | 5.0291 | 6
HGB | A | 3 | 15 | 9 | 4.5889 | 5.8033 | 6
HGB | A | 3 | 15 | 10 | 4.3027 | 4.995 | 6
HGB | A | 3 | 15 | 11 | 4.2506 | 5.5377 | 6
HGB | A | 3 | 15 | 12 | 4.0181 | 5.3236 | 6
HGB | A | 3 | 15 | 13 | 4.116 | 5.3334 | 6
HGB | A | 3 | 15 | 14 | 4.2069 | 5.1299 | 6
HGB | A | 3 | 15 | 15 | 4.3004 | 5.4823 | 6
HGB | A | 3 | 15 | 16 | 4.3328 | 5.3861 | 6
HGB | A | 3 | 20 | 1 | 3.5239 | 4.1855 | 6
HGB | A | 3 | 20 | 2 | 3.7111 | 4.439 | 6
HGB | A | 3 | 20 | 3 | 3.6401 | 4.3509 | 6
HGB | A | 3 | 20 | 4 | 3.7282 | 4.4993 | 6
HGB | A | 3 | 20 | 5 | 3.6423 | 4.5468 | 6
HGB | A | 3 | 20 | 6 | 3.7896 | 4.5701 | 6
HGB | A | 3 | 20 | 7 | 3.7625 | 4.4372 | 6
HGB | A | 3 | 20 | 8 | 3.6966 | 4.5748 | 6
HGB | A | 3 | 20 | 9 | 3.7508 | 4.6562 | 6
HGB | A | 3 | 20 | 10 | 3.613 | 4.5115 | 6
HGB | A | 3 | 20 | 11 | 3.7288 | 4.4654 | 6
HGB | A | 3 | 20 | 12 | 3.4834 | 4.4043 | 6
HGB | A | 3 | 20 | 13 | 3.536 | 4.0001 | 6
HGB | A | 3 | 20 | 14 | 3.6465 | 4.4745 | 6
HGB | A | 3 | 20 | 15 | 3.6422 | 4.2935 | 6
HGB | A | 3 | 20 | 16 | 3.6379 | 4.3754 | 6
HGB | A | 3 | 26 | 1 | 2.887 | 3.4652 | 6
HGB | A | 3 | 26 | 2 | 2.8026 | 3.2092 | 6
HGB | A | 3 | 26 | 3 | 2.9876 | 3.4171 | 6
HGB | A | 3 | 26 | 4 | 2.7591 | 3.3857 | 6
HGB | A | 3 | 26 | 5 | 2.8332 | 3.5343 | 6
HGB | A | 3 | 26 | 6 | 2.8782 | 3.4736 | 6
HGB | A | 3 | 26 | 7 | 2.6547 | 3.3901 | 6
HGB | A | 3 | 26 | 8 | 2.854 | 3.3178 | 6
HGB | A | 3 | 26 | 9 | 2.9901 | 3.4534 | 6
HGB | A | 3 | 26 | 10 | 2.7762 | 3.5174 | 6
HGB | A | 3 | 26 | 11 | 2.8643 | 3.5279 | 6
HGB | A | 3 | 26 | 12 | 2.9342 | 3.4324 | 6
HGB | A | 3 | 26 | 13 | 2.7475 | 3.379 | 6
HGB | A | 3 | 26 | 14 | 2.8786 | 3.1906 | 6
HGB | A | 3 | 26 | 15 | 2.8717 | 3.4235 | 6
HGB | A | 3 | 26 | 16 | 2.7537 | 3.3793 | 6
HGB | A | 3 | 32 | 1 | 1.8331 | 2.0178 | 6
HGB | A | 3 | 32 | 2 | 1.9718 | 2.0518 | 6
HGB | A | 3 | 32 | 3 | 1.8273 | 1.9695 | 6
HGB | A | 3 | 32 | 4 | 2.3015 | 2.7468 | 6
HGB | A | 3 | 32 | 5 | 1.9455 | 2.277 | 6
HGB | A | 3 | 32 | 6 | 1.9738 | 2.2767 | 6
HGB | A | 3 | 32 | 7 | 2.2353 | 2.6156 | 6
HGB | A | 3 | 32 | 8 | 2.2121 | 2.4619 | 6
HGB | A | 3 | 32 | 9 | 2.1924 | 2.5525 | 6
HGB | A | 3 | 32 | 10 | 2.1535 | 2.4855 | 6
HGB | A | 3 | 32 | 11 | 1.9367 | 2.2077 | 6
HGB | A | 3 | 32 | 12 | 1.9201 | 2.0797 | 6
HGB | A | 3 | 32 | 13 | 2.3311 | 2.7033 | 6
HGB | A | 3 | 32 | 14 | 1.8401 | 2.0032 | 6
HGB | A | 3 | 32 | 15 | 1.8675 | 2.1181 | 6
HGB | A | 3 | 32 | 16 | 1.8836 | 1.9485 | 6
HGB | A | 3 | 34 | 1 | 9.8417 | 14.7117 | 6
HGB | A | 3 | 34 | 2 | 9.929 | 15.0078 | 6
HGB | A | 3 | 34 | 3 | 9.1769 | 13.537 | 6
HGB | A | 3 | 34 | 4 | 9.8534 | 15.8666 | 6
HGB | A | 3 | 34 | 5 | 9.1055 | 15.1079 | 6
HGB | A | 3 | 34 | 6 | 14.0923 | 14.5958 | 6
HGB | A | 3 | 34 | 7 | 8.0805 | 15.6165 | 6
HGB | A | 3 | 34 | 8 | 11.16 | 14.278 | 6
HGB | A | 3 | 34 | 9 | 10.8008 | 13.44 | 6
HGB | A | 3 | 34 | 10 | 10.4138 | 14.9215 | 6
HGB | A | 3 | 34 | 11 | 29.7635 | 15.35 | 6
HGB | A | 3 | 34 | 12 | 14.2683 | 15.0704 | 6
HGB | A | 3 | 34 | 13 | 14.2532 | 13.9269 | 6
HGB | A | 3 | 34 | 14 | 10.2554 | 15.456 | 6
HGB | A | 3 | 34 | 15 | 9.4771 | 15.5562 | 6
HGB | A | 3 | 34 | 16 | 10.1062 | 14.4638 | 6
HGB | B | 3 | 1 | 1 | 6.9385 | 9.89 | 6
HGB | B | 3 | 1 | 2 | 7.0967 | 10.1722 | 6
HGB | B | 3 | 1 | 3 | 7.1029 | 9.2849 | 6
HGB | B | 3 | 1 | 4 | 7.2663 | 9.9661 | 6
HGB | B | 3 | 1 | 5 | 6.9671 | 10.1783 | 6
HGB | B | 3 | 1 | 6 | 7.216 | 9.6851 | 6
HGB | B | 3 | 1 | 7 | 7.2162 | 10.4997 | 6
HGB | B | 3 | 1 | 8 | 7.1946 | 10.3097 | 6
HGB | B | 3 | 1 | 9 | 6.9881 | 10.3589 | 6
HGB | B | 3 | 1 | 10 | 6.7674 | 10.0352 | 6
HGB | B | 3 | 1 | 11 | 6.7486 | 10.6889 | 6
HGB | B | 3 | 1 | 12 | 7.2777 | 10.2149 | 6
HGB | B | 3 | 1 | 13 | 7.4676 | 10.1802 | 6
HGB | B | 3 | 1 | 14 | 7.0585 | 9.6912 | 6
HGB | B | 3 | 1 | 15 | 6.8602 | 10.0496 | 6
HGB | B | 3 | 1 | 16 | 7.2255 | 10.524 | 6
HGB | B | 3 | 7 | 1 | 5.9021 | 7.529 | 6
HGB | B | 3 | 7 | 2 | 6.1076 | 7.3686 | 6
HGB | B | 3 | 7 | 3 | 6.6906 | 8.1924 | 6
HGB | B | 3 | 7 | 4 | 6.268 | 8.2894 | 6
HGB | B | 3 | 7 | 5 | 6.5513 | 7.7968 | 6
HGB | B | 3 | 7 | 6 | 6.4726 | 7.5927 | 6
HGB | B | 3 | 7 | 7 | 6.7117 | 7.8277 | 6
HGB | B | 3 | 7 | 8 | 5.9549 | 7.6157 | 6
HGB | B | 3 | 7 | 9 | 6.274 | 7.5761 | 6
HGB | B | 3 | 7 | 10 | 6.4302 | 8.3953 | 6
HGB | B | 3 | 7 | 11 | 6.1873 | 7.6926 | 6
HGB | B | 3 | 7 | 12 | 6.2453 | 8.387 | 6
HGB | B | 3 | 7 | 13 | 6.6459 | 7.3104 | 6
HGB | B | 3 | 7 | 14 | 6.2661 | 8.5325 | 6
HGB | B | 3 | 7 | 15 | 6.3469 | 8.2678 | 6
HGB | B | 3 | 7 | 16 | 6.3369 | 7.7931 | 6
HGB | B | 3 | 11 | 1 | 5.2453 | 6.5775 | 6
HGB | B | 3 | 11 | 2 | 5.2861 | 6.9426 | 6
HGB | B | 3 | 11 | 3 | 5.2767 | 6.6775 | 6
HGB | B | 3 | 11 | 4 | 5.3288 | 6.9046 | 6
HGB | B | 3 | 11 | 5 | 5.1061 | 6.7958 | 6
HGB | B | 3 | 11 | 6 | 5.1521 | 6.2844 | 6
HGB | B | 3 | 11 | 7 | 5.1156 | 7.0665 | 6
HGB | B | 3 | 11 | 8 | 5.1766 | 6.7118 | 6
HGB | B | 3 | 11 | 9 | 5.1983 | 7.2016 | 6
HGB | B | 3 | 11 | 10 | 4.8777 | 6.6469 | 6
HGB | B | 3 | 11 | 11 | 5.2294 | 6.2693 | 6
HGB | B | 3 | 11 | 12 | 4.9817 | 6.4524 | 6
HGB | B | 3 | 11 | 13 | 5.2813 | 6.5875 | 6
HGB | B | 3 | 11 | 14 | 5.4723 | 6.312 | 6
HGB | B | 3 | 11 | 15 | 5.27 | 6.8295 | 6
HGB | B | 3 | 11 | 16 | 5.5035 | 6.5453 | 6
HGB | B | 3 | 15 | 1 | 4.4372 | 4.8265 | 6
HGB | B | 3 | 15 | 2 | 4.1336 | 5.3237 | 6
HGB | B | 3 | 15 | 3 | 4.3878 | 5.4439 | 6
HGB | B | 3 | 15 | 4 | 4.2137 | 5.529 | 6
HGB | B | 3 | 15 | 5 | 4.228 | 5.5521 | 6
HGB | B | 3 | 15 | 6 | 4.1968 | 5.633 | 6
HGB | B | 3 | 15 | 7 | 4.5842 | 5.6607 | 6
HGB | B | 3 | 15 | 8 | 4.6554 | 5.7425 | 6
HGB | B | 3 | 15 | 9 | 4.6727 | 5.23 | 6
HGB | B | 3 | 15 | 10 | 4.4819 | 5.4482 | 6
HGB | B | 3 | 15 | 11 | 4.1208 | 5.0835 | 6
HGB | B | 3 | 15 | 12 | 4.4388 | 5.4242 | 6
HGB | B | 3 | 15 | 13 | 4.3211 | 5.0674 | 6
HGB | B | 3 | 15 | 14 | 4.3091 | 5.3548 | 6
HGB | B | 3 | 15 | 15 | 4.5655 | 5.5334 | 6
HGB | B | 3 | 15 | 16 | 4.4481 | 5.5767 | 6
HGB | B | 3 | 20 | 1 | 3.4846 | 4.1436 | 6
HGB | B | 3 | 20 | 2 | 3.5792 | 4.2618 | 6
HGB | B | 3 | 20 | 3 | 3.6708 | 4.6007 | 6
HGB | B | 3 | 20 | 4 | 3.683 | 4.5785 | 6
HGB | B | 3 | 20 | 5 | 3.7728 | 4.4223 | 6
HGB | B | 3 | 20 | 6 | 3.5673 | 4.398 | 6
HGB | B | 3 | 20 | 7 | 3.6064 | 4.3921 | 6
HGB | B | 3 | 20 | 8 | 3.6419 | 4.4006 | 6
HGB | B | 3 | 20 | 9 | 3.711 | 4.562 | 6
HGB | B | 3 | 20 | 10 | 3.7142 | 4.6087 | 6
HGB | B | 3 | 20 | 11 | 3.562 | 4.6421 | 6
HGB | B | 3 | 20 | 12 | 3.5877 | 4.6255 | 6
HGB | B | 3 | 20 | 13 | 3.6317 | 4.5769 | 6
HGB | B | 3 | 20 | 14 | 3.6671 | 4.5743 | 6
HGB | B | 3 | 20 | 15 | 3.5147 | 4.5295 | 6
HGB | B | 3 | 20 | 16 | 3.6035 | 4.5018 | 6
HGB | B | 3 | 26 | 1 | 2.8152 | 3.1781 | 6
HGB | B | 3 | 26 | 2 | 2.8378 | 3.4795 | 6
HGB | B | 3 | 26 | 3 | 2.8942 | 3.3094 | 6
HGB | B | 3 | 26 | 4 | 2.9985 | 3.4454 | 6
HGB | B | 3 | 26 | 5 | 3.02 | 3.6017 | 6
HGB | B | 3 | 26 | 6 | 2.9738 | 3.537 | 6
HGB | B | 3 | 26 | 7 | 2.853 | 3.4944 | 6
HGB | B | 3 | 26 | 8 | 2.8963 | 3.7705 | 6
HGB | B | 3 | 26 | 9 | 2.9193 | 3.3656 | 6
HGB | B | 3 | 26 | 10 | 2.7632 | 3.3291 | 6
HGB | B | 3 | 26 | 11 | 2.9622 | 3.4892 | 6
HGB | B | 3 | 26 | 12 | 2.8444 | 3.427 | 6
HGB | B | 3 | 26 | 13 | 2.9203 | 3.353 | 6
HGB | B | 3 | 26 | 14 | 2.8602 | 3.2439 | 6
HGB | B | 3 | 26 | 15 | 2.856 | 3.4233 | 6
HGB | B | 3 | 26 | 16 | 2.8378 | 3.3866 | 6
HGB | B | 3 | 32 | 1 | 1.8072 | 2.042 | 5
HGB | B | 3 | 32 | 2 | 1.8147 | 2.0586 | 5
HGB | B | 3 | 32 | 3 | 1.8652 | 2.0972 | 5
HGB | B | 3 | 32 | 4 | 2.2432 | 2.8071 | 5
HGB | B | 3 | 32 | 5 | 1.9301 | 2.2328 | 5
HGB | B | 3 | 32 | 6 | 2.0621 | 2.2603 | 5
HGB | B | 3 | 32 | 7 | 2.0833 | 2.4009 | 5
HGB | B | 3 | 32 | 8 | 2.241 | 2.6232 | 5
HGB | B | 3 | 32 | 9 | 2.1453 | 2.6081 | 5
HGB | B | 3 | 32 | 10 | 2.1422 | 2.3661 | 5
HGB | B | 3 | 32 | 11 | 1.951 | 2.3064 | 5
HGB | B | 3 | 32 | 12 | 1.9118 | 2.1654 | 5
HGB | B | 3 | 32 | 13 | 2.2112 | 2.574 | 5
HGB | B | 3 | 32 | 14 | 1.8862 | 2.1076 | 5
HGB | B | 3 | 32 | 15 | 1.9026 | 2.0581 | 5
HGB | B | 3 | 32 | 16 | 1.8154 | 1.9592 | 5
HGB | B | 3 | 34 | 1 | 10.2201 | 15.1455 | ~
HGB | B | 3 | 34 | 2 | 10.1561 | 14.1864 | ~
HGB | B | 3 | 34 | 3 | 9.8827 | 13.0318 | ~
HGB | B | 3 | 34 | 4 | 9.8385 | 14.3814 | ~
HGB | B | 3 | 34 | 5 | 10.3764 | 15.7646 | ~
HGB | B | 3 | 34 | 6 | 16.2231 | 13.9794 | ~
HGB | B | 3 | 34 | 7 | 9.766 | 15.3207 | ~
HGB | B | 3 | 34 | 8 | 10.7047 | 15.0144 | ~
HGB | B | 3 | 34 | 9 | 10.0682 | 13.0421 | ~
HGB | B | 3 | 34 | 10 | 14.2523 | 16.3019 | ~
HGB | B | 3 | 34 | 11 | 13.9218 | 14.021 | ~
HGB | B | 3 | 34 | 12 | 13.8318 | 13.9929 | ~
HGB | B | 3 | 34 | 13 | 14.2673 | 12.9199 | ~
HGB | B | 3 | 34 | 14 | 10.587 | 15.8111 | ~
HGB | B | 3 | 34 | 15 | 10.1436 | 13.8464 | ~
HGB | B | 3 | 34 | 16 | 10.7315 | 15.1283 | ~

## TEB saturation
Band | HAM | Subsample | Gain State | Band Median |  | Band Average |  | Tmax
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | Tsat | Tsat | Tsat | Tsat
I4 | A | 1 | H | 348.806 | 349.9174 | 348.7391 | 349.9503 | 353
I4 | A | 2 | H | 349.4774 | 351.311 | 349.3351 | 351.3232 | 353
I4 | B | 1 | H | 348.8093 | 349.9246 | 348.7496 | 349.9625 | 353
I4 | B | 2 | H | 349.4369 | 351.1202 | 349.3195 | 351.1855 | 353
I5 | A | 1 | H | 369.7993 | 373.8215 | 370.4755 | 374.4488 | 340
I5 | A | 2 | H | 369.514 | 373.6201 | 370.2384 | 374.2511 | 340
I5 | B | 1 | H | 370.0017 | 374.1845 | 370.5783 | 374.7833 | 340
I5 | B | 2 | H | 369.7592 | 374.2053 | 370.3776 | 374.5828 | 340
M12 | A | 1 | H | 357.1049 | 357.3455 | 357.1002 | 357.3532 | 353
M12 | B | 1 | H | 357.1053 | 357.3705 | 357.1002 | 357.3771 | 353
M13 | A | 1 | H | 350.6962 | 349.4265 | 350.851 | 349.5532 | 343
M13 | B | 1 | H | 350.5882 | 349.4393 | 350.7461 | 349.5669 | 343
M14 | A | 1 | H | 348.7085 | 352.0082 | 348.902 | 352.3494 | 336
M14 | B | 1 | H | 348.698 | 351.994 | 348.8797 | 352.3403 | 336
M15 | A | 1 | H | 358.0674 | 360.4138 | 357.4933 | 359.6346 | 343
M15 | B | 1 | H | 358.275 | 360.7188 | 357.708 | 359.9067 | 343
M16A | A | 1 | H | 368.2274 | 371.7328 | 368.3785 | 372.1552 | 340
M16A | B | 1 | H | 368.4895 | 372.0875 | 368.5917 | 372.4935 | 340
M16B | A | 1 | H | 369.2489 | 373.5661 | 369.2914 | 373.5176 | 340
M16B | B | 1 | H | 369.4968 | 373.8362 | 369.4336 | 373.7736 | 340
M13 | A | 1 | L | 657.4777 | 658.5386 | 657.6471 | 658.7832 | 634
M13 | B | 1 | L | 657.4927 | 658.5866 | 657.6628 | 658.8309 | 634
Band | HAM | Detector | Subsample | Gain State | Detector Dependent |  | Tmax |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 TEB saturation.
[5 empty] | Tsat | Tsat |  |  | 2. The results were derived from RC-1 using the TMC source.
I4 | A | 1 | 1 | H | 349.2388 | 350.4485 | 353 |  | 3. Saturation was determined from partial views of the TMC source.
I4 | A | 2 | 1 | H | 348.7916 | 350.0681 | 353 |  | 4. Tsat and Tmax values are in K.
I4 | A | 3 | 1 | H | 347.9993 | 349.4268 | 353 |  | 5. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I4 | A | 4 | 1 | H | 348.3879 | 349.9339 | 353 |  | 6. The test was set to operational day mode.
I4 | A | 5 | 1 | H | 348.3109 | 349.8628 | 353
I4 | A | 6 | 1 | H | 349.5057 | 350.6718 | 353
I4 | A | 7 | 1 | H | 348.6062 | 349.8237 | 353
I4 | A | 8 | 1 | H | 349.1538 | 350.3349 | 353
I4 | A | 9 | 1 | H | 348.8736 | 350.0376 | 353
I4 | A | 10 | 1 | H | 349.0774 | 350.4209 | 353
I4 | A | 11 | 1 | H | 349.0291 | 350.1628 | 353
I4 | A | 12 | 1 | H | 348.9906 | 350.0618 | 353
I4 | A | 13 | 1 | H | 349.0581 | 350.0762 | 353
I4 | A | 14 | 1 | H | 348.7028 | 349.828 | 353
I4 | A | 15 | 1 | H | 348.838 | 349.8783 | 353
I4 | A | 16 | 1 | H | 348.7118 | 349.8684 | 353
I4 | A | 17 | 1 | H | 348.8717 | 350.0492 | 353
I4 | A | 18 | 1 | H | 348.6434 | 349.7961 | 353
I4 | A | 19 | 1 | H | 348.7428 | 349.8454 | 353
I4 | A | 20 | 1 | H | 349.2369 | 350.3585 | 353
I4 | A | 21 | 1 | H | 348.8205 | 349.9009 | 353
I4 | A | 22 | 1 | H | 348.7047 | 349.8952 | 353
I4 | A | 23 | 1 | H | 349.0844 | 350.1779 | 353
I4 | A | 24 | 1 | H | 348.1081 | 349.3966 | 353
I4 | A | 25 | 1 | H | 348.6671 | 349.8697 | 353
I4 | A | 26 | 1 | H | 348.3845 | 349.6029 | 353
I4 | A | 27 | 1 | H | 348.15 | 349.4381 | 353
I4 | A | 28 | 1 | H | 347.8046 | 349.2044 | 353
I4 | A | 29 | 1 | H | 347.9478 | 349.3012 | 353
I4 | A | 30 | 1 | H | 349.0476 | 350.2109 | 353
I4 | A | 31 | 1 | H | 349.1535 | 350.2706 | 353
I4 | A | 32 | 1 | H | 349.0082 | 350.186 | 353
I4 | A | 1 | 2 | H | 349.2561 | 351.6716 | 353
I4 | A | 2 | 2 | H | 348.908 | 351.3609 | 353
I4 | A | 3 | 2 | H | 348.49 | 350.923 | 353
I4 | A | 4 | 2 | H | 349.0721 | 351.1822 | 353
I4 | A | 5 | 2 | H | 348.8619 | 351.1353 | 353
I4 | A | 6 | 2 | H | 349.9532 | 351.8873 | 353
I4 | A | 7 | 2 | H | 349.1774 | 351.2413 | 353
I4 | A | 8 | 2 | H | 349.8109 | 351.5489 | 353
I4 | A | 9 | 2 | H | 349.4134 | 351.3684 | 353
I4 | A | 10 | 2 | H | 349.7977 | 351.4825 | 353
I4 | A | 11 | 2 | H | 349.6818 | 351.4575 | 353
I4 | A | 12 | 2 | H | 349.7098 | 351.3268 | 353
I4 | A | 13 | 2 | H | 349.6711 | 351.3932 | 353
I4 | A | 14 | 2 | H | 349.5194 | 351.1537 | 353
I4 | A | 15 | 2 | H | 349.566 | 351.2139 | 353
I4 | A | 16 | 2 | H | 349.4972 | 351.2923 | 353
I4 | A | 17 | 2 | H | 349.5583 | 351.3268 | 353
I4 | A | 18 | 2 | H | 349.5094 | 351.2292 | 353
I4 | A | 19 | 2 | H | 349.5084 | 351.1849 | 353
I4 | A | 20 | 2 | H | 349.9066 | 351.6041 | 353
I4 | A | 21 | 2 | H | 349.4951 | 351.2791 | 353
I4 | A | 22 | 2 | H | 349.4597 | 351.2952 | 353
I4 | A | 23 | 2 | H | 349.5646 | 351.5094 | 353
I4 | A | 24 | 2 | H | 348.9509 | 350.9958 | 353
I4 | A | 25 | 2 | H | 349.3734 | 351.3443 | 353
I4 | A | 26 | 2 | H | 349.0981 | 351.168 | 353
I4 | A | 27 | 2 | H | 348.9464 | 350.9839 | 353
I4 | A | 28 | 2 | H | 348.5721 | 350.7637 | 353
I4 | A | 29 | 2 | H | 348.3993 | 350.8898 | 353
I4 | A | 30 | 2 | H | 349.031 | 351.7353 | 353
I4 | A | 31 | 2 | H | 349.3233 | 351.7515 | 353
I4 | A | 32 | 2 | H | 349.6413 | 351.6438 | 353
I4 | B | 1 | 1 | H | 349.2922 | 350.4537 | 353
I4 | B | 2 | 1 | H | 348.8002 | 350.1119 | 353
I4 | B | 3 | 1 | H | 348.0313 | 349.4866 | 353
I4 | B | 4 | 1 | H | 348.3865 | 349.8129 | 353
I4 | B | 5 | 1 | H | 348.3158 | 349.8209 | 353
I4 | B | 6 | 1 | H | 349.5475 | 350.6813 | 353
I4 | B | 7 | 1 | H | 348.5889 | 349.8578 | 353
I4 | B | 8 | 1 | H | 349.1472 | 350.3097 | 353
I4 | B | 9 | 1 | H | 348.9153 | 350.0771 | 353
I4 | B | 10 | 1 | H | 349.0472 | 350.1847 | 353
I4 | B | 11 | 1 | H | 349.0501 | 350.1682 | 353
I4 | B | 12 | 1 | H | 348.9826 | 350.0694 | 353
I4 | B | 13 | 1 | H | 349.082 | 350.0834 | 353
I4 | B | 14 | 1 | H | 348.7153 | 349.8656 | 353
I4 | B | 15 | 1 | H | 348.8686 | 349.8922 | 353
I4 | B | 16 | 1 | H | 348.7059 | 349.8971 | 353
I4 | B | 17 | 1 | H | 348.8942 | 350.0612 | 353
I4 | B | 18 | 1 | H | 348.6648 | 349.8562 | 353
I4 | B | 19 | 1 | H | 348.776 | 349.8945 | 353
I4 | B | 20 | 1 | H | 349.2544 | 350.367 | 353
I4 | B | 21 | 1 | H | 348.8184 | 349.9109 | 353
I4 | B | 22 | 1 | H | 348.6938 | 349.8962 | 353
I4 | B | 23 | 1 | H | 349.0802 | 350.2089 | 353
I4 | B | 24 | 1 | H | 348.0885 | 349.4645 | 353
I4 | B | 25 | 1 | H | 348.6813 | 349.9383 | 353
I4 | B | 26 | 1 | H | 348.377 | 349.632 | 353
I4 | B | 27 | 1 | H | 348.1662 | 349.4298 | 353
I4 | B | 28 | 1 | H | 347.8002 | 349.2653 | 353
I4 | B | 29 | 1 | H | 347.9724 | 349.3195 | 353
I4 | B | 30 | 1 | H | 349.0491 | 350.2241 | 353
I4 | B | 31 | 1 | H | 349.1663 | 350.3075 | 353
I4 | B | 32 | 1 | H | 349.0284 | 350.2522 | 353
I4 | B | 1 | 2 | H | 349.2543 | 351.4484 | 353
I4 | B | 2 | 2 | H | 349.0954 | 351.2114 | 353
I4 | B | 3 | 2 | H | 348.5436 | 350.8244 | 353
I4 | B | 4 | 2 | H | 349.0054 | 351.0889 | 353
I4 | B | 5 | 2 | H | 348.8329 | 350.9911 | 353
I4 | B | 6 | 2 | H | 349.9943 | 351.5975 | 353
I4 | B | 7 | 2 | H | 349.2118 | 351.0713 | 353
I4 | B | 8 | 2 | H | 349.7557 | 351.4185 | 353
I4 | B | 9 | 2 | H | 349.4318 | 351.1577 | 353
I4 | B | 10 | 2 | H | 349.754 | 351.3206 | 353
I4 | B | 11 | 2 | H | 349.5742 | 351.2473 | 353
I4 | B | 12 | 2 | H | 349.7039 | 351.2094 | 353
I4 | B | 13 | 2 | H | 349.6515 | 351.0867 | 353
I4 | B | 14 | 2 | H | 349.5338 | 351.0217 | 353
I4 | B | 15 | 2 | H | 349.5335 | 351.0067 | 353
I4 | B | 16 | 2 | H | 349.5064 | 351.1109 | 353
I4 | B | 17 | 2 | H | 349.5299 | 351.0688 | 353
I4 | B | 18 | 2 | H | 349.468 | 351.1277 | 353
I4 | B | 19 | 2 | H | 349.4419 | 351.1126 | 353
I4 | B | 20 | 2 | H | 349.8911 | 351.4706 | 353
I4 | B | 21 | 2 | H | 349.4574 | 351.0987 | 353
I4 | B | 22 | 2 | H | 349.4806 | 351.1068 | 353
I4 | B | 23 | 2 | H | 349.6364 | 351.3054 | 353
I4 | B | 24 | 2 | H | 349.0694 | 350.8604 | 353
I4 | B | 25 | 2 | H | 349.3541 | 351.1802 | 353
I4 | B | 26 | 2 | H | 349.1327 | 351.1451 | 353
I4 | B | 27 | 2 | H | 348.8992 | 350.9637 | 353
I4 | B | 28 | 2 | H | 348.6669 | 350.7348 | 353
I4 | B | 29 | 2 | H | 348.5752 | 350.9148 | 353
I4 | B | 30 | 2 | H | 349.0518 | 351.6404 | 353
I4 | B | 31 | 2 | H | 349.1683 | 351.713 | 353
I4 | B | 32 | 2 | H | 349.0198 | 351.6793 | 353
I5 | A | 1 | 1 | H | 366.2724 | 369.4031 | 340
I5 | A | 2 | 1 | H | 372.813 | 376.3768 | 340
I5 | A | 3 | 1 | H | 365.3924 | 369.3029 | 340
I5 | A | 4 | 1 | H | 368.9824 | 372.9603 | 340
I5 | A | 5 | 1 | H | 363.3377 | 367.2003 | 340
I5 | A | 6 | 1 | H | 367.6314 | 371.1557 | 340
I5 | A | 7 | 1 | H | 362.7147 | 366.2076 | 340
I5 | A | 8 | 1 | H | 368.4948 | 372.0963 | 340
I5 | A | 9 | 1 | H | 363.9953 | 367.0725 | 340
I5 | A | 10 | 1 | H | 369.4388 | 373.2391 | 340
I5 | A | 11 | 1 | H | 361.1718 | 364.7621 | 340
I5 | A | 12 | 1 | H | 369.8813 | 373.6034 | 340
I5 | A | 13 | 1 | H | 363.8419 | 367.1028 | 340
I5 | A | 14 | 1 | H | 373.9378 | 376.4874 | 340
I5 | A | 15 | 1 | H | 363.9269 | 368.2327 | 340
I5 | A | 16 | 1 | H | 372.0536 | 376.176 | 340
I5 | A | 17 | 1 | H | 369.1281 | 372.5852 | 340
I5 | A | 18 | 1 | H | 372.3574 | 376.5396 | 340
I5 | A | 19 | 1 | H | 368.8437 | 373.3672 | 340
I5 | A | 20 | 1 | H | 373.8642 | 378.5169 | 340
I5 | A | 21 | 1 | H | 367.9452 | 371.9249 | 340
I5 | A | 22 | 1 | H | 378.194 | 382.2377 | 340
I5 | A | 23 | 1 | H | 369.7173 | 374.0395 | 340
I5 | A | 24 | 1 | H | 375.3846 | 381.0823 | 340
I5 | A | 25 | 1 | H | 371.9305 | 376.8925 | 340
I5 | A | 26 | 1 | H | 376.7308 | 380.7897 | 340
I5 | A | 27 | 1 | H | 372.2738 | 376.5715 | 340
I5 | A | 28 | 1 | H | 377.4214 | 382.4379 | 340
I5 | A | 29 | 1 | H | 373.7696 | 377.8293 | 340
I5 | A | 30 | 1 | H | 378.7912 | 383.3446 | 340
I5 | A | 31 | 1 | H | 375.349 | 379.5863 | 340
I5 | A | 32 | 1 | H | 379.6277 | 383.2379 | 340
I5 | A | 1 | 2 | H | 366.2145 | 369.4569 | 340
I5 | A | 2 | 2 | H | 372.5274 | 376.067 | 340
I5 | A | 3 | 2 | H | 365.2348 | 369.0761 | 340
I5 | A | 4 | 2 | H | 368.6194 | 372.526 | 340
I5 | A | 5 | 2 | H | 363.1533 | 367.1839 | 340
I5 | A | 6 | 2 | H | 367.2846 | 370.9643 | 340
I5 | A | 7 | 2 | H | 362.6317 | 366.1998 | 340
I5 | A | 8 | 2 | H | 368.203 | 371.9089 | 340
I5 | A | 9 | 2 | H | 363.9411 | 367.0266 | 340
I5 | A | 10 | 2 | H | 369.2122 | 372.9067 | 340
I5 | A | 11 | 2 | H | 361.1945 | 364.618 | 340
I5 | A | 12 | 2 | H | 369.5158 | 373.2476 | 340
I5 | A | 13 | 2 | H | 363.8172 | 366.959 | 340
I5 | A | 14 | 2 | H | 373.5801 | 376.1462 | 340
I5 | A | 15 | 2 | H | 363.9354 | 368.1777 | 340
I5 | A | 16 | 2 | H | 371.8832 | 375.8381 | 340
I5 | A | 17 | 2 | H | 368.9973 | 372.5003 | 340
I5 | A | 18 | 2 | H | 371.9772 | 376.2687 | 340
I5 | A | 19 | 2 | H | 368.5738 | 373.0912 | 340
I5 | A | 20 | 2 | H | 373.3639 | 378.1779 | 340
I5 | A | 21 | 2 | H | 367.9947 | 371.7112 | 340
I5 | A | 22 | 2 | H | 377.6881 | 382.0277 | 340
I5 | A | 23 | 2 | H | 369.5122 | 373.9925 | 340
I5 | A | 24 | 2 | H | 375.0154 | 380.6621 | 340
I5 | A | 25 | 2 | H | 371.9427 | 376.8832 | 340
I5 | A | 26 | 2 | H | 376.248 | 380.3744 | 340
I5 | A | 27 | 2 | H | 372.1433 | 376.5732 | 340
I5 | A | 28 | 2 | H | 377.1169 | 382.1916 | 340
I5 | A | 29 | 2 | H | 373.628 | 377.9153 | 340
I5 | A | 30 | 2 | H | 378.4694 | 382.8288 | 340
I5 | A | 31 | 2 | H | 375.0599 | 379.5928 | 340
I5 | A | 32 | 2 | H | 378.95 | 382.9423 | 340
I5 | B | 1 | 1 | H | 366.0244 | 369.5885 | 340
I5 | B | 2 | 1 | H | 373.3265 | 376.9153 | 340
I5 | B | 3 | 1 | H | 365.7469 | 369.5255 | 340
I5 | B | 4 | 1 | H | 369.45 | 373.311 | 340
I5 | B | 5 | 1 | H | 363.1763 | 367.3044 | 340
I5 | B | 6 | 1 | H | 367.9174 | 371.5153 | 340
I5 | B | 7 | 1 | H | 363.1752 | 366.4189 | 340
I5 | B | 8 | 1 | H | 368.7087 | 371.9801 | 340
I5 | B | 9 | 1 | H | 364.2835 | 367.7636 | 340
I5 | B | 10 | 1 | H | 369.5431 | 373.6511 | 340
I5 | B | 11 | 1 | H | 361.3635 | 365.0307 | 340
I5 | B | 12 | 1 | H | 370.2327 | 374.0581 | 340
I5 | B | 13 | 1 | H | 363.6065 | 366.9876 | 340
I5 | B | 14 | 1 | H | 373.93 | 377.0814 | 340
I5 | B | 15 | 1 | H | 364.0449 | 368.1677 | 340
I5 | B | 16 | 1 | H | 372.3307 | 376.7286 | 340
I5 | B | 17 | 1 | H | 369.2399 | 374.096 | 340
I5 | B | 18 | 1 | H | 372.7799 | 376.9047 | 340
I5 | B | 19 | 1 | H | 369.0607 | 373.529 | 340
I5 | B | 20 | 1 | H | 374.2125 | 378.7539 | 340
I5 | B | 21 | 1 | H | 367.8836 | 372.8257 | 340
I5 | B | 22 | 1 | H | 377.8183 | 383.1301 | 340
I5 | B | 23 | 1 | H | 369.7706 | 374.2731 | 340
I5 | B | 24 | 1 | H | 375.4952 | 381.061 | 340
I5 | B | 25 | 1 | H | 371.87 | 376.7104 | 340
I5 | B | 26 | 1 | H | 376.8328 | 381.1824 | 340
I5 | B | 27 | 1 | H | 372.0061 | 376.7833 | 340
I5 | B | 28 | 1 | H | 377.9651 | 382.5532 | 340
I5 | B | 29 | 1 | H | 373.7067 | 378.6944 | 340
I5 | B | 30 | 1 | H | 378.7468 | 382.9758 | 340
I5 | B | 31 | 1 | H | 375.2961 | 379.7617 | 340
I5 | B | 32 | 1 | H | 378.9623 | 383.8038 | 340
I5 | B | 1 | 2 | H | 365.9351 | 369.6917 | 340
I5 | B | 2 | 2 | H | 373.0045 | 376.5044 | 340
I5 | B | 3 | 2 | H | 365.9046 | 369.5337 | 340
I5 | B | 4 | 2 | H | 369.0551 | 372.9154 | 340
I5 | B | 5 | 2 | H | 363.1225 | 367.0316 | 340
I5 | B | 6 | 2 | H | 367.5905 | 371.1633 | 340
I5 | B | 7 | 2 | H | 362.9502 | 366.3479 | 340
I5 | B | 8 | 2 | H | 368.2364 | 371.5812 | 340
I5 | B | 9 | 2 | H | 364.1478 | 367.5247 | 340
I5 | B | 10 | 2 | H | 369.3737 | 373.2115 | 340
I5 | B | 11 | 2 | H | 361.2595 | 365.0481 | 340
I5 | B | 12 | 2 | H | 369.7565 | 373.7879 | 340
I5 | B | 13 | 2 | H | 363.4469 | 366.9417 | 340
I5 | B | 14 | 2 | H | 373.6255 | 376.9104 | 340
I5 | B | 15 | 2 | H | 364.2634 | 368.2079 | 340
I5 | B | 16 | 2 | H | 372.0009 | 376.418 | 340
I5 | B | 17 | 2 | H | 369.0539 | 374.052 | 340
I5 | B | 18 | 2 | H | 372.5113 | 376.4328 | 340
I5 | B | 19 | 2 | H | 368.9682 | 373.5522 | 340
I5 | B | 20 | 2 | H | 373.9542 | 378.4178 | 340
I5 | B | 21 | 2 | H | 368.0343 | 372.7445 | 340
I5 | B | 22 | 2 | H | 377.2477 | 382.6412 | 340
I5 | B | 23 | 2 | H | 369.7619 | 374.3586 | 340
I5 | B | 24 | 2 | H | 375.1593 | 380.7112 | 340
I5 | B | 25 | 2 | H | 371.8828 | 376.5116 | 340
I5 | B | 26 | 2 | H | 376.3558 | 380.8203 | 340
I5 | B | 27 | 2 | H | 371.8662 | 376.803 | 340
I5 | B | 28 | 2 | H | 377.5042 | 382.1966 | 340
I5 | B | 29 | 2 | H | 373.6125 | 378.4842 | 340
I5 | B | 30 | 2 | H | 378.6732 | 382.8064 | 340
I5 | B | 31 | 2 | H | 375.1623 | 379.8938 | 340
I5 | B | 32 | 2 | H | 378.6638 | 383.4025 | 340
M12 | A | 1 | 1 | H | 357.186 | 357.367 | 353
M12 | A | 2 | 1 | H | 356.7018 | 356.9688 | 353
M12 | A | 3 | 1 | H | 356.587 | 356.8101 | 353
M12 | A | 4 | 1 | H | 356.9062 | 357.1524 | 353
M12 | A | 5 | 1 | H | 356.8959 | 357.1424 | 353
M12 | A | 6 | 1 | H | 357.3869 | 357.6228 | 353
M12 | A | 7 | 1 | H | 357.1383 | 357.3603 | 353
M12 | A | 8 | 1 | H | 357.3757 | 357.6102 | 353
M12 | A | 9 | 1 | H | 356.9654 | 357.2242 | 353
M12 | A | 10 | 1 | H | 357.3053 | 357.5596 | 353
M12 | A | 11 | 1 | H | 357.0714 | 357.3307 | 353
M12 | A | 12 | 1 | H | 357.3962 | 357.6566 | 353
M12 | A | 13 | 1 | H | 356.9724 | 357.2654 | 353
M12 | A | 14 | 1 | H | 357.2359 | 357.5202 | 353
M12 | A | 15 | 1 | H | 356.8794 | 357.1692 | 353
M12 | A | 16 | 1 | H | 357.5988 | 357.8909 | 353
M12 | B | 1 | 1 | H | 357.1933 | 357.3895 | 353
M12 | B | 2 | 1 | H | 356.6971 | 356.9875 | 353
M12 | B | 3 | 1 | H | 356.5864 | 356.8401 | 353
M12 | B | 4 | 1 | H | 356.9125 | 357.1861 | 353
M12 | B | 5 | 1 | H | 356.8914 | 357.1633 | 353
M12 | B | 6 | 1 | H | 357.3888 | 357.6488 | 353
M12 | B | 7 | 1 | H | 357.1459 | 357.3857 | 353
M12 | B | 8 | 1 | H | 357.3745 | 357.6327 | 353
M12 | B | 9 | 1 | H | 356.9668 | 357.2537 | 353
M12 | B | 10 | 1 | H | 357.305 | 357.5835 | 353
M12 | B | 11 | 1 | H | 357.0648 | 357.3553 | 353
M12 | B | 12 | 1 | H | 357.3971 | 357.6756 | 353
M12 | B | 13 | 1 | H | 356.9738 | 357.2775 | 353
M12 | B | 14 | 1 | H | 357.2332 | 357.5552 | 353
M12 | B | 15 | 1 | H | 356.8836 | 357.1912 | 353
M12 | B | 16 | 1 | H | 357.5896 | 357.9079 | 353
M13 | A | 1 | 1 | H | 349.4265 | 348.1597 | 343
M13 | A | 2 | 1 | H | 351.3087 | 350.0411 | 343
M13 | A | 3 | 1 | H | 349.0044 | 347.796 | 343
M13 | A | 4 | 1 | H | 351.4363 | 350.1588 | 343
M13 | A | 5 | 1 | H | 349.3656 | 348.1698 | 343
M13 | A | 6 | 1 | H | 352.0453 | 350.7425 | 343
M13 | A | 7 | 1 | H | 349.7628 | 348.526 | 343
M13 | A | 8 | 1 | H | 352.2973 | 350.954 | 343
M13 | A | 9 | 1 | H | 349.8583 | 348.5905 | 343
M13 | A | 10 | 1 | H | 351.9951 | 350.6443 | 343
M13 | A | 11 | 1 | H | 350.0837 | 348.8119 | 343
M13 | A | 12 | 1 | H | 352.2491 | 350.8981 | 343
M13 | A | 13 | 1 | H | 349.8425 | 348.5216 | 343
M13 | A | 14 | 1 | H | 352.1676 | 350.7721 | 343
M13 | A | 15 | 1 | H | 349.9403 | 348.6102 | 343
M13 | A | 16 | 1 | H | 352.8327 | 351.4545 | 343
M13 | B | 1 | 1 | H | 349.2923 | 348.1681 | 343
M13 | B | 2 | 1 | H | 351.1886 | 350.0563 | 343
M13 | B | 3 | 1 | H | 348.8738 | 347.8114 | 343
M13 | B | 4 | 1 | H | 351.307 | 350.1751 | 343
M13 | B | 5 | 1 | H | 349.2381 | 348.1778 | 343
M13 | B | 6 | 1 | H | 351.9309 | 350.7645 | 343
M13 | B | 7 | 1 | H | 349.6195 | 348.5359 | 343
M13 | B | 8 | 1 | H | 352.1811 | 350.9706 | 343
M13 | B | 9 | 1 | H | 349.7477 | 348.6042 | 343
M13 | B | 10 | 1 | H | 351.8894 | 350.6623 | 343
M13 | B | 11 | 1 | H | 349.9877 | 348.8222 | 343
M13 | B | 12 | 1 | H | 352.1597 | 350.9124 | 343
M13 | B | 13 | 1 | H | 349.7448 | 348.5392 | 343
M13 | B | 14 | 1 | H | 352.0905 | 350.774 | 343
M13 | B | 15 | 1 | H | 349.8807 | 348.6247 | 343
M13 | B | 16 | 1 | H | 352.806 | 351.4713 | 343
M14 | A | 1 | 1 | H | 349.1364 | 352.3971 | 336
M14 | A | 2 | 1 | H | 346.5348 | 349.3427 | 336
M14 | A | 3 | 1 | H | 347.9463 | 351.3212 | 336
M14 | A | 4 | 1 | H | 350.9639 | 354.1773 | 336
M14 | A | 5 | 1 | H | 347.9076 | 351.0989 | 336
M14 | A | 6 | 1 | H | 347.9651 | 351.4018 | 336
M14 | A | 7 | 1 | H | 349.5808 | 353.3296 | 336
M14 | A | 8 | 1 | H | 348.2805 | 351.4892 | 336
M14 | A | 9 | 1 | H | 349.2831 | 352.8494 | 336
M14 | A | 10 | 1 | H | 347.9321 | 351.1883 | 336
M14 | A | 11 | 1 | H | 351.1924 | 355.3192 | 336
M14 | A | 12 | 1 | H | 349.2866 | 352.901 | 336
M14 | A | 13 | 1 | H | 348.2224 | 351.6193 | 336
M14 | A | 14 | 1 | H | 347.7115 | 351.0274 | 336
M14 | A | 15 | 1 | H | 349.4686 | 353.0315 | 336
M14 | A | 16 | 1 | H | 351.0201 | 355.0957 | 336
M14 | B | 1 | 1 | H | 349.1306 | 352.4068 | 336
M14 | B | 2 | 1 | H | 346.5165 | 349.3422 | 336
M14 | B | 3 | 1 | H | 347.9417 | 351.3046 | 336
M14 | B | 4 | 1 | H | 350.8935 | 354.1539 | 336
M14 | B | 5 | 1 | H | 347.8835 | 351.0956 | 336
M14 | B | 6 | 1 | H | 347.944 | 351.4412 | 336
M14 | B | 7 | 1 | H | 349.5531 | 353.342 | 336
M14 | B | 8 | 1 | H | 348.2654 | 351.4946 | 336
M14 | B | 9 | 1 | H | 349.2679 | 352.8355 | 336
M14 | B | 10 | 1 | H | 347.9107 | 351.1333 | 336
M14 | B | 11 | 1 | H | 351.1783 | 355.3412 | 336
M14 | B | 12 | 1 | H | 349.2444 | 352.8915 | 336
M14 | B | 13 | 1 | H | 348.1924 | 351.5812 | 336
M14 | B | 14 | 1 | H | 347.756 | 351.009 | 336
M14 | B | 15 | 1 | H | 349.4313 | 353.0428 | 336
M14 | B | 16 | 1 | H | 350.9664 | 355.0298 | 336
M15 | A | 1 | 1 | H | 341.6069 | 345.1809 | 343
M15 | A | 2 | 1 | H | 358.94 | 360.7418 | 343
M15 | A | 3 | 1 | H | 357.273 | 359.0653 | 343
M15 | A | 4 | 1 | H | 359.4279 | 361.0596 | 343
M15 | A | 5 | 1 | H | 357.5892 | 359.5128 | 343
M15 | A | 6 | 1 | H | 357.849 | 359.7395 | 343
M15 | A | 7 | 1 | H | 357.5768 | 359.5884 | 343
M15 | A | 8 | 1 | H | 358.9956 | 360.9999 | 343
M15 | A | 9 | 1 | H | 357.8655 | 360.1925 | 343
M15 | A | 10 | 1 | H | 359.1688 | 361.2536 | 343
M15 | A | 11 | 1 | H | 358.0656 | 359.8345 | 343
M15 | A | 12 | 1 | H | 359.3084 | 361.6215 | 343
M15 | A | 13 | 1 | H | 357.7318 | 359.8319 | 343
M15 | A | 14 | 1 | H | 359.0869 | 361.252 | 343
M15 | A | 15 | 1 | H | 358.0692 | 360.6351 | 343
M15 | A | 16 | 1 | H | 361.3381 | 363.6438 | 343
M15 | B | 1 | 1 | H | 341.8394 | 345.488 | 343
M15 | B | 2 | 1 | H | 359.1137 | 361.084 | 343
M15 | B | 3 | 1 | H | 357.5039 | 359.3767 | 343
M15 | B | 4 | 1 | H | 359.6262 | 361.3251 | 343
M15 | B | 5 | 1 | H | 357.805 | 359.7669 | 343
M15 | B | 6 | 1 | H | 358.0931 | 359.9814 | 343
M15 | B | 7 | 1 | H | 357.7868 | 359.8427 | 343
M15 | B | 8 | 1 | H | 359.1855 | 361.2888 | 343
M15 | B | 9 | 1 | H | 358.2427 | 360.4833 | 343
M15 | B | 10 | 1 | H | 359.4071 | 361.5576 | 343
M15 | B | 11 | 1 | H | 358.2452 | 360.0741 | 343
M15 | B | 12 | 1 | H | 359.4674 | 361.8479 | 343
M15 | B | 13 | 1 | H | 357.9087 | 360.1096 | 343
M15 | B | 14 | 1 | H | 359.2847 | 361.4237 | 343
M15 | B | 15 | 1 | H | 358.3048 | 360.9543 | 343
M15 | B | 16 | 1 | H | 361.5143 | 363.9022 | 343
M16A | A | 1 | 1 | H | 368.9055 | 371.8267 | 340
M16A | A | 2 | 1 | H | 366.1279 | 370.0279 | 340
M16A | A | 3 | 1 | H | 365.1637 | 368.0776 | 340
M16A | A | 4 | 1 | H | 366.3754 | 370.0247 | 340
M16A | A | 5 | 1 | H | 366.8446 | 370.1453 | 340
M16A | A | 6 | 1 | H | 366.9845 | 370.9983 | 340
M16A | A | 7 | 1 | H | 367.3571 | 370.9121 | 340
M16A | A | 8 | 1 | H | 367.8787 | 371.3349 | 340
M16A | A | 9 | 1 | H | 367.9595 | 371.6389 | 340
M16A | A | 10 | 1 | H | 369.1944 | 373.3124 | 340
M16A | A | 11 | 1 | H | 368.4952 | 372.3429 | 340
M16A | A | 12 | 1 | H | 370.5298 | 375.0103 | 340
M16A | A | 13 | 1 | H | 369.9579 | 374.241 | 340
M16A | A | 14 | 1 | H | 368.534 | 372.8549 | 340
M16A | A | 15 | 1 | H | 371.9113 | 375.8642 | 340
M16A | A | 16 | 1 | H | 371.8369 | 375.8708 | 340
M16A | B | 1 | 1 | H | 369.1423 | 372.1651 | 340
M16A | B | 2 | 1 | H | 366.2672 | 370.3171 | 340
M16A | B | 3 | 1 | H | 365.4478 | 368.5644 | 340
M16A | B | 4 | 1 | H | 366.5744 | 370.3396 | 340
M16A | B | 5 | 1 | H | 367.0795 | 370.5349 | 340
M16A | B | 6 | 1 | H | 366.9676 | 371.1969 | 340
M16A | B | 7 | 1 | H | 367.4086 | 371.2077 | 340
M16A | B | 8 | 1 | H | 368.1165 | 371.7122 | 340
M16A | B | 9 | 1 | H | 368.1349 | 372.01 | 340
M16A | B | 10 | 1 | H | 369.4199 | 373.6715 | 340
M16A | B | 11 | 1 | H | 368.8442 | 372.8377 | 340
M16A | B | 12 | 1 | H | 370.7755 | 375.2564 | 340
M16A | B | 13 | 1 | H | 370.1351 | 374.3953 | 340
M16A | B | 14 | 1 | H | 368.8644 | 373.2132 | 340
M16A | B | 15 | 1 | H | 372.1397 | 376.2634 | 340
M16A | B | 16 | 1 | H | 372.149 | 376.2101 | 340
M16B | A | 1 | 1 | H | 367.2041 | 370.6598 | 340
M16B | A | 2 | 1 | H | 367.3775 | 371.174 | 340
M16B | A | 3 | 1 | H | 369.7977 | 373.9955 | 340
M16B | A | 4 | 1 | H | 366.5499 | 370.7345 | 340
M16B | A | 5 | 1 | H | 368.3757 | 372.4347 | 340
M16B | A | 6 | 1 | H | 368.5405 | 372.6084 | 340
M16B | A | 7 | 1 | H | 368.926 | 372.7885 | 340
M16B | A | 8 | 1 | H | 368.5407 | 373.1368 | 340
M16B | A | 9 | 1 | H | 370.4675 | 374.4058 | 340
M16B | A | 10 | 1 | H | 368.0431 | 372.1876 | 340
M16B | A | 11 | 1 | H | 370.3438 | 374.496 | 340
M16B | A | 12 | 1 | H | 370.1849 | 374.8357 | 340
M16B | A | 13 | 1 | H | 369.5717 | 374.4232 | 340
M16B | A | 14 | 1 | H | 370.2476 | 375.1837 | 340
M16B | A | 15 | 1 | H | 372.4701 | 377.329 | 340
M16B | A | 16 | 1 | H | 372.0213 | 375.8885 | 340
M16B | B | 1 | 1 | H | 367.4211 | 370.9667 | 340
M16B | B | 2 | 1 | H | 367.39 | 371.5641 | 340
M16B | B | 3 | 1 | H | 370.0171 | 374.4221 | 340
M16B | B | 4 | 1 | H | 366.7429 | 370.9532 | 340
M16B | B | 5 | 1 | H | 368.5046 | 372.7782 | 340
M16B | B | 6 | 1 | H | 368.7315 | 372.9112 | 340
M16B | B | 7 | 1 | H | 368.9826 | 373.0955 | 340
M16B | B | 8 | 1 | H | 368.7398 | 373.2502 | 340
M16B | B | 9 | 1 | H | 370.5145 | 374.6653 | 340
M16B | B | 10 | 1 | H | 368.022 | 372.4155 | 340
M16B | B | 11 | 1 | H | 370.3267 | 374.7613 | 340
M16B | B | 12 | 1 | H | 370.0571 | 374.9715 | 340
M16B | B | 13 | 1 | H | 370.0111 | 374.6205 | 340
M16B | B | 14 | 1 | H | 370.6033 | 375.283 | 340
M16B | B | 15 | 1 | H | 372.4975 | 377.5491 | 340
M16B | B | 16 | 1 | H | 372.3757 | 376.1707 | 340
M13 | A | 1 | 1 | L | 654.2278 | 655.688 | 634
M13 | A | 2 | 1 | L | 658.668 | 659.7715 | 634
M13 | A | 3 | 1 | L | 653.8524 | 654.968 | 634
M13 | A | 4 | 1 | L | 659.0044 | 660.1354 | 634
M13 | A | 5 | 1 | L | 654.1479 | 655.2943 | 634
M13 | A | 6 | 1 | L | 660.4117 | 661.5061 | 634
M13 | A | 7 | 1 | L | 654.967 | 656.0302 | 634
M13 | A | 8 | 1 | L | 660.1322 | 661.1763 | 634
M13 | A | 9 | 1 | L | 655.4036 | 656.457 | 634
M13 | A | 10 | 1 | L | 659.976 | 661.0113 | 634
M13 | A | 11 | 1 | L | 655.6609 | 656.7045 | 634
M13 | A | 12 | 1 | L | 660.77 | 661.8509 | 634
M13 | A | 13 | 1 | L | 655.5522 | 656.6009 | 634
M13 | A | 14 | 1 | L | 661.1069 | 662.1833 | 634
M13 | A | 15 | 1 | L | 656.2874 | 657.3056 | 634
M13 | A | 16 | 1 | L | 662.1858 | 663.8474 | 634
M13 | B | 1 | 1 | L | 654.2425 | 655.7421 | 634
M13 | B | 2 | 1 | L | 658.6881 | 659.8235 | 634
M13 | B | 3 | 1 | L | 653.8628 | 655.02 | 634
M13 | B | 4 | 1 | L | 659.027 | 660.186 | 634
M13 | B | 5 | 1 | L | 654.1614 | 655.3377 | 634
M13 | B | 6 | 1 | L | 660.4366 | 661.5532 | 634
M13 | B | 7 | 1 | L | 654.9803 | 656.0743 | 634
M13 | B | 8 | 1 | L | 660.1526 | 661.2219 | 634
M13 | B | 9 | 1 | L | 655.4173 | 656.5055 | 634
M13 | B | 10 | 1 | L | 659.9914 | 661.0616 | 634
M13 | B | 11 | 1 | L | 655.6722 | 656.75 | 634
M13 | B | 12 | 1 | L | 660.7898 | 661.9024 | 634
M13 | B | 13 | 1 | L | 655.559 | 656.6429 | 634
M13 | B | 14 | 1 | L | 661.1258 | 662.2296 | 634
M13 | B | 15 | 1 | L | 656.2974 | 657.3497 | 634
M13 | B | 16 | 1 | L | 662.201 | 663.8948 | 634

## RSB saturation
Band | HAM | Subsample | Gain State | Band Median |  | Band Average |  | Lmax
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | Lsat | Lsat | Lsat | Lsat
I1 | A | 1 | H | 840.405 | 856.015 | 840.6481 | 856.1534 | 718
I1 | A | 2 | H | 840.51 | 856.16 | 840.6713 | 856.1934 | 718
I1 | B | 1 | H | 837.075 | 852.75 | 837.2731 | 852.895 | 718
I1 | B | 2 | H | 837.055 | 852.905 | 837.2597 | 852.9681 | 718
I2 | A | 1 | H | 413.325 | 422.38 | 411.9581 | 420.8622 | 349
I2 | A | 2 | H | 413.5 | 422.265 | 412.0381 | 420.9103 | 349
I2 | B | 1 | H | 413.755 | 422.665 | 412.3006 | 421.1266 | 349
I2 | B | 2 | H | 413.845 | 422.5 | 412.3941 | 421.1578 | 349
I3 | A | 1 | H | 68.64 | 70.835 | 69.4363 | 72.6075 | 72.5
I3 | B | 2 | H | 68.62 | 70.825 | 69.4119 | 72.635 | 72.5
I3 | A | 1 | H | 68.75 | 70.965 | 69.5362 | 72.8037 | 72.5
I3 | B | 2 | H | 68.745 | 70.94 | 69.5222 | 72.7469 | 72.5
M1 | A | 1 | H | 184.385 | 193.34 | 185.2388 | 194.2763 | 135
M1 | B | 1 | H | 188.4 | 197.62 | 189.2712 | 198.5894 | 135
M2 | A | 1 | H | 178.59 | 184.4 | 179.0756 | 184.5587 | 127
M2 | B | 1 | H | 180.27 | 186.075 | 180.7794 | 186.3306 | 127
M3 | A | 1 | H | 135.795 | 139.86 | 135.8244 | 139.9281 | 107
M3 | B | 1 | H | 136.215 | 140.325 | 136.2669 | 140.3988 | 107
M4 | A | 1 | H | 110.9 | 119.64 | 111.2313 | 119.9487 | 78
M4 | B | 1 | H | 110.76 | 119.525 | 111.12 | 119.8381 | 78
M5 | A | 1 | H | 72.005 | 74.8 | 72.0062 | 74.7969 | 59
M5 | B | 1 | H | 71.675 | 74.445 | 71.68 | 74.4663 | 59
M6 | A | 1 | H | 47.48 | 49.25 | 47.4613 | 49.2512 | 41
M6 | B | 1 | H | 47.26 | 49.065 | 47.2587 | 49.0538 | 41
M7 | A | 1 | H | 36.98 | 38.15 | 36.9937 | 38.1569 | 29
M7 | B | 1 | H | 37.01 | 38.175 | 37.0262 | 38.1788 | 29
M8 | A | 1 | H | 121.495 | 124.255 | 121.2194 | 124.125 | 164.9
M8 | B | 1 | H | 121.92 | 124.7 | 121.6419 | 124.5606 | 164.9
M9 | A | 1 | H | 92.85 | 93.455 | 92.7381 | 93.2913 | 77.1
M9 | B | 1 | H | 93.095 | 93.735 | 92.9931 | 93.5494 | 77.1
M10 | A | 1 | H | 82.365 | 83.4 | 81.9469 | 83.0531 | 71.2
M10 | B | 1 | H | 82.5 | 83.535 | 82.0737 | 83.1919 | 71.2
M11 | A | 1 | H | 33.49 | 33.165 | 33.5494 | 33.2644 | 31.8
M11 | B | 1 | H | 33.51 | 33.215 | 33.5688 | 33.2875 | 31.8
M1 | A | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 615
M1 | B | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 615
M2 | A | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 687
M2 | B | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 687
M3 | A | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 702
M3 | B | 1 | L | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 702
M4 | A | 1 | L | 864.23 | 884.17 | 866.0025 | 886.3825 | 667
M4 | B | 1 | L | 863.34 | 883.515 | 865.16 | 885.7319 | 667
M5 | A | 1 | L | 723.81 | 740.085 | 723.6944 | 740.0875 | 651
M5 | B | 1 | L | 719.365 | 736.735 | 720.0119 | 736.8044 | 651
M7 | A | 1 | L | 416.45 | 426.225 | 416.3706 | 426.1862 | 349
M7 | B | 1 | L | 416.8 | 426.475 | 416.7288 | 426.4475 | 349
Band | HAM | Detector | Subsample | Gain State | Detector Dependent |  | Lmax |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 RSB saturation.
[5 empty] | Lsat | Lsat |  |  | 2. The results were derived from RC-1 using the SIS-100 source (ScMA for M1, M2, and M3 low gain).
I1 | A | 1 | 1 | H | 825.37 | 840.77 | 718 |  | 3. Saturation was determined from partial views of the SIS-100 source (data not available for M1, M2, or M3).
I1 | A | 2 | 1 | H | 827.33 | 842.45 | 718 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | A | 3 | 1 | H | 828.75 | 844.52 | 718 |  | 5. The test was set to operational day mode.
I1 | A | 4 | 1 | H | 832.06 | 847.67 | 718
I1 | A | 5 | 1 | H | 831.38 | 847.45 | 718
I1 | A | 6 | 1 | H | 830.98 | 846.17 | 718
I1 | A | 7 | 1 | H | 834.49 | 849.8 | 718
I1 | A | 8 | 1 | H | 830.91 | 846.16 | 718
I1 | A | 9 | 1 | H | 828.53 | 844.22 | 718
I1 | A | 10 | 1 | H | 829.19 | 844.45 | 718
I1 | A | 11 | 1 | H | 844.18 | 859.95 | 718
I1 | A | 12 | 1 | H | 843.69 | 859.48 | 718
I1 | A | 13 | 1 | H | 840 | 855.66 | 718
I1 | A | 14 | 1 | H | 841.35 | 856.55 | 718
I1 | A | 15 | 1 | H | 847.76 | 863.05 | 718
I1 | A | 16 | 1 | H | 853.51 | 868.92 | 718
I1 | A | 17 | 1 | H | 845.84 | 861.55 | 718
I1 | A | 18 | 1 | H | 842.03 | 857.02 | 718
I1 | A | 19 | 1 | H | 836.48 | 851.79 | 718
I1 | A | 20 | 1 | H | 836.13 | 851.29 | 718
I1 | A | 21 | 1 | H | 838.37 | 853.66 | 718
I1 | A | 22 | 1 | H | 839.38 | 854.93 | 718
I1 | A | 23 | 1 | H | 844.92 | 860.2 | 718
I1 | A | 24 | 1 | H | 855.31 | 870.81 | 718
I1 | A | 25 | 1 | H | 859.36 | 875.09 | 718
I1 | A | 26 | 1 | H | 856.11 | 871.75 | 718
I1 | A | 27 | 1 | H | 854.29 | 870.07 | 718
I1 | A | 28 | 1 | H | 840.81 | 856.37 | 718
I1 | A | 29 | 1 | H | 839.76 | 855.64 | 718
I1 | A | 30 | 1 | H | 843.27 | 859.05 | 718
I1 | A | 31 | 1 | H | 855.75 | 871.21 | 718
I1 | A | 32 | 1 | H | 843.45 | 859.21 | 718
I1 | A | 1 | 2 | H | 825.13 | 840.92 | 718
I1 | A | 2 | 2 | H | 827.24 | 842.48 | 718
I1 | A | 3 | 2 | H | 828.67 | 844.6 | 718
I1 | A | 4 | 2 | H | 832.08 | 847.75 | 718
I1 | A | 5 | 2 | H | 831.37 | 847.34 | 718
I1 | A | 6 | 2 | H | 830.96 | 846.29 | 718
I1 | A | 7 | 2 | H | 834.5 | 849.85 | 718
I1 | A | 8 | 2 | H | 830.95 | 846.17 | 718
I1 | A | 9 | 2 | H | 828.33 | 844.31 | 718
I1 | A | 10 | 2 | H | 829.06 | 844.48 | 718
I1 | A | 11 | 2 | H | 844.29 | 859.8 | 718
I1 | A | 12 | 2 | H | 843.78 | 859.42 | 718
I1 | A | 13 | 2 | H | 840.1 | 855.72 | 718
I1 | A | 14 | 2 | H | 841.41 | 856.62 | 718
I1 | A | 15 | 2 | H | 847.79 | 862.97 | 718
I1 | A | 16 | 2 | H | 853.54 | 868.75 | 718
I1 | A | 17 | 2 | H | 846.03 | 861.39 | 718
I1 | A | 18 | 2 | H | 842.17 | 857.24 | 718
I1 | A | 19 | 2 | H | 836.44 | 851.87 | 718
I1 | A | 20 | 2 | H | 836.28 | 851.46 | 718
I1 | A | 21 | 2 | H | 838.45 | 853.88 | 718
I1 | A | 22 | 2 | H | 839.65 | 855.17 | 718
I1 | A | 23 | 2 | H | 845.07 | 860.42 | 718
I1 | A | 24 | 2 | H | 855.38 | 870.64 | 718
I1 | A | 25 | 2 | H | 859.61 | 875.22 | 718
I1 | A | 26 | 2 | H | 855.99 | 871.73 | 718
I1 | A | 27 | 2 | H | 854.05 | 869.94 | 718
I1 | A | 28 | 2 | H | 840.92 | 856.43 | 718
I1 | A | 29 | 2 | H | 839.91 | 855.89 | 718
I1 | A | 30 | 2 | H | 843.5 | 858.97 | 718
I1 | A | 31 | 2 | H | 855.45 | 871.25 | 718
I1 | A | 32 | 2 | H | 843.38 | 859.22 | 718
I1 | B | 1 | 1 | H | 822.18 | 837.58 | 718
I1 | B | 2 | 1 | H | 824.14 | 839.37 | 718
I1 | B | 3 | 1 | H | 825.41 | 841.45 | 718
I1 | B | 4 | 1 | H | 828.81 | 844.65 | 718
I1 | B | 5 | 1 | H | 828.1 | 844.26 | 718
I1 | B | 6 | 1 | H | 827.66 | 842.98 | 718
I1 | B | 7 | 1 | H | 831.12 | 846.66 | 718
I1 | B | 8 | 1 | H | 827.43 | 842.96 | 718
I1 | B | 9 | 1 | H | 825.26 | 841.06 | 718
I1 | B | 10 | 1 | H | 825.84 | 841.31 | 718
I1 | B | 11 | 1 | H | 840.99 | 856.62 | 718
I1 | B | 12 | 1 | H | 840.37 | 856.04 | 718
I1 | B | 13 | 1 | H | 836.7 | 852.39 | 718
I1 | B | 14 | 1 | H | 838.02 | 853.21 | 718
I1 | B | 15 | 1 | H | 844.37 | 859.75 | 718
I1 | B | 16 | 1 | H | 850.16 | 865.46 | 718
I1 | B | 17 | 1 | H | 842.56 | 858.07 | 718
I1 | B | 18 | 1 | H | 838.71 | 853.94 | 718
I1 | B | 19 | 1 | H | 833.12 | 848.68 | 718
I1 | B | 20 | 1 | H | 832.74 | 847.98 | 718
I1 | B | 21 | 1 | H | 834.87 | 850.45 | 718
I1 | B | 22 | 1 | H | 836.01 | 851.7 | 718
I1 | B | 23 | 1 | H | 841.53 | 856.86 | 718
I1 | B | 24 | 1 | H | 851.74 | 867.44 | 718
I1 | B | 25 | 1 | H | 855.93 | 871.73 | 718
I1 | B | 26 | 1 | H | 852.64 | 868.42 | 718
I1 | B | 27 | 1 | H | 850.51 | 866.77 | 718
I1 | B | 28 | 1 | H | 837.45 | 853.08 | 718
I1 | B | 29 | 1 | H | 836.43 | 852.42 | 718
I1 | B | 30 | 1 | H | 839.84 | 855.68 | 718
I1 | B | 31 | 1 | H | 852.18 | 867.93 | 718
I1 | B | 32 | 1 | H | 839.92 | 855.74 | 718
I1 | B | 1 | 2 | H | 821.97 | 837.84 | 718
I1 | B | 2 | 2 | H | 823.85 | 839.51 | 718
I1 | B | 3 | 2 | H | 825.13 | 841.61 | 718
I1 | B | 4 | 2 | H | 828.72 | 844.77 | 718
I1 | B | 5 | 2 | H | 828.02 | 844.41 | 718
I1 | B | 6 | 2 | H | 827.55 | 843.21 | 718
I1 | B | 7 | 2 | H | 831.1 | 846.74 | 718
I1 | B | 8 | 2 | H | 827.27 | 843 | 718
I1 | B | 9 | 2 | H | 825.12 | 841.08 | 718
I1 | B | 10 | 2 | H | 825.86 | 841.3 | 718
I1 | B | 11 | 2 | H | 841.07 | 856.6 | 718
I1 | B | 12 | 2 | H | 840.53 | 856 | 718
I1 | B | 13 | 2 | H | 836.68 | 852.66 | 718
I1 | B | 14 | 2 | H | 838 | 853.47 | 718
I1 | B | 15 | 2 | H | 844.3 | 859.7 | 718
I1 | B | 16 | 2 | H | 850.13 | 865.29 | 718
I1 | B | 17 | 2 | H | 842.66 | 858.15 | 718
I1 | B | 18 | 2 | H | 838.71 | 854.09 | 718
I1 | B | 19 | 2 | H | 833.02 | 848.57 | 718
I1 | B | 20 | 2 | H | 832.79 | 848.22 | 718
I1 | B | 21 | 2 | H | 835.11 | 850.59 | 718
I1 | B | 22 | 2 | H | 836.26 | 851.97 | 718
I1 | B | 23 | 2 | H | 841.66 | 857.05 | 718
I1 | B | 24 | 2 | H | 851.73 | 867.4 | 718
I1 | B | 25 | 2 | H | 856.01 | 871.72 | 718
I1 | B | 26 | 2 | H | 852.54 | 868.42 | 718
I1 | B | 27 | 2 | H | 850.56 | 866.63 | 718
I1 | B | 28 | 2 | H | 837.43 | 853.15 | 718
I1 | B | 29 | 2 | H | 836.56 | 852.58 | 718
I1 | B | 30 | 2 | H | 839.96 | 855.54 | 718
I1 | B | 31 | 2 | H | 852.08 | 867.97 | 718
I1 | B | 32 | 2 | H | 839.93 | 855.74 | 718
I2 | A | 1 | 1 | H | 403.06 | 411.75 | 349
I2 | A | 2 | 1 | H | 406.45 | 414.83 | 349
I2 | A | 3 | 1 | H | 408.71 | 417.7 | 349
I2 | A | 4 | 1 | H | 404.84 | 414.01 | 349
I2 | A | 5 | 1 | H | 404.44 | 413.65 | 349
I2 | A | 6 | 1 | H | 404.82 | 413.55 | 349
I2 | A | 7 | 1 | H | 403.93 | 412.6 | 349
I2 | A | 8 | 1 | H | 396.34 | 405.19 | 349
I2 | A | 9 | 1 | H | 396.69 | 405.69 | 349
I2 | A | 10 | 1 | H | 397.15 | 405.97 | 349
I2 | A | 11 | 1 | H | 402.93 | 411.84 | 349
I2 | A | 12 | 1 | H | 404.13 | 413 | 349
I2 | A | 13 | 1 | H | 406.59 | 415.38 | 349
I2 | A | 14 | 1 | H | 409.98 | 419.16 | 349
I2 | A | 15 | 1 | H | 413.65 | 422.69 | 349
I2 | A | 16 | 1 | H | 415.27 | 424.02 | 349
I2 | A | 17 | 1 | H | 416.2 | 425.13 | 349
I2 | A | 18 | 1 | H | 413 | 422.07 | 349
I2 | A | 19 | 1 | H | 414.62 | 423.6 | 349
I2 | A | 20 | 1 | H | 412.59 | 421.51 | 349
I2 | A | 21 | 1 | H | 415.69 | 424.54 | 349
I2 | A | 22 | 1 | H | 414.2 | 423.39 | 349
I2 | A | 23 | 1 | H | 419.22 | 427.87 | 349
I2 | A | 24 | 1 | H | 420.62 | 429.64 | 349
I2 | A | 25 | 1 | H | 426.38 | 435.45 | 349
I2 | A | 26 | 1 | H | 426.03 | 435.03 | 349
I2 | A | 27 | 1 | H | 418.27 | 426.9 | 349
I2 | A | 28 | 1 | H | 416.03 | 425.16 | 349
I2 | A | 29 | 1 | H | 417.77 | 426.79 | 349
I2 | A | 30 | 1 | H | 415.7 | 424.57 | 349
I2 | A | 31 | 1 | H | 429.32 | 438.17 | 349
I2 | A | 32 | 1 | H | 428.04 | 436.74 | 349
I2 | A | 1 | 2 | H | 403.14 | 411.76 | 349
I2 | A | 2 | 2 | H | 406.66 | 414.92 | 349
I2 | A | 3 | 2 | H | 408.83 | 417.68 | 349
I2 | A | 4 | 2 | H | 404.95 | 413.98 | 349
I2 | A | 5 | 2 | H | 404.57 | 413.72 | 349
I2 | A | 6 | 2 | H | 404.82 | 413.6 | 349
I2 | A | 7 | 2 | H | 403.85 | 412.7 | 349
I2 | A | 8 | 2 | H | 396.42 | 405.17 | 349
I2 | A | 9 | 2 | H | 396.78 | 405.76 | 349
I2 | A | 10 | 2 | H | 397.26 | 406.04 | 349
I2 | A | 11 | 2 | H | 402.79 | 411.83 | 349
I2 | A | 12 | 2 | H | 404.08 | 413.07 | 349
I2 | A | 13 | 2 | H | 406.53 | 415.54 | 349
I2 | A | 14 | 2 | H | 410.24 | 419.36 | 349
I2 | A | 15 | 2 | H | 413.77 | 422.45 | 349
I2 | A | 16 | 2 | H | 415.35 | 424.1 | 349
I2 | A | 17 | 2 | H | 416.24 | 425.02 | 349
I2 | A | 18 | 2 | H | 413.23 | 422.08 | 349
I2 | A | 19 | 2 | H | 414.66 | 423.54 | 349
I2 | A | 20 | 2 | H | 412.75 | 421.65 | 349
I2 | A | 21 | 2 | H | 415.9 | 424.7 | 349
I2 | A | 22 | 2 | H | 414.4 | 423.48 | 349
I2 | A | 23 | 2 | H | 419.31 | 427.97 | 349
I2 | A | 24 | 2 | H | 420.88 | 429.82 | 349
I2 | A | 25 | 2 | H | 426.36 | 435.47 | 349
I2 | A | 26 | 2 | H | 426.09 | 435.15 | 349
I2 | A | 27 | 2 | H | 418.27 | 427.07 | 349
I2 | A | 28 | 2 | H | 416.15 | 425.11 | 349
I2 | A | 29 | 2 | H | 417.8 | 426.75 | 349
I2 | A | 30 | 2 | H | 415.87 | 424.58 | 349
I2 | A | 31 | 2 | H | 429.22 | 438.16 | 349
I2 | A | 32 | 2 | H | 428.05 | 436.9 | 349
I2 | B | 1 | 1 | H | 403.25 | 411.95 | 349
I2 | B | 2 | 1 | H | 406.74 | 415.02 | 349
I2 | B | 3 | 1 | H | 409.02 | 417.89 | 349
I2 | B | 4 | 1 | H | 405.06 | 414.18 | 349
I2 | B | 5 | 1 | H | 404.7 | 413.94 | 349
I2 | B | 6 | 1 | H | 405.04 | 413.67 | 349
I2 | B | 7 | 1 | H | 404.19 | 412.82 | 349
I2 | B | 8 | 1 | H | 396.6 | 405.38 | 349
I2 | B | 9 | 1 | H | 396.96 | 405.89 | 349
I2 | B | 10 | 1 | H | 397.55 | 406.22 | 349
I2 | B | 11 | 1 | H | 403.21 | 412.08 | 349
I2 | B | 12 | 1 | H | 404.46 | 413.31 | 349
I2 | B | 13 | 1 | H | 406.87 | 415.66 | 349
I2 | B | 14 | 1 | H | 410.36 | 419.44 | 349
I2 | B | 15 | 1 | H | 414.1 | 422.98 | 349
I2 | B | 16 | 1 | H | 415.59 | 424.35 | 349
I2 | B | 17 | 1 | H | 416.55 | 425.35 | 349
I2 | B | 18 | 1 | H | 413.41 | 422.35 | 349
I2 | B | 19 | 1 | H | 414.96 | 423.87 | 349
I2 | B | 20 | 1 | H | 412.96 | 421.82 | 349
I2 | B | 21 | 1 | H | 416.09 | 424.76 | 349
I2 | B | 22 | 1 | H | 414.62 | 423.62 | 349
I2 | B | 23 | 1 | H | 419.58 | 428.17 | 349
I2 | B | 24 | 1 | H | 421.09 | 429.89 | 349
I2 | B | 25 | 1 | H | 426.77 | 435.67 | 349
I2 | B | 26 | 1 | H | 426.47 | 435.42 | 349
I2 | B | 27 | 1 | H | 418.57 | 427.35 | 349
I2 | B | 28 | 1 | H | 416.41 | 425.46 | 349
I2 | B | 29 | 1 | H | 418.2 | 427.08 | 349
I2 | B | 30 | 1 | H | 416.11 | 424.95 | 349
I2 | B | 31 | 1 | H | 429.71 | 438.48 | 349
I2 | B | 32 | 1 | H | 428.42 | 437.03 | 349
I2 | B | 1 | 2 | H | 403.42 | 411.96 | 349
I2 | B | 2 | 2 | H | 407 | 415.13 | 349
I2 | B | 3 | 2 | H | 409.16 | 417.83 | 349
I2 | B | 4 | 2 | H | 405.28 | 414.2 | 349
I2 | B | 5 | 2 | H | 404.83 | 414 | 349
I2 | B | 6 | 2 | H | 405.16 | 413.82 | 349
I2 | B | 7 | 2 | H | 404.18 | 412.97 | 349
I2 | B | 8 | 2 | H | 396.65 | 405.43 | 349
I2 | B | 9 | 2 | H | 397.09 | 405.96 | 349
I2 | B | 10 | 2 | H | 397.61 | 406.33 | 349
I2 | B | 11 | 2 | H | 403.13 | 412.12 | 349
I2 | B | 12 | 2 | H | 404.46 | 413.33 | 349
I2 | B | 13 | 2 | H | 406.93 | 415.77 | 349
I2 | B | 14 | 2 | H | 410.53 | 419.56 | 349
I2 | B | 15 | 2 | H | 414.16 | 422.67 | 349
I2 | B | 16 | 2 | H | 415.71 | 424.27 | 349
I2 | B | 17 | 2 | H | 416.62 | 425.19 | 349
I2 | B | 18 | 2 | H | 413.53 | 422.33 | 349
I2 | B | 19 | 2 | H | 415.03 | 423.81 | 349
I2 | B | 20 | 2 | H | 413.16 | 421.91 | 349
I2 | B | 21 | 2 | H | 416.21 | 424.95 | 349
I2 | B | 22 | 2 | H | 414.77 | 423.74 | 349
I2 | B | 23 | 2 | H | 419.66 | 428.3 | 349
I2 | B | 24 | 2 | H | 421.34 | 430.06 | 349
I2 | B | 25 | 2 | H | 426.79 | 435.69 | 349
I2 | B | 26 | 2 | H | 426.45 | 435.44 | 349
I2 | B | 27 | 2 | H | 418.62 | 427.33 | 349
I2 | B | 28 | 2 | H | 416.56 | 425.39 | 349
I2 | B | 29 | 2 | H | 418.26 | 427.05 | 349
I2 | B | 30 | 2 | H | 416.28 | 424.96 | 349
I2 | B | 31 | 2 | H | 429.61 | 438.35 | 349
I2 | B | 32 | 2 | H | 428.42 | 437.2 | 349
I3 | A | 1 | 1 | H | 69.1 | 71.71 | 72.5
I3 | A | 2 | 1 | H | 67.75 | 70.01 | 72.5
I3 | A | 3 | 1 | H | 67.67 | 70.11 | 72.5
I3 | A | 4 | 1 | H | 96.9 | 131.64 | 72.5
I3 | A | 5 | 1 | H | 66.94 | 69.3 | 72.5
I3 | A | 6 | 1 | H | 66.87 | 69.08 | 72.5
I3 | A | 7 | 1 | H | 66.79 | 68.99 | 72.5
I3 | A | 8 | 1 | H | 67.71 | 69.74 | 72.5
I3 | A | 9 | 1 | H | 68.47 | 70.43 | 72.5
I3 | A | 10 | 1 | H | 68.74 | 70.74 | 72.5
I3 | A | 11 | 1 | H | 68.51 | 70.55 | 72.5
I3 | A | 12 | 1 | H | 67.62 | 69.63 | 72.5
I3 | A | 13 | 1 | H | 67.47 | 69.5 | 72.5
I3 | A | 14 | 1 | H | 67.94 | 69.91 | 72.5
I3 | A | 15 | 1 | H | 67.8 | 69.91 | 72.5
I3 | A | 16 | 1 | H | 67.92 | 69.9 | 72.5
I3 | A | 17 | 1 | H | 68.45 | 70.47 | 72.5
I3 | A | 18 | 1 | H | 68.11 | 70.17 | 72.5
I3 | A | 19 | 1 | H | 68.83 | 70.99 | 72.5
I3 | A | 20 | 1 | H | 69.19 | 71.24 | 72.5
I3 | A | 21 | 1 | H | 69.07 | 71.11 | 72.5
I3 | A | 22 | 1 | H | 69.32 | 71.34 | 72.5
I3 | A | 23 | 1 | H | 69.26 | 71.43 | 72.5
I3 | A | 24 | 1 | H | 68.84 | 70.98 | 72.5
I3 | A | 25 | 1 | H | 69.5 | 71.65 | 72.5
I3 | A | 26 | 1 | H | 68.8 | 70.94 | 72.5
I3 | A | 27 | 1 | H | 68.62 | 70.93 | 72.5
I3 | A | 28 | 1 | H | 68.66 | 70.93 | 72.5
I3 | A | 29 | 1 | H | 69.22 | 71.56 | 72.5
I3 | A | 30 | 1 | H | 70.42 | 72.55 | 72.5
I3 | A | 31 | 1 | H | 70.4 | 72.67 | 72.5
I3 | A | 32 | 1 | H | 71.07 | 73.33 | 72.5
I3 | A | 1 | 2 | H | 69.06 | 71.7 | 72.5
I3 | A | 2 | 2 | H | 67.73 | 69.97 | 72.5
I3 | A | 3 | 2 | H | 67.73 | 70.09 | 72.5
I3 | A | 4 | 2 | H | 96.7 | 132.89 | 72.5
I3 | A | 5 | 2 | H | 66.9 | 69.32 | 72.5
I3 | A | 6 | 2 | H | 66.9 | 69.09 | 72.5
I3 | A | 7 | 2 | H | 66.76 | 68.93 | 72.5
I3 | A | 8 | 2 | H | 67.59 | 69.71 | 72.5
I3 | A | 9 | 2 | H | 68.43 | 70.47 | 72.5
I3 | A | 10 | 2 | H | 68.69 | 70.74 | 72.5
I3 | A | 11 | 2 | H | 68.49 | 70.56 | 72.5
I3 | A | 12 | 2 | H | 67.54 | 69.57 | 72.5
I3 | A | 13 | 2 | H | 67.41 | 69.47 | 72.5
I3 | A | 14 | 2 | H | 67.87 | 69.95 | 72.5
I3 | A | 15 | 2 | H | 67.89 | 69.87 | 72.5
I3 | A | 16 | 2 | H | 67.87 | 69.93 | 72.5
I3 | A | 17 | 2 | H | 68.37 | 70.49 | 72.5
I3 | A | 18 | 2 | H | 68.14 | 70.15 | 72.5
I3 | A | 19 | 2 | H | 68.83 | 70.95 | 72.5
I3 | A | 20 | 2 | H | 69.2 | 71.2 | 72.5
I3 | A | 21 | 2 | H | 69.07 | 71.1 | 72.5
I3 | A | 22 | 2 | H | 69.32 | 71.35 | 72.5
I3 | A | 23 | 2 | H | 69.25 | 71.39 | 72.5
I3 | A | 24 | 2 | H | 68.81 | 70.99 | 72.5
I3 | A | 25 | 2 | H | 69.49 | 71.61 | 72.5
I3 | A | 26 | 2 | H | 68.76 | 70.91 | 72.5
I3 | A | 27 | 2 | H | 68.58 | 70.94 | 72.5
I3 | A | 28 | 2 | H | 68.66 | 70.96 | 72.5
I3 | A | 29 | 2 | H | 69.25 | 71.54 | 72.5
I3 | A | 30 | 2 | H | 70.4 | 72.56 | 72.5
I3 | A | 31 | 2 | H | 70.41 | 72.64 | 72.5
I3 | A | 32 | 2 | H | 71.08 | 73.28 | 72.5
I3 | A | 1 | 1 | H | 69.18 | 71.84 | 72.5
I3 | A | 2 | 1 | H | 67.82 | 70.15 | 72.5
I3 | B | 3 | 1 | H | 67.8 | 70.21 | 72.5
I3 | B | 4 | 1 | H | 97.09 | 133.99 | 72.5
I3 | B | 5 | 1 | H | 67.04 | 69.42 | 72.5
I3 | B | 6 | 1 | H | 66.94 | 69.2 | 72.5
I3 | B | 7 | 1 | H | 66.88 | 69.11 | 72.5
I3 | B | 8 | 1 | H | 67.82 | 69.8 | 72.5
I3 | B | 9 | 1 | H | 68.55 | 70.59 | 72.5
I3 | B | 10 | 1 | H | 68.82 | 70.87 | 72.5
I3 | B | 11 | 1 | H | 68.62 | 70.71 | 72.5
I3 | B | 12 | 1 | H | 67.73 | 69.73 | 72.5
I3 | B | 13 | 1 | H | 67.56 | 69.61 | 72.5
I3 | B | 14 | 1 | H | 68.01 | 70.08 | 72.5
I3 | B | 15 | 1 | H | 67.92 | 70.05 | 72.5
I3 | B | 16 | 1 | H | 68.02 | 70.05 | 72.5
I3 | B | 17 | 1 | H | 68.54 | 70.64 | 72.5
I3 | B | 18 | 1 | H | 68.19 | 70.33 | 72.5
I3 | B | 19 | 1 | H | 68.89 | 71.09 | 72.5
I3 | B | 20 | 1 | H | 69.3 | 71.33 | 72.5
I3 | B | 21 | 1 | H | 69.19 | 71.23 | 72.5
I3 | B | 22 | 1 | H | 69.44 | 71.49 | 72.5
I3 | B | 23 | 1 | H | 69.37 | 71.53 | 72.5
I3 | B | 24 | 1 | H | 68.95 | 71.1 | 72.5
I3 | B | 25 | 1 | H | 69.58 | 71.76 | 72.5
I3 | B | 26 | 1 | H | 68.88 | 71.06 | 72.5
I3 | B | 27 | 1 | H | 68.73 | 71.07 | 72.5
I3 | B | 28 | 1 | H | 68.77 | 71.06 | 72.5
I3 | B | 29 | 1 | H | 69.34 | 71.69 | 72.5
I3 | B | 30 | 1 | H | 70.52 | 72.68 | 72.5
I3 | B | 31 | 1 | H | 70.49 | 72.78 | 72.5
I3 | B | 32 | 1 | H | 71.18 | 73.47 | 72.5
I3 | B | 1 | 2 | H | 69.18 | 71.81 | 72.5
I3 | B | 2 | 2 | H | 67.87 | 70.04 | 72.5
I3 | B | 3 | 2 | H | 67.81 | 70.22 | 72.5
I3 | B | 4 | 2 | H | 96.94 | 133.13 | 72.5
I3 | B | 5 | 2 | H | 66.97 | 69.4 | 72.5
I3 | B | 6 | 2 | H | 67 | 69.17 | 72.5
I3 | B | 7 | 2 | H | 66.86 | 69.03 | 72.5
I3 | B | 8 | 2 | H | 67.74 | 69.83 | 72.5
I3 | B | 9 | 2 | H | 68.55 | 70.56 | 72.5
I3 | B | 10 | 2 | H | 68.81 | 70.85 | 72.5
I3 | B | 11 | 2 | H | 68.59 | 70.66 | 72.5
I3 | B | 12 | 2 | H | 67.62 | 69.69 | 72.5
I3 | B | 13 | 2 | H | 67.55 | 69.59 | 72.5
I3 | B | 14 | 2 | H | 67.99 | 70.06 | 72.5
I3 | B | 15 | 2 | H | 67.98 | 69.95 | 72.5
I3 | B | 16 | 2 | H | 68 | 70.05 | 72.5
I3 | B | 17 | 2 | H | 68.51 | 70.63 | 72.5
I3 | B | 18 | 2 | H | 68.24 | 70.25 | 72.5
I3 | B | 19 | 2 | H | 68.94 | 71.08 | 72.5
I3 | B | 20 | 2 | H | 69.29 | 71.34 | 72.5
I3 | B | 21 | 2 | H | 69.16 | 71.23 | 72.5
I3 | B | 22 | 2 | H | 69.41 | 71.47 | 72.5
I3 | B | 23 | 2 | H | 69.35 | 71.51 | 72.5
I3 | B | 24 | 2 | H | 68.92 | 71.08 | 72.5
I3 | B | 25 | 2 | H | 69.6 | 71.74 | 72.5
I3 | B | 26 | 2 | H | 68.85 | 71.04 | 72.5
I3 | B | 27 | 2 | H | 68.72 | 71.03 | 72.5
I3 | B | 28 | 2 | H | 68.77 | 71.06 | 72.5
I3 | B | 29 | 2 | H | 69.33 | 71.65 | 72.5
I3 | B | 30 | 2 | H | 70.49 | 72.69 | 72.5
I3 | B | 31 | 2 | H | 70.51 | 72.71 | 72.5
I3 | B | 32 | 2 | H | 71.16 | 73.35 | 72.5
M1 | A | 1 | 1 | H | 188.93 | 198.14 | 135
M1 | A | 2 | 1 | H | 189.59 | 198.88 | 135
M1 | A | 3 | 1 | H | 187.18 | 196.36 | 135
M1 | A | 4 | 1 | H | 185.32 | 194.41 | 135
M1 | A | 5 | 1 | H | 184.02 | 193.14 | 135
M1 | A | 6 | 1 | H | 184.09 | 193.11 | 135
M1 | A | 7 | 1 | H | 184.36 | 193.36 | 135
M1 | A | 8 | 1 | H | 183.84 | 192.83 | 135
M1 | A | 9 | 1 | H | 183.4 | 192.4 | 135
M1 | A | 10 | 1 | H | 183.12 | 192.11 | 135
M1 | A | 11 | 1 | H | 183.53 | 192.48 | 135
M1 | A | 12 | 1 | H | 184.7 | 193.73 | 135
M1 | A | 13 | 1 | H | 184.41 | 193.32 | 135
M1 | A | 14 | 1 | H | 184.28 | 193.31 | 135
M1 | A | 15 | 1 | H | 186.05 | 195 | 135
M1 | A | 16 | 1 | H | 187 | 195.84 | 135
M1 | B | 1 | 1 | H | 193.23 | 202.77 | 135
M1 | B | 2 | 1 | H | 193.84 | 203.4 | 135
M1 | B | 3 | 1 | H | 191.34 | 200.87 | 135
M1 | B | 4 | 1 | H | 189.39 | 198.86 | 135
M1 | B | 5 | 1 | H | 188.15 | 197.44 | 135
M1 | B | 6 | 1 | H | 188.14 | 197.4 | 135
M1 | B | 7 | 1 | H | 188.41 | 197.64 | 135
M1 | B | 8 | 1 | H | 187.81 | 197.04 | 135
M1 | B | 9 | 1 | H | 187.37 | 196.53 | 135
M1 | B | 10 | 1 | H | 187.08 | 196.38 | 135
M1 | B | 11 | 1 | H | 187.4 | 196.67 | 135
M1 | B | 12 | 1 | H | 188.68 | 197.97 | 135
M1 | B | 13 | 1 | H | 188.39 | 197.6 | 135
M1 | B | 14 | 1 | H | 188.23 | 197.55 | 135
M1 | B | 15 | 1 | H | 190.02 | 199.25 | 135
M1 | B | 16 | 1 | H | 190.86 | 200.06 | 135
M2 | A | 1 | 1 | H | 181 | 186.45 | 127
M2 | A | 2 | 1 | H | 181.99 | 186.97 | 127
M2 | A | 3 | 1 | H | 180.8 | 185.76 | 127
M2 | A | 4 | 1 | H | 179.81 | 185.79 | 127
M2 | A | 5 | 1 | H | 180.03 | 184.92 | 127
M2 | A | 6 | 1 | H | 178.62 | 184.56 | 127
M2 | A | 7 | 1 | H | 178.38 | 184.24 | 127
M2 | A | 8 | 1 | H | 177.78 | 183.59 | 127
M2 | A | 9 | 1 | H | 178 | 182.92 | 127
M2 | A | 10 | 1 | H | 177.84 | 183.58 | 127
M2 | A | 11 | 1 | H | 178.65 | 184.63 | 127
M2 | A | 12 | 1 | H | 179.06 | 184.47 | 127
M2 | A | 13 | 1 | H | 178.45 | 184.33 | 127
M2 | A | 14 | 1 | H | 178.56 | 183.57 | 127
M2 | A | 15 | 1 | H | 177.96 | 183.85 | 127
M2 | A | 16 | 1 | H | 178.28 | 183.31 | 127
M2 | B | 1 | 1 | H | 182.86 | 188.31 | 127
M2 | B | 2 | 1 | H | 183.77 | 188.91 | 127
M2 | B | 3 | 1 | H | 182.54 | 187.65 | 127
M2 | B | 4 | 1 | H | 181.54 | 187.62 | 127
M2 | B | 5 | 1 | H | 181.78 | 186.73 | 127
M2 | B | 6 | 1 | H | 180.32 | 186.3 | 127
M2 | B | 7 | 1 | H | 180.07 | 185.93 | 127
M2 | B | 8 | 1 | H | 179.46 | 185.37 | 127
M2 | B | 9 | 1 | H | 179.7 | 184.73 | 127
M2 | B | 10 | 1 | H | 179.49 | 185.3 | 127
M2 | B | 11 | 1 | H | 180.32 | 186.28 | 127
M2 | B | 12 | 1 | H | 180.77 | 186.16 | 127
M2 | B | 13 | 1 | H | 180.11 | 185.99 | 127
M2 | B | 14 | 1 | H | 180.22 | 185.33 | 127
M2 | B | 15 | 1 | H | 179.58 | 185.62 | 127
M2 | B | 16 | 1 | H | 179.94 | 185.06 | 127
M3 | A | 1 | 1 | H | 135.92 | 140.11 | 107
M3 | A | 2 | 1 | H | 135.93 | 140.18 | 107
M3 | A | 3 | 1 | H | 135.26 | 139.36 | 107
M3 | A | 4 | 1 | H | 134.59 | 138.71 | 107
M3 | A | 5 | 1 | H | 134.41 | 138.51 | 107
M3 | A | 6 | 1 | H | 134.81 | 138.92 | 107
M3 | A | 7 | 1 | H | 135.51 | 139.63 | 107
M3 | A | 8 | 1 | H | 135.48 | 139.55 | 107
M3 | A | 9 | 1 | H | 135.89 | 139.95 | 107
M3 | A | 10 | 1 | H | 135.7 | 139.77 | 107
M3 | A | 11 | 1 | H | 135.42 | 139.46 | 107
M3 | A | 12 | 1 | H | 136.03 | 140.09 | 107
M3 | A | 13 | 1 | H | 137 | 141.07 | 107
M3 | A | 14 | 1 | H | 136.82 | 140.94 | 107
M3 | A | 15 | 1 | H | 136.79 | 140.88 | 107
M3 | A | 16 | 1 | H | 137.63 | 141.72 | 107
M3 | B | 1 | 1 | H | 136.4 | 140.67 | 107
M3 | B | 2 | 1 | H | 136.42 | 140.71 | 107
M3 | B | 3 | 1 | H | 135.7 | 139.82 | 107
M3 | B | 4 | 1 | H | 135.06 | 139.2 | 107
M3 | B | 5 | 1 | H | 134.86 | 138.97 | 107
M3 | B | 6 | 1 | H | 135.28 | 139.38 | 107
M3 | B | 7 | 1 | H | 135.97 | 140.1 | 107
M3 | B | 8 | 1 | H | 135.93 | 140.03 | 107
M3 | B | 9 | 1 | H | 136.3 | 140.4 | 107
M3 | B | 10 | 1 | H | 136.13 | 140.25 | 107
M3 | B | 11 | 1 | H | 135.85 | 139.96 | 107
M3 | B | 12 | 1 | H | 136.44 | 140.53 | 107
M3 | B | 13 | 1 | H | 137.41 | 141.51 | 107
M3 | B | 14 | 1 | H | 137.28 | 141.4 | 107
M3 | B | 15 | 1 | H | 137.22 | 141.29 | 107
M3 | B | 16 | 1 | H | 138.02 | 142.16 | 107
M4 | A | 1 | 1 | H | 109.96 | 118.5 | 78
M4 | A | 2 | 1 | H | 109.97 | 118.69 | 78
M4 | A | 3 | 1 | H | 110.17 | 118.25 | 78
M4 | A | 4 | 1 | H | 110.29 | 118.94 | 78
M4 | A | 5 | 1 | H | 110.06 | 118.61 | 78
M4 | A | 6 | 1 | H | 110.32 | 119.01 | 78
M4 | A | 7 | 1 | H | 110.36 | 119 | 78
M4 | A | 8 | 1 | H | 110.86 | 119.69 | 78
M4 | A | 9 | 1 | H | 110.94 | 119.59 | 78
M4 | A | 10 | 1 | H | 111.74 | 120.1 | 78
M4 | A | 11 | 1 | H | 111.18 | 120.08 | 78
M4 | A | 12 | 1 | H | 111.9 | 120.89 | 78
M4 | A | 13 | 1 | H | 112.49 | 121.62 | 78
M4 | A | 14 | 1 | H | 113.09 | 122.02 | 78
M4 | A | 15 | 1 | H | 113.35 | 122.26 | 78
M4 | A | 16 | 1 | H | 113.02 | 121.93 | 78
M4 | B | 1 | 1 | H | 109.87 | 118.41 | 78
M4 | B | 2 | 1 | H | 109.87 | 118.6 | 78
M4 | B | 3 | 1 | H | 110.07 | 118.15 | 78
M4 | B | 4 | 1 | H | 110.17 | 118.83 | 78
M4 | B | 5 | 1 | H | 109.96 | 118.5 | 78
M4 | B | 6 | 1 | H | 110.23 | 118.9 | 78
M4 | B | 7 | 1 | H | 110.26 | 118.89 | 78
M4 | B | 8 | 1 | H | 110.73 | 119.57 | 78
M4 | B | 9 | 1 | H | 110.79 | 119.48 | 78
M4 | B | 10 | 1 | H | 111.64 | 119.98 | 78
M4 | B | 11 | 1 | H | 111.07 | 119.97 | 78
M4 | B | 12 | 1 | H | 111.8 | 120.76 | 78
M4 | B | 13 | 1 | H | 112.38 | 121.49 | 78
M4 | B | 14 | 1 | H | 112.97 | 121.91 | 78
M4 | B | 15 | 1 | H | 113.21 | 122.16 | 78
M4 | B | 16 | 1 | H | 112.9 | 121.81 | 78
M5 | A | 1 | 1 | H | 71.84 | 74.69 | 59
M5 | A | 2 | 1 | H | 72.21 | 75.12 | 59
M5 | A | 3 | 1 | H | 72.25 | 75.07 | 59
M5 | A | 4 | 1 | H | 71.58 | 74.38 | 59
M5 | A | 5 | 1 | H | 71.22 | 74 | 59
M5 | A | 6 | 1 | H | 71.64 | 74.4 | 59
M5 | A | 7 | 1 | H | 71.99 | 74.79 | 59
M5 | A | 8 | 1 | H | 71.57 | 74.35 | 59
M5 | A | 9 | 1 | H | 71.94 | 74.71 | 59
M5 | A | 10 | 1 | H | 71.9 | 74.68 | 59
M5 | A | 11 | 1 | H | 72.02 | 74.81 | 59
M5 | A | 12 | 1 | H | 72.07 | 74.82 | 59
M5 | A | 13 | 1 | H | 72.23 | 75 | 59
M5 | A | 14 | 1 | H | 72.47 | 75.27 | 59
M5 | A | 15 | 1 | H | 72.84 | 75.59 | 59
M5 | A | 16 | 1 | H | 72.33 | 75.07 | 59
M5 | B | 1 | 1 | H | 71.52 | 74.36 | 59
M5 | B | 2 | 1 | H | 71.88 | 74.8 | 59
M5 | B | 3 | 1 | H | 71.92 | 74.74 | 59
M5 | B | 4 | 1 | H | 71.27 | 74.05 | 59
M5 | B | 5 | 1 | H | 70.9 | 73.67 | 59
M5 | B | 6 | 1 | H | 71.3 | 74.09 | 59
M5 | B | 7 | 1 | H | 71.66 | 74.43 | 59
M5 | B | 8 | 1 | H | 71.27 | 74.03 | 59
M5 | B | 9 | 1 | H | 71.6 | 74.36 | 59
M5 | B | 10 | 1 | H | 71.58 | 74.34 | 59
M5 | B | 11 | 1 | H | 71.69 | 74.46 | 59
M5 | B | 12 | 1 | H | 71.71 | 74.49 | 59
M5 | B | 13 | 1 | H | 71.9 | 74.69 | 59
M5 | B | 14 | 1 | H | 72.15 | 74.92 | 59
M5 | B | 15 | 1 | H | 72.52 | 75.26 | 59
M5 | B | 16 | 1 | H | 72.01 | 74.77 | 59
M6 | A | 1 | 1 | H | 46.68 | 48.42 | 41
M6 | A | 2 | 1 | H | 47.11 | 48.91 | 41
M6 | A | 3 | 1 | H | 47.21 | 49 | 41
M6 | A | 4 | 1 | H | 46.76 | 48.51 | 41
M6 | A | 5 | 1 | H | 46.96 | 48.75 | 41
M6 | A | 6 | 1 | H | 47.32 | 49.1 | 41
M6 | A | 7 | 1 | H | 47.13 | 48.96 | 41
M6 | A | 8 | 1 | H | 47.2 | 49.04 | 41
M6 | A | 9 | 1 | H | 47.64 | 49.4 | 41
M6 | A | 10 | 1 | H | 47.66 | 49.43 | 41
M6 | A | 11 | 1 | H | 47.66 | 49.44 | 41
M6 | A | 12 | 1 | H | 47.7 | 49.49 | 41
M6 | A | 13 | 1 | H | 48.2 | 50.01 | 41
M6 | A | 14 | 1 | H | 48.21 | 50.03 | 41
M6 | A | 15 | 1 | H | 48 | 49.8 | 41
M6 | A | 16 | 1 | H | 47.94 | 49.73 | 41
M6 | B | 1 | 1 | H | 46.5 | 48.22 | 41
M6 | B | 2 | 1 | H | 46.89 | 48.71 | 41
M6 | B | 3 | 1 | H | 47 | 48.81 | 41
M6 | B | 4 | 1 | H | 46.56 | 48.3 | 41
M6 | B | 5 | 1 | H | 46.78 | 48.55 | 41
M6 | B | 6 | 1 | H | 47.11 | 48.92 | 41
M6 | B | 7 | 1 | H | 46.96 | 48.75 | 41
M6 | B | 8 | 1 | H | 47.04 | 48.83 | 41
M6 | B | 9 | 1 | H | 47.41 | 49.21 | 41
M6 | B | 10 | 1 | H | 47.45 | 49.24 | 41
M6 | B | 11 | 1 | H | 47.45 | 49.26 | 41
M6 | B | 12 | 1 | H | 47.49 | 49.32 | 41
M6 | B | 13 | 1 | H | 48.01 | 49.8 | 41
M6 | B | 14 | 1 | H | 48 | 49.8 | 41
M6 | B | 15 | 1 | H | 47.79 | 49.59 | 41
M6 | B | 16 | 1 | H | 47.7 | 49.55 | 41
M7 | A | 1 | 1 | H | 36.72 | 37.87 | 29
M7 | A | 2 | 1 | H | 36.97 | 38.16 | 29
M7 | A | 3 | 1 | H | 36.8 | 37.96 | 29
M7 | A | 4 | 1 | H | 36.64 | 37.8 | 29
M7 | A | 5 | 1 | H | 36.74 | 37.9 | 29
M7 | A | 6 | 1 | H | 36.76 | 37.92 | 29
M7 | A | 7 | 1 | H | 36.94 | 38.1 | 29
M7 | A | 8 | 1 | H | 37.08 | 38.24 | 29
M7 | A | 9 | 1 | H | 37.3 | 38.46 | 29
M7 | A | 10 | 1 | H | 36.99 | 38.14 | 29
M7 | A | 11 | 1 | H | 36.94 | 38.1 | 29
M7 | A | 12 | 1 | H | 37.05 | 38.2 | 29
M7 | A | 13 | 1 | H | 37.34 | 38.52 | 29
M7 | A | 14 | 1 | H | 37.27 | 38.44 | 29
M7 | A | 15 | 1 | H | 37.27 | 38.45 | 29
M7 | A | 16 | 1 | H | 37.09 | 38.25 | 29
M7 | B | 1 | 1 | H | 36.75 | 37.89 | 29
M7 | B | 2 | 1 | H | 37 | 38.18 | 29
M7 | B | 3 | 1 | H | 36.83 | 37.98 | 29
M7 | B | 4 | 1 | H | 36.68 | 37.81 | 29
M7 | B | 5 | 1 | H | 36.77 | 37.92 | 29
M7 | B | 6 | 1 | H | 36.8 | 37.94 | 29
M7 | B | 7 | 1 | H | 36.97 | 38.12 | 29
M7 | B | 8 | 1 | H | 37.11 | 38.26 | 29
M7 | B | 9 | 1 | H | 37.33 | 38.49 | 29
M7 | B | 10 | 1 | H | 37.02 | 38.17 | 29
M7 | B | 11 | 1 | H | 36.98 | 38.12 | 29
M7 | B | 12 | 1 | H | 37.08 | 38.23 | 29
M7 | B | 13 | 1 | H | 37.37 | 38.54 | 29
M7 | B | 14 | 1 | H | 37.3 | 38.46 | 29
M7 | B | 15 | 1 | H | 37.3 | 38.47 | 29
M7 | B | 16 | 1 | H | 37.13 | 38.28 | 29
M8 | A | 1 | 1 | H | 114.51 | 117.32 | 164.9
M8 | A | 2 | 1 | H | 115.06 | 117.87 | 164.9
M8 | A | 3 | 1 | H | 115.91 | 118.77 | 164.9
M8 | A | 4 | 1 | H | 118.26 | 121.22 | 164.9
M8 | A | 5 | 1 | H | 118.4 | 121.33 | 164.9
M8 | A | 6 | 1 | H | 120.08 | 122.56 | 164.9
M8 | A | 7 | 1 | H | 120.72 | 123.67 | 164.9
M8 | A | 8 | 1 | H | 121.39 | 123.84 | 164.9
M8 | A | 9 | 1 | H | 122.81 | 125.28 | 164.9
M8 | A | 10 | 1 | H | 121.6 | 124.67 | 164.9
M8 | A | 11 | 1 | H | 123.49 | 126.66 | 164.9
M8 | A | 12 | 1 | H | 123.4 | 126.65 | 164.9
M8 | A | 13 | 1 | H | 125.4 | 128.66 | 164.9
M8 | A | 14 | 1 | H | 124.71 | 128 | 164.9
M8 | A | 15 | 1 | H | 126.83 | 129.53 | 164.9
M8 | A | 16 | 1 | H | 126.94 | 129.97 | 164.9
M8 | B | 1 | 1 | H | 114.92 | 117.74 | 164.9
M8 | B | 2 | 1 | H | 115.45 | 118.31 | 164.9
M8 | B | 3 | 1 | H | 116.32 | 119.2 | 164.9
M8 | B | 4 | 1 | H | 118.67 | 121.66 | 164.9
M8 | B | 5 | 1 | H | 118.82 | 121.77 | 164.9
M8 | B | 6 | 1 | H | 120.49 | 123.01 | 164.9
M8 | B | 7 | 1 | H | 121.14 | 124.12 | 164.9
M8 | B | 8 | 1 | H | 121.82 | 124.29 | 164.9
M8 | B | 9 | 1 | H | 123.24 | 125.75 | 164.9
M8 | B | 10 | 1 | H | 122.02 | 125.11 | 164.9
M8 | B | 11 | 1 | H | 123.92 | 127.13 | 164.9
M8 | B | 12 | 1 | H | 123.84 | 127.11 | 164.9
M8 | B | 13 | 1 | H | 125.84 | 129.12 | 164.9
M8 | B | 14 | 1 | H | 125.14 | 128.48 | 164.9
M8 | B | 15 | 1 | H | 127.27 | 130 | 164.9
M8 | B | 16 | 1 | H | 127.37 | 130.17 | 164.9
M9 | A | 1 | 1 | H | 85.4 | 85.76 | 77.1
M9 | A | 2 | 1 | H | 85.6 | 86.03 | 77.1
M9 | A | 3 | 1 | H | 86.71 | 87.02 | 77.1
M9 | A | 4 | 1 | H | 89.24 | 89.83 | 77.1
M9 | A | 5 | 1 | H | 88.83 | 89.25 | 77.1
M9 | A | 6 | 1 | H | 90.01 | 90.59 | 77.1
M9 | A | 7 | 1 | H | 91.48 | 92.07 | 77.1
M9 | A | 8 | 1 | H | 92.36 | 93 | 77.1
M9 | A | 9 | 1 | H | 93.34 | 93.91 | 77.1
M9 | A | 10 | 1 | H | 94.48 | 95.14 | 77.1
M9 | A | 11 | 1 | H | 94.52 | 95.19 | 77.1
M9 | A | 12 | 1 | H | 96.23 | 96.85 | 77.1
M9 | A | 13 | 1 | H | 96.56 | 97.17 | 77.1
M9 | A | 14 | 1 | H | 98.5 | 99.17 | 77.1
M9 | A | 15 | 1 | H | 98.89 | 99.41 | 77.1
M9 | A | 16 | 1 | H | 101.66 | 102.27 | 77.1
M9 | B | 1 | 1 | H | 85.65 | 86 | 77.1
M9 | B | 2 | 1 | H | 85.8 | 86.27 | 77.1
M9 | B | 3 | 1 | H | 86.95 | 87.24 | 77.1
M9 | B | 4 | 1 | H | 89.5 | 90.08 | 77.1
M9 | B | 5 | 1 | H | 89.08 | 89.52 | 77.1
M9 | B | 6 | 1 | H | 90.29 | 90.85 | 77.1
M9 | B | 7 | 1 | H | 91.68 | 92.32 | 77.1
M9 | B | 8 | 1 | H | 92.64 | 93.29 | 77.1
M9 | B | 9 | 1 | H | 93.55 | 94.18 | 77.1
M9 | B | 10 | 1 | H | 94.72 | 95.41 | 77.1
M9 | B | 11 | 1 | H | 94.79 | 95.47 | 77.1
M9 | B | 12 | 1 | H | 96.49 | 97.1 | 77.1
M9 | B | 13 | 1 | H | 96.8 | 97.43 | 77.1
M9 | B | 14 | 1 | H | 98.8 | 99.46 | 77.1
M9 | B | 15 | 1 | H | 99.19 | 99.68 | 77.1
M9 | B | 16 | 1 | H | 101.96 | 102.49 | 77.1
M10 | A | 1 | 1 | H | 80.01 | 81.35 | 71.2
M10 | A | 2 | 1 | H | 78.92 | 80.4 | 71.2
M10 | A | 3 | 1 | H | 79.37 | 80.77 | 71.2
M10 | A | 4 | 1 | H | 79.94 | 81.01 | 71.2
M10 | A | 5 | 1 | H | 80.38 | 81.21 | 71.2
M10 | A | 6 | 1 | H | 81.67 | 82.55 | 71.2
M10 | A | 7 | 1 | H | 81.27 | 82.26 | 71.2
M10 | A | 8 | 1 | H | 82.93 | 83.99 | 71.2
M10 | A | 9 | 1 | H | 82.17 | 83.27 | 71.2
M10 | A | 10 | 1 | H | 82.56 | 83.53 | 71.2
M10 | A | 11 | 1 | H | 82.65 | 83.66 | 71.2
M10 | A | 12 | 1 | H | 83.69 | 84.83 | 71.2
M10 | A | 13 | 1 | H | 82.66 | 83.54 | 71.2
M10 | A | 14 | 1 | H | 83.9 | 85.19 | 71.2
M10 | A | 15 | 1 | H | 83.99 | 85.14 | 71.2
M10 | A | 16 | 1 | H | 85.04 | 86.15 | 71.2
M10 | B | 1 | 1 | H | 80.12 | 81.49 | 71.2
M10 | B | 2 | 1 | H | 79.02 | 80.53 | 71.2
M10 | B | 3 | 1 | H | 79.51 | 80.91 | 71.2
M10 | B | 4 | 1 | H | 80.08 | 81.15 | 71.2
M10 | B | 5 | 1 | H | 80.5 | 81.34 | 71.2
M10 | B | 6 | 1 | H | 81.8 | 82.69 | 71.2
M10 | B | 7 | 1 | H | 81.39 | 82.4 | 71.2
M10 | B | 8 | 1 | H | 83.06 | 84.13 | 71.2
M10 | B | 9 | 1 | H | 82.3 | 83.41 | 71.2
M10 | B | 10 | 1 | H | 82.7 | 83.66 | 71.2
M10 | B | 11 | 1 | H | 82.78 | 83.8 | 71.2
M10 | B | 12 | 1 | H | 83.81 | 84.98 | 71.2
M10 | B | 13 | 1 | H | 82.8 | 83.68 | 71.2
M10 | B | 14 | 1 | H | 84.03 | 85.34 | 71.2
M10 | B | 15 | 1 | H | 84.11 | 85.27 | 71.2
M10 | B | 16 | 1 | H | 85.17 | 86.29 | 71.2
M11 | A | 1 | 1 | H | 34.25 | 34.03 | 31.8
M11 | A | 2 | 1 | H | 32.79 | 32.6 | 31.8
M11 | A | 3 | 1 | H | 34.25 | 34.01 | 31.8
M11 | A | 4 | 1 | H | 34.22 | 33.83 | 31.8
M11 | A | 5 | 1 | H | 32.9 | 32.67 | 31.8
M11 | A | 6 | 1 | H | 33.2 | 32.8 | 31.8
M11 | A | 7 | 1 | H | 33.42 | 33.02 | 31.8
M11 | A | 8 | 1 | H | 34.36 | 34.12 | 31.8
M11 | A | 9 | 1 | H | 34.02 | 33.75 | 31.8
M11 | A | 10 | 1 | H | 33.56 | 33.31 | 31.8
M11 | A | 11 | 1 | H | 33.03 | 32.79 | 31.8
M11 | A | 12 | 1 | H | 33.07 | 32.83 | 31.8
M11 | A | 13 | 1 | H | 32.81 | 32.46 | 31.8
M11 | A | 14 | 1 | H | 33.17 | 32.94 | 31.8
M11 | A | 15 | 1 | H | 33.59 | 33.35 | 31.8
M11 | A | 16 | 1 | H | 34.15 | 33.72 | 31.8
M11 | B | 1 | 1 | H | 34.27 | 34.06 | 31.8
M11 | B | 2 | 1 | H | 32.81 | 32.63 | 31.8
M11 | B | 3 | 1 | H | 34.27 | 34.04 | 31.8
M11 | B | 4 | 1 | H | 34.24 | 33.85 | 31.8
M11 | B | 5 | 1 | H | 32.92 | 32.69 | 31.8
M11 | B | 6 | 1 | H | 33.22 | 32.82 | 31.8
M11 | B | 7 | 1 | H | 33.44 | 33.1 | 31.8
M11 | B | 8 | 1 | H | 34.38 | 34.14 | 31.8
M11 | B | 9 | 1 | H | 34.04 | 33.69 | 31.8
M11 | B | 10 | 1 | H | 33.58 | 33.33 | 31.8
M11 | B | 11 | 1 | H | 33.05 | 32.81 | 31.8
M11 | B | 12 | 1 | H | 33.09 | 32.86 | 31.8
M11 | B | 13 | 1 | H | 32.82 | 32.49 | 31.8
M11 | B | 14 | 1 | H | 33.19 | 32.97 | 31.8
M11 | B | 15 | 1 | H | 33.61 | 33.37 | 31.8
M11 | B | 16 | 1 | H | 34.17 | 33.75 | 31.8
M1 | A | 1 | 1 | L | NaN | NaN | 615
M1 | A | 2 | 1 | L | NaN | NaN | 615
M1 | A | 3 | 1 | L | NaN | NaN | 615
M1 | A | 4 | 1 | L | NaN | NaN | 615
M1 | A | 5 | 1 | L | NaN | NaN | 615
M1 | A | 6 | 1 | L | NaN | NaN | 615
M1 | A | 7 | 1 | L | NaN | NaN | 615
M1 | A | 8 | 1 | L | NaN | NaN | 615
M1 | A | 9 | 1 | L | NaN | NaN | 615
M1 | A | 10 | 1 | L | NaN | NaN | 615
M1 | A | 11 | 1 | L | NaN | NaN | 615
M1 | A | 12 | 1 | L | NaN | NaN | 615
M1 | A | 13 | 1 | L | NaN | NaN | 615
M1 | A | 14 | 1 | L | NaN | NaN | 615
M1 | A | 15 | 1 | L | NaN | NaN | 615
M1 | A | 16 | 1 | L | NaN | NaN | 615
M1 | B | 1 | 1 | L | NaN | NaN | 615
M1 | B | 2 | 1 | L | NaN | NaN | 615
M1 | B | 3 | 1 | L | NaN | NaN | 615
M1 | B | 4 | 1 | L | NaN | NaN | 615
M1 | B | 5 | 1 | L | NaN | NaN | 615
M1 | B | 6 | 1 | L | NaN | NaN | 615
M1 | B | 7 | 1 | L | NaN | NaN | 615
M1 | B | 8 | 1 | L | NaN | NaN | 615
M1 | B | 9 | 1 | L | NaN | NaN | 615
M1 | B | 10 | 1 | L | NaN | NaN | 615
M1 | B | 11 | 1 | L | NaN | NaN | 615
M1 | B | 12 | 1 | L | NaN | NaN | 615
M1 | B | 13 | 1 | L | NaN | NaN | 615
M1 | B | 14 | 1 | L | NaN | NaN | 615
M1 | B | 15 | 1 | L | NaN | NaN | 615
M1 | B | 16 | 1 | L | NaN | NaN | 615
M2 | A | 1 | 1 | L | NaN | NaN | 687
M2 | A | 2 | 1 | L | NaN | NaN | 687
M2 | A | 3 | 1 | L | NaN | NaN | 687
M2 | A | 4 | 1 | L | NaN | NaN | 687
M2 | A | 5 | 1 | L | NaN | NaN | 687
M2 | A | 6 | 1 | L | NaN | NaN | 687
M2 | A | 7 | 1 | L | NaN | NaN | 687
M2 | A | 8 | 1 | L | NaN | NaN | 687
M2 | A | 9 | 1 | L | NaN | NaN | 687
M2 | A | 10 | 1 | L | NaN | NaN | 687
M2 | A | 11 | 1 | L | NaN | NaN | 687
M2 | A | 12 | 1 | L | NaN | NaN | 687
M2 | A | 13 | 1 | L | NaN | NaN | 687
M2 | A | 14 | 1 | L | NaN | NaN | 687
M2 | A | 15 | 1 | L | NaN | NaN | 687
M2 | A | 16 | 1 | L | NaN | NaN | 687
M2 | B | 1 | 1 | L | NaN | NaN | 687
M2 | B | 2 | 1 | L | NaN | NaN | 687
M2 | B | 3 | 1 | L | NaN | NaN | 687
M2 | B | 4 | 1 | L | NaN | NaN | 687
M2 | B | 5 | 1 | L | NaN | NaN | 687
M2 | B | 6 | 1 | L | NaN | NaN | 687
M2 | B | 7 | 1 | L | NaN | NaN | 687
M2 | B | 8 | 1 | L | NaN | NaN | 687
M2 | B | 9 | 1 | L | NaN | NaN | 687
M2 | B | 10 | 1 | L | NaN | NaN | 687
M2 | B | 11 | 1 | L | NaN | NaN | 687
M2 | B | 12 | 1 | L | NaN | NaN | 687
M2 | B | 13 | 1 | L | NaN | NaN | 687
M2 | B | 14 | 1 | L | NaN | NaN | 687
M2 | B | 15 | 1 | L | NaN | NaN | 687
M2 | B | 16 | 1 | L | NaN | NaN | 687
M3 | A | 1 | 1 | L | NaN | NaN | 702
M3 | A | 2 | 1 | L | NaN | NaN | 702
M3 | A | 3 | 1 | L | NaN | NaN | 702
M3 | A | 4 | 1 | L | NaN | NaN | 702
M3 | A | 5 | 1 | L | NaN | NaN | 702
M3 | A | 6 | 1 | L | NaN | NaN | 702
M3 | A | 7 | 1 | L | NaN | NaN | 702
M3 | A | 8 | 1 | L | NaN | NaN | 702
M3 | A | 9 | 1 | L | NaN | NaN | 702
M3 | A | 10 | 1 | L | NaN | NaN | 702
M3 | A | 11 | 1 | L | NaN | NaN | 702
M3 | A | 12 | 1 | L | NaN | NaN | 702
M3 | A | 13 | 1 | L | NaN | NaN | 702
M3 | A | 14 | 1 | L | NaN | NaN | 702
M3 | A | 15 | 1 | L | NaN | NaN | 702
M3 | A | 16 | 1 | L | NaN | NaN | 702
M3 | B | 1 | 1 | L | NaN | NaN | 702
M3 | B | 2 | 1 | L | NaN | NaN | 702
M3 | B | 3 | 1 | L | NaN | NaN | 702
M3 | B | 4 | 1 | L | NaN | NaN | 702
M3 | B | 5 | 1 | L | NaN | NaN | 702
M3 | B | 6 | 1 | L | NaN | NaN | 702
M3 | B | 7 | 1 | L | NaN | NaN | 702
M3 | B | 8 | 1 | L | NaN | NaN | 702
M3 | B | 9 | 1 | L | NaN | NaN | 702
M3 | B | 10 | 1 | L | NaN | NaN | 702
M3 | B | 11 | 1 | L | NaN | NaN | 702
M3 | B | 12 | 1 | L | NaN | NaN | 702
M3 | B | 13 | 1 | L | NaN | NaN | 702
M3 | B | 14 | 1 | L | NaN | NaN | 702
M3 | B | 15 | 1 | L | NaN | NaN | 702
M3 | B | 16 | 1 | L | NaN | NaN | 702
M4 | A | 1 | 1 | L | 857.93 | 877.97 | 667
M4 | A | 2 | 1 | L | 859.8 | 880.22 | 667
M4 | A | 3 | 1 | L | 860.01 | 880.47 | 667
M4 | A | 4 | 1 | L | 858.18 | 878.58 | 667
M4 | A | 5 | 1 | L | 855.22 | 875.57 | 667
M4 | A | 6 | 1 | L | 857.13 | 877.53 | 667
M4 | A | 7 | 1 | L | 861.34 | 881.7 | 667
M4 | A | 8 | 1 | L | 863.68 | 883.9 | 667
M4 | A | 9 | 1 | L | 864.78 | 884.44 | 667
M4 | A | 10 | 1 | L | 867.12 | 887.59 | 667
M4 | A | 11 | 1 | L | 864.88 | 885.31 | 667
M4 | A | 12 | 1 | L | 874.73 | 895.37 | 667
M4 | A | 13 | 1 | L | 879.31 | 899.92 | 667
M4 | A | 14 | 1 | L | 878.83 | 899.51 | 667
M4 | A | 15 | 1 | L | 877.3 | 897.79 | 667
M4 | A | 16 | 1 | L | 875.8 | 896.25 | 667
M4 | B | 1 | 1 | L | 857.31 | 877.53 | 667
M4 | B | 2 | 1 | L | 859.05 | 879.71 | 667
M4 | B | 3 | 1 | L | 859.28 | 879.91 | 667
M4 | B | 4 | 1 | L | 857.42 | 878.03 | 667
M4 | B | 5 | 1 | L | 854.45 | 874.96 | 667
M4 | B | 6 | 1 | L | 856.34 | 876.83 | 667
M4 | B | 7 | 1 | L | 860.45 | 881.05 | 667
M4 | B | 8 | 1 | L | 862.76 | 883.29 | 667
M4 | B | 9 | 1 | L | 863.92 | 883.74 | 667
M4 | B | 10 | 1 | L | 866.2 | 886.89 | 667
M4 | B | 11 | 1 | L | 864.1 | 884.56 | 667
M4 | B | 12 | 1 | L | 873.85 | 894.64 | 667
M4 | B | 13 | 1 | L | 878.39 | 899.18 | 667
M4 | B | 14 | 1 | L | 877.88 | 898.78 | 667
M4 | B | 15 | 1 | L | 876.3 | 897.08 | 667
M4 | B | 16 | 1 | L | 874.86 | 895.53 | 667
M5 | A | 1 | 1 | L | 721.45 | 738.08 | 651
M5 | A | 2 | 1 | L | 724.82 | 741.64 | 651
M5 | A | 3 | 1 | L | 724.32 | 740.51 | 651
M5 | A | 4 | 1 | L | 718.52 | 735.55 | 651
M5 | A | 5 | 1 | L | 717.61 | 734.94 | 651
M5 | A | 6 | 1 | L | 721.23 | 737.26 | 651
M5 | A | 7 | 1 | L | 724.55 | 740.5 | 651
M5 | A | 8 | 1 | L | 721.79 | 737.74 | 651
M5 | A | 9 | 1 | L | 721.22 | 738.63 | 651
M5 | A | 10 | 1 | L | 724.06 | 740.07 | 651
M5 | A | 11 | 1 | L | 722.4 | 738.64 | 651
M5 | A | 12 | 1 | L | 723.56 | 740.1 | 651
M5 | A | 13 | 1 | L | 725.97 | 742.09 | 651
M5 | A | 14 | 1 | L | 730.55 | 746.65 | 651
M5 | A | 15 | 1 | L | 729.91 | 746.07 | 651
M5 | A | 16 | 1 | L | 727.15 | 742.93 | 651
M5 | B | 1 | 1 | L | 716.73 | 734.84 | 651
M5 | B | 2 | 1 | L | 720.82 | 738.4 | 651
M5 | B | 3 | 1 | L | 720.24 | 737.33 | 651
M5 | B | 4 | 1 | L | 715.49 | 732.38 | 651
M5 | B | 5 | 1 | L | 714.96 | 731.69 | 651
M5 | B | 6 | 1 | L | 717.85 | 733.94 | 651
M5 | B | 7 | 1 | L | 721.26 | 737.21 | 651
M5 | B | 8 | 1 | L | 718.17 | 734.47 | 651
M5 | B | 9 | 1 | L | 718.81 | 735.25 | 651
M5 | B | 10 | 1 | L | 718.85 | 736.67 | 651
M5 | B | 11 | 1 | L | 717.77 | 735.39 | 651
M5 | B | 12 | 1 | L | 719.88 | 736.8 | 651
M5 | B | 13 | 1 | L | 722.8 | 738.83 | 651
M5 | B | 14 | 1 | L | 727.1 | 743.2 | 651
M5 | B | 15 | 1 | L | 726.63 | 742.74 | 651
M5 | B | 16 | 1 | L | 722.83 | 739.73 | 651
M7 | A | 1 | 1 | L | 414.95 | 425.04 | 349
M7 | A | 2 | 1 | L | 416.5 | 426.69 | 349
M7 | A | 3 | 1 | L | 414.75 | 424.76 | 349
M7 | A | 4 | 1 | L | 413.12 | 422.96 | 349
M7 | A | 5 | 1 | L | 413.99 | 423.85 | 349
M7 | A | 6 | 1 | L | 415.09 | 424.91 | 349
M7 | A | 7 | 1 | L | 415.76 | 425.49 | 349
M7 | A | 8 | 1 | L | 417.59 | 427.29 | 349
M7 | A | 9 | 1 | L | 418.04 | 427.66 | 349
M7 | A | 10 | 1 | L | 416.49 | 426.23 | 349
M7 | A | 11 | 1 | L | 415.57 | 425.41 | 349
M7 | A | 12 | 1 | L | 416.41 | 426.22 | 349
M7 | A | 13 | 1 | L | 418.94 | 428.73 | 349
M7 | A | 14 | 1 | L | 419.36 | 429.06 | 349
M7 | A | 15 | 1 | L | 418.31 | 427.94 | 349
M7 | A | 16 | 1 | L | 417.06 | 426.74 | 349
M7 | B | 1 | 1 | L | 415.28 | 425.26 | 349
M7 | B | 2 | 1 | L | 416.83 | 426.88 | 349
M7 | B | 3 | 1 | L | 415.07 | 425 | 349
M7 | B | 4 | 1 | L | 413.47 | 423.21 | 349
M7 | B | 5 | 1 | L | 414.37 | 424.09 | 349
M7 | B | 6 | 1 | L | 415.44 | 425.17 | 349
M7 | B | 7 | 1 | L | 416.08 | 425.77 | 349
M7 | B | 8 | 1 | L | 417.97 | 427.57 | 349
M7 | B | 9 | 1 | L | 418.39 | 427.93 | 349
M7 | B | 10 | 1 | L | 416.87 | 426.45 | 349
M7 | B | 11 | 1 | L | 415.93 | 425.65 | 349
M7 | B | 12 | 1 | L | 416.77 | 426.5 | 349
M7 | B | 13 | 1 | L | 419.31 | 429 | 349
M7 | B | 14 | 1 | L | 419.73 | 429.37 | 349
M7 | B | 15 | 1 | L | 418.7 | 428.27 | 349
M7 | B | 16 | 1 | L | 417.45 | 427.04 | 349

## DNB saturation
Band | HAM | Gain State | Agg Zone | Band Median |  | Band Average |  | Lmax
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | Lsat | Lsat | Lsat | Lsat
LGS | A | 1 | 1 | 0.018849 | 0.019534 | 0.018839 | 0.019523 | 0.02
LGS | A | 1 | 7 | 0.028841 | 0.030154 | 0.028826 | 0.030087 | 0.02
LGS | A | 1 | 11 | 0.041711 | 0.043202 | 0.041687 | 0.043191 | 0.02
LGS | A | 1 | 15 | 0.064513 | 0.066793 | 0.064417 | 0.066721 | 0.02
LGS | A | 1 | 20 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | A | 1 | 26 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | A | 1 | 32 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | A | 1 | 34 | 0.0097873 | 0.010149 | 0.009782 | 0.010131 | 0.02
LGS | B | 1 | 1 | 0.018811 | 0.019522 | 0.018803 | 0.019474 | 0.02
LGS | B | 1 | 7 | 0.028786 | 0.030077 | 0.028772 | 0.030057 | 0.02
LGS | B | 1 | 11 | 0.041636 | 0.043124 | 0.04161 | 0.043133 | 0.02
LGS | B | 1 | 15 | 0.064391 | 0.066714 | 0.064297 | 0.066662 | 0.02
LGS | B | 1 | 20 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | B | 1 | 26 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | B | 1 | 32 | #NUM! | #NUM! | #DIV/0! | #DIV/0! | 0.02
LGS | B | 1 | 34 | 0.0097685 | 0.010134 | 0.0097639 | 0.010117 | 0.02
MGS | A | 2 | 1 | 5.7226e-05 | 5.7995e-05 | 5.7354e-05 | 5.8143e-05 | ~
MGS | A | 2 | 7 | 8.743e-05 | 8.9445e-05 | 8.7675e-05 | 8.9639e-05 | ~
MGS | A | 2 | 11 | 0.00012806 | 0.00013163 | 0.00012844 | 0.00013202 | ~
MGS | A | 2 | 15 | 0.00019849 | 0.00020317 | 0.00019896 | 0.00020427 | ~
MGS | A | 2 | 20 | 1.1348e-07 | 1.1255e-07 | 8.0608e-05 | 8.3158e-05 | ~
MGS | A | 2 | 26 | 0.00046933 | 0.00048537 | 0.00047073 | 0.00048658 | ~
MGS | A | 2 | 32 | 0.00073858 | 0.00076453 | 0.00074023 | 0.000766 | ~
MGS | A | 2 | 34 | 2.9916e-05 | 3.0265e-05 | 2.999e-05 | 3.041e-05 | ~
MGS | B | 2 | 1 | 5.7115e-05 | 5.7914e-05 | 5.7242e-05 | 5.8052e-05 | ~
MGS | B | 2 | 7 | 8.7255e-05 | 8.9283e-05 | 8.7494e-05 | 8.9478e-05 | ~
MGS | B | 2 | 11 | 0.00012779 | 0.00013139 | 0.00012817 | 0.00013175 | ~
MGS | B | 2 | 15 | 0.00019783 | 0.00020317 | 0.00019843 | 0.00020387 | ~
MGS | B | 2 | 20 | 0.00029522 | 0.00030319 | 0.00029597 | 0.00030419 | ~
MGS | B | 2 | 26 | 0.0004687 | 0.00048456 | 0.00046975 | 0.00048577 | ~
MGS | B | 2 | 32 | 0.00073694 | 0.00076293 | 0.00073859 | 0.00076439 | ~
MGS | B | 2 | 34 | 2.9865e-05 | 3.025e-05 | 2.9937e-05 | 3.034e-05 | ~
HGA | A | 3 | 1 | 1.1302e-07 | 1.0048e-07 | 1.1312e-07 | 1.0027e-07 | ~
HGA | A | 3 | 7 | 1.1303e-07 | 1.0553e-07 | 1.1265e-07 | 1.0541e-07 | ~
HGA | A | 3 | 11 | 1.1302e-07 | 1.0845e-07 | 1.1265e-07 | 1.0838e-07 | ~
HGA | A | 3 | 15 | 1.1303e-07 | 1.106e-07 | 1.1285e-07 | 1.1056e-07 | ~
HGA | A | 3 | 20 | 1.1304e-07 | 1.12e-07 | 1.132e-07 | 1.1199e-07 | ~
HGA | A | 3 | 26 | 1.1315e-07 | 1.1297e-07 | 1.1353e-07 | 1.1295e-07 | ~
HGA | A | 3 | 32 | 1.1305e-07 | 1.127e-07 | 1.1269e-07 | 1.1242e-07 | ~
HGA | A | 3 | 34 | 1.135e-07 | 1.1382e-07 | 1.1414e-07 | 1.144e-07 | ~
HGA | B | 3 | 1 | 1.1302e-07 | 1.0048e-07 | 1.1312e-07 | 1.0026e-07 | ~
HGA | B | 3 | 7 | 1.1303e-07 | 1.0554e-07 | 1.1265e-07 | 1.0542e-07 | ~
HGA | B | 3 | 11 | 1.1303e-07 | 1.0846e-07 | 1.1265e-07 | 1.0838e-07 | ~
HGA | B | 3 | 15 | 1.1304e-07 | 1.106e-07 | 1.1286e-07 | 1.1057e-07 | ~
HGA | B | 3 | 20 | 1.1305e-07 | 1.1199e-07 | 1.132e-07 | 1.1199e-07 | ~
HGA | B | 3 | 26 | 1.1315e-07 | 1.1297e-07 | 1.1353e-07 | 1.1295e-07 | ~
HGA | B | 3 | 32 | 1.1306e-07 | 1.127e-07 | 1.1269e-07 | 1.1243e-07 | ~
HGA | B | 3 | 34 | 1.1352e-07 | 1.1381e-07 | 1.1414e-07 | 1.144e-07 | ~
HGB | A | 3 | 1 | 1.131e-07 | 1.0038e-07 | 1.1312e-07 | 1.003e-07 | ~
HGB | A | 3 | 7 | 1.131e-07 | 1.0541e-07 | 1.126e-07 | 1.0537e-07 | ~
HGB | A | 3 | 11 | 1.1308e-07 | 1.083e-07 | 1.126e-07 | 1.0832e-07 | ~
HGB | A | 3 | 15 | 1.1303e-07 | 1.1068e-07 | 1.1283e-07 | 1.1065e-07 | ~
HGB | A | 3 | 20 | 1.1307e-07 | 1.1207e-07 | 1.1318e-07 | 1.1206e-07 | ~
HGB | A | 3 | 26 | 1.131e-07 | 1.1305e-07 | 1.1353e-07 | 1.1306e-07 | ~
HGB | A | 3 | 32 | 1.1304e-07 | 1.127e-07 | 1.1262e-07 | 1.1253e-07 | ~
HGB | A | 3 | 34 | 1.1359e-07 | 1.14e-07 | 1.1418e-07 | 1.1451e-07 | ~
HGB | B | 3 | 1 | 1.1309e-07 | 1.0036e-07 | 1.1312e-07 | 1.003e-07 | ~
HGB | B | 3 | 7 | 1.131e-07 | 1.0542e-07 | 1.126e-07 | 1.0538e-07 | ~
HGB | B | 3 | 11 | 1.1308e-07 | 1.083e-07 | 1.126e-07 | 1.0831e-07 | ~
HGB | B | 3 | 15 | 1.13e-07 | 1.1068e-07 | 1.1283e-07 | 1.1065e-07 | ~
HGB | B | 3 | 20 | 1.1308e-07 | 1.1208e-07 | 1.1318e-07 | 1.1206e-07 | ~
HGB | B | 3 | 26 | 1.131e-07 | 1.1305e-07 | 1.1352e-07 | 1.1306e-07 | ~
HGB | B | 3 | 32 | 1.1303e-07 | 1.1269e-07 | 1.1262e-07 | 1.1253e-07 | ~
HGB | B | 3 | 34 | 1.136e-07 | 1.1399e-07 | 1.1419e-07 | 1.1451e-07 | ~
Band | HAM | Gain State | Agg Zone | Detector | Detector Dependent |  | Lmax |  | Notes:
[5 empty] | Eside A | Eside B |  |  | 1. This table provides VIIRS F2 DNB radiometric saturation.
[5 empty] | Lsat | Lsat |  |  | 2. The results were derived from RC-1 using the SIS-100 source.
LGS | A | 1 | 1 | 1 | 0.018814 | 0.019687 | 0.02 |  | 3. Linear fits of the SIS-100 radiance vs dn were used to determine the radiometric coefficients.
LGS | A | 1 | 1 | 2 | 0.018857 | 0.019688 | 0.02 |  | 4. The testing was conducted at Raytheon El Segundo, Fall/Winter 2013.
LGS | A | 1 | 1 | 3 | 0.019006 | 0.019843 | 0.02 |  | 5. The test was set to diagnostic mode.
LGS | A | 1 | 1 | 4 | 0.018936 | 0.01968 | 0.02 |  | 6. Radiance units are W cm-2 sr-1
LGS | A | 1 | 1 | 5 | 0.018829 | 0.019523 | 0.02 |  | 7. Note only Low Gain Agg modes 1, 7, 11, 15, and 34 saturate at the highest lamp level
LGS | A | 1 | 1 | 6 | 0.01888 | 0.019679 | 0.02 |  | 8. Alhough Agg modes 20, 26, and 32 don't saturate during the test based on their gains they are not expected to presaturate
LGS | A | 1 | 1 | 7 | 0.018904 | 0.019678 | 0.02 |  | 9. Super agg mode 34 not used in EV
LGS | A | 1 | 1 | 8 | 0.018798 | 0.019522 | 0.02 |  | 10. DNB not at operational temperature during ambient testing
LGS | A | 1 | 1 | 9 | 0.018644 | 0.019224 | 0.02
LGS | A | 1 | 1 | 10 | 0.01872 | 0.019225 | 0.02
LGS | A | 1 | 1 | 11 | 0.01877 | 0.019382 | 0.02
LGS | A | 1 | 1 | 12 | 0.018733 | 0.019382 | 0.02
LGS | A | 1 | 1 | 13 | 0.018848 | 0.019382 | 0.02
LGS | A | 1 | 1 | 14 | 0.018877 | 0.019545 | 0.02
LGS | A | 1 | 1 | 15 | 0.018849 | 0.019386 | 0.02
LGS | A | 1 | 1 | 16 | 0.018961 | 0.019546 | 0.02
LGS | A | 1 | 7 | 1 | 0.028933 | 0.0303 | 0.02
LGS | A | 1 | 7 | 2 | 0.029039 | 0.030456 | 0.02
LGS | A | 1 | 7 | 3 | 0.028987 | 0.030299 | 0.02
LGS | A | 1 | 7 | 4 | 0.028888 | 0.030299 | 0.02
LGS | A | 1 | 7 | 5 | 0.028893 | 0.030299 | 0.02
LGS | A | 1 | 7 | 6 | 0.028827 | 0.030299 | 0.02
LGS | A | 1 | 7 | 7 | 0.028932 | 0.030299 | 0.02
LGS | A | 1 | 7 | 8 | 0.028785 | 0.030143 | 0.02
LGS | A | 1 | 7 | 9 | 0.028561 | 0.029698 | 0.02
LGS | A | 1 | 7 | 10 | 0.028632 | 0.029698 | 0.02
LGS | A | 1 | 7 | 11 | 0.028628 | 0.029699 | 0.02
LGS | A | 1 | 7 | 12 | 0.02877 | 0.029855 | 0.02
LGS | A | 1 | 7 | 13 | 0.028748 | 0.029855 | 0.02
LGS | A | 1 | 7 | 14 | 0.028797 | 0.030011 | 0.02
LGS | A | 1 | 7 | 15 | 0.028854 | 0.030011 | 0.02
LGS | A | 1 | 7 | 16 | 0.028948 | 0.030165 | 0.02
LGS | A | 1 | 11 | 1 | 0.041894 | 0.043573 | 0.02
LGS | A | 1 | 11 | 2 | 0.041985 | 0.043729 | 0.02
LGS | A | 1 | 11 | 3 | 0.041963 | 0.043729 | 0.02
LGS | A | 1 | 11 | 4 | 0.041718 | 0.043417 | 0.02
LGS | A | 1 | 11 | 5 | 0.041755 | 0.043418 | 0.02
LGS | A | 1 | 11 | 6 | 0.041866 | 0.043576 | 0.02
LGS | A | 1 | 11 | 7 | 0.041725 | 0.043419 | 0.02
LGS | A | 1 | 11 | 8 | 0.041704 | 0.043263 | 0.02
LGS | A | 1 | 11 | 9 | 0.04142 | 0.042671 | 0.02
LGS | A | 1 | 11 | 10 | 0.041241 | 0.042515 | 0.02
LGS | A | 1 | 11 | 11 | 0.041537 | 0.042828 | 0.02
LGS | A | 1 | 11 | 12 | 0.041592 | 0.042984 | 0.02
LGS | A | 1 | 11 | 13 | 0.041486 | 0.042828 | 0.02
LGS | A | 1 | 11 | 14 | 0.041757 | 0.04314 | 0.02
LGS | A | 1 | 11 | 15 | 0.041689 | 0.042984 | 0.02
LGS | A | 1 | 11 | 16 | 0.041665 | 0.042984 | 0.02
LGS | A | 1 | 15 | 1 | 0.064816 | 0.067457 | 0.02
LGS | A | 1 | 15 | 2 | 0.064863 | 0.067461 | 0.02
LGS | A | 1 | 15 | 3 | 0.064838 | 0.067463 | 0.02
LGS | A | 1 | 15 | 4 | 0.064562 | 0.067201 | 0.02
LGS | A | 1 | 15 | 5 | 0.064523 | 0.067158 | 0.02
LGS | A | 1 | 15 | 6 | 0.064508 | 0.067158 | 0.02
LGS | A | 1 | 15 | 7 | 0.064539 | 0.067003 | 0.02
LGS | A | 1 | 15 | 8 | 0.064558 | 0.067004 | 0.02
LGS | A | 1 | 15 | 9 | 0.064155 | 0.066113 | 0.02
LGS | A | 1 | 15 | 10 | 0.063772 | 0.0658 | 0.02
LGS | A | 1 | 15 | 11 | 0.063952 | 0.065957 | 0.02
LGS | A | 1 | 15 | 12 | 0.064122 | 0.066116 | 0.02
LGS | A | 1 | 15 | 13 | 0.064191 | 0.066207 | 0.02
LGS | A | 1 | 15 | 14 | 0.064413 | 0.066426 | 0.02
LGS | A | 1 | 15 | 15 | 0.064519 | 0.066583 | 0.02
LGS | A | 1 | 15 | 16 | 0.064348 | 0.066427 | 0.02
LGS | A | 1 | 20 | 1 |  |  | 0.02
LGS | A | 1 | 20 | 2 |  |  | 0.02
LGS | A | 1 | 20 | 3 |  |  | 0.02
LGS | A | 1 | 20 | 4 |  |  | 0.02
LGS | A | 1 | 20 | 5 |  |  | 0.02
LGS | A | 1 | 20 | 6 |  |  | 0.02
LGS | A | 1 | 20 | 7 |  |  | 0.02
LGS | A | 1 | 20 | 8 |  |  | 0.02
LGS | A | 1 | 20 | 9 |  |  | 0.02
LGS | A | 1 | 20 | 10 |  |  | 0.02
LGS | A | 1 | 20 | 11 |  |  | 0.02
LGS | A | 1 | 20 | 12 |  |  | 0.02
LGS | A | 1 | 20 | 13 |  |  | 0.02
LGS | A | 1 | 20 | 14 |  |  | 0.02
LGS | A | 1 | 20 | 15 |  |  | 0.02
LGS | A | 1 | 20 | 16 |  |  | 0.02
LGS | A | 1 | 26 | 1 |  |  | 0.02
LGS | A | 1 | 26 | 2 |  |  | 0.02
LGS | A | 1 | 26 | 3 |  |  | 0.02
LGS | A | 1 | 26 | 4 |  |  | 0.02
LGS | A | 1 | 26 | 5 |  |  | 0.02
LGS | A | 1 | 26 | 6 |  |  | 0.02
LGS | A | 1 | 26 | 7 |  |  | 0.02
LGS | A | 1 | 26 | 8 |  |  | 0.02
LGS | A | 1 | 26 | 9 |  |  | 0.02
LGS | A | 1 | 26 | 10 |  |  | 0.02
LGS | A | 1 | 26 | 11 |  |  | 0.02
LGS | A | 1 | 26 | 12 |  |  | 0.02
LGS | A | 1 | 26 | 13 |  |  | 0.02
LGS | A | 1 | 26 | 14 |  |  | 0.02
LGS | A | 1 | 26 | 15 |  |  | 0.02
LGS | A | 1 | 26 | 16 |  |  | 0.02
LGS | A | 1 | 32 | 1 |  |  | 0.02
LGS | A | 1 | 32 | 2 |  |  | 0.02
LGS | A | 1 | 32 | 3 |  |  | 0.02
LGS | A | 1 | 32 | 4 |  |  | 0.02
LGS | A | 1 | 32 | 5 |  |  | 0.02
LGS | A | 1 | 32 | 6 |  |  | 0.02
LGS | A | 1 | 32 | 7 |  |  | 0.02
LGS | A | 1 | 32 | 8 |  |  | 0.02
LGS | A | 1 | 32 | 9 |  |  | 0.02
LGS | A | 1 | 32 | 10 |  |  | 0.02
LGS | A | 1 | 32 | 11 |  |  | 0.02
LGS | A | 1 | 32 | 12 |  |  | 0.02
LGS | A | 1 | 32 | 13 |  |  | 0.02
LGS | A | 1 | 32 | 14 |  |  | 0.02
LGS | A | 1 | 32 | 15 |  |  | 0.02
LGS | A | 1 | 32 | 16 |  |  | 0.02
LGS | A | 1 | 34 | 1 | 0.009769 | 0.010182 | 0.02
LGS | A | 1 | 34 | 2 | 0.0097896 | 0.010182 | 0.02
LGS | A | 1 | 34 | 3 | 0.0098656 | 0.010259 | 0.02
LGS | A | 1 | 34 | 4 | 0.0098322 | 0.010228 | 0.02
LGS | A | 1 | 34 | 5 | 0.0097767 | 0.010165 | 0.02
LGS | A | 1 | 34 | 6 | 0.0098035 | 0.010196 | 0.02
LGS | A | 1 | 34 | 7 | 0.0098159 | 0.010212 | 0.02
LGS | A | 1 | 34 | 8 | 0.0097599 | 0.010133 | 0.02
LGS | A | 1 | 34 | 9 | 0.0096812 | 0.0099688 | 0.02
LGS | A | 1 | 34 | 10 | 0.0097216 | 0.010016 | 0.02
LGS | A | 1 | 34 | 11 | 0.009747 | 0.010048 | 0.02
LGS | A | 1 | 34 | 12 | 0.0097269 | 0.010033 | 0.02
LGS | A | 1 | 34 | 13 | 0.0097874 | 0.010095 | 0.02
LGS | A | 1 | 34 | 14 | 0.0098007 | 0.01011 | 0.02
LGS | A | 1 | 34 | 15 | 0.0097871 | 0.010092 | 0.02
LGS | A | 1 | 34 | 16 | 0.0098472 | 0.010173 | 0.02
LGS | B | 1 | 1 | 1 | 0.018779 | 0.019531 | 0.02
LGS | B | 1 | 1 | 2 | 0.018824 | 0.019531 | 0.02
LGS | B | 1 | 1 | 3 | 0.018971 | 0.019687 | 0.02
LGS | B | 1 | 1 | 4 | 0.018901 | 0.01968 | 0.02
LGS | B | 1 | 1 | 5 | 0.018791 | 0.019523 | 0.02
LGS | B | 1 | 1 | 6 | 0.018841 | 0.019679 | 0.02
LGS | B | 1 | 1 | 7 | 0.018868 | 0.019678 | 0.02
LGS | B | 1 | 1 | 8 | 0.018764 | 0.019522 | 0.02
LGS | B | 1 | 1 | 9 | 0.018608 | 0.019224 | 0.02
LGS | B | 1 | 1 | 10 | 0.018686 | 0.019225 | 0.02
LGS | B | 1 | 1 | 11 | 0.018735 | 0.019381 | 0.02
LGS | B | 1 | 1 | 12 | 0.018697 | 0.019225 | 0.02
LGS | B | 1 | 1 | 13 | 0.01881 | 0.019382 | 0.02
LGS | B | 1 | 1 | 14 | 0.018841 | 0.019389 | 0.02
LGS | B | 1 | 1 | 15 | 0.018812 | 0.019386 | 0.02
LGS | B | 1 | 1 | 16 | 0.018924 | 0.019545 | 0.02
LGS | B | 1 | 7 | 1 | 0.028882 | 0.030301 | 0.02
LGS | B | 1 | 7 | 2 | 0.028985 | 0.030456 | 0.02
LGS | B | 1 | 7 | 3 | 0.028933 | 0.0303 | 0.02
LGS | B | 1 | 7 | 4 | 0.028829 | 0.030299 | 0.02
LGS | B | 1 | 7 | 5 | 0.028838 | 0.030299 | 0.02
LGS | B | 1 | 7 | 6 | 0.028771 | 0.030143 | 0.02
LGS | B | 1 | 7 | 7 | 0.02888 | 0.030299 | 0.02
LGS | B | 1 | 7 | 8 | 0.028732 | 0.030143 | 0.02
LGS | B | 1 | 7 | 9 | 0.028506 | 0.029698 | 0.02
LGS | B | 1 | 7 | 10 | 0.028575 | 0.029698 | 0.02
LGS | B | 1 | 7 | 11 | 0.028577 | 0.029698 | 0.02
LGS | B | 1 | 7 | 12 | 0.028715 | 0.029855 | 0.02
LGS | B | 1 | 7 | 13 | 0.028692 | 0.029855 | 0.02
LGS | B | 1 | 7 | 14 | 0.028745 | 0.029855 | 0.02
LGS | B | 1 | 7 | 15 | 0.028802 | 0.030011 | 0.02
LGS | B | 1 | 7 | 16 | 0.028893 | 0.030008 | 0.02
LGS | B | 1 | 11 | 1 | 0.04182 | 0.043573 | 0.02
LGS | B | 1 | 11 | 2 | 0.041904 | 0.043572 | 0.02
LGS | B | 1 | 11 | 3 | 0.041883 | 0.043573 | 0.02
LGS | B | 1 | 11 | 4 | 0.041644 | 0.043417 | 0.02
LGS | B | 1 | 11 | 5 | 0.041682 | 0.043418 | 0.02
LGS | B | 1 | 11 | 6 | 0.041787 | 0.043575 | 0.02
LGS | B | 1 | 11 | 7 | 0.04165 | 0.043263 | 0.02
LGS | B | 1 | 11 | 8 | 0.041628 | 0.043263 | 0.02
LGS | B | 1 | 11 | 9 | 0.041342 | 0.042671 | 0.02
LGS | B | 1 | 11 | 10 | 0.041167 | 0.042514 | 0.02
LGS | B | 1 | 11 | 11 | 0.041458 | 0.042828 | 0.02
LGS | B | 1 | 11 | 12 | 0.041512 | 0.042828 | 0.02
LGS | B | 1 | 11 | 13 | 0.041405 | 0.042671 | 0.02
LGS | B | 1 | 11 | 14 | 0.041679 | 0.042984 | 0.02
LGS | B | 1 | 11 | 15 | 0.041609 | 0.042985 | 0.02
LGS | B | 1 | 11 | 16 | 0.041588 | 0.042984 | 0.02
LGS | B | 1 | 15 | 1 | 0.064692 | 0.067302 | 0.02
LGS | B | 1 | 15 | 2 | 0.064746 | 0.06746 | 0.02
LGS | B | 1 | 15 | 3 | 0.064708 | 0.067462 | 0.02
LGS | B | 1 | 15 | 4 | 0.06445 | 0.067045 | 0.02
LGS | B | 1 | 15 | 5 | 0.064404 | 0.067158 | 0.02
LGS | B | 1 | 15 | 6 | 0.064388 | 0.067002 | 0.02
LGS | B | 1 | 15 | 7 | 0.06442 | 0.067003 | 0.02
LGS | B | 1 | 15 | 8 | 0.064442 | 0.067004 | 0.02
LGS | B | 1 | 15 | 9 | 0.064033 | 0.066113 | 0.02
LGS | B | 1 | 15 | 10 | 0.063651 | 0.065644 | 0.02
LGS | B | 1 | 15 | 11 | 0.063837 | 0.065801 | 0.02
LGS | B | 1 | 15 | 12 | 0.063998 | 0.066115 | 0.02
LGS | B | 1 | 15 | 13 | 0.064069 | 0.066206 | 0.02
LGS | B | 1 | 15 | 14 | 0.064292 | 0.066425 | 0.02
LGS | B | 1 | 15 | 15 | 0.064395 | 0.066426 | 0.02
LGS | B | 1 | 15 | 16 | 0.064228 | 0.066426 | 0.02
LGS | B | 1 | 20 | 1 |  |  | 0.02
LGS | B | 1 | 20 | 2 |  |  | 0.02
LGS | B | 1 | 20 | 3 |  |  | 0.02
LGS | B | 1 | 20 | 4 |  |  | 0.02
LGS | B | 1 | 20 | 5 |  |  | 0.02
LGS | B | 1 | 20 | 6 |  |  | 0.02
LGS | B | 1 | 20 | 7 |  |  | 0.02
LGS | B | 1 | 20 | 8 |  |  | 0.02
LGS | B | 1 | 20 | 9 |  |  | 0.02
LGS | B | 1 | 20 | 10 |  |  | 0.02
LGS | B | 1 | 20 | 11 |  |  | 0.02
LGS | B | 1 | 20 | 12 |  |  | 0.02
LGS | B | 1 | 20 | 13 |  |  | 0.02
LGS | B | 1 | 20 | 14 |  |  | 0.02
LGS | B | 1 | 20 | 15 |  |  | 0.02
LGS | B | 1 | 20 | 16 |  |  | 0.02
LGS | B | 1 | 26 | 1 |  |  | 0.02
LGS | B | 1 | 26 | 2 |  |  | 0.02
LGS | B | 1 | 26 | 3 |  |  | 0.02
LGS | B | 1 | 26 | 4 |  |  | 0.02
LGS | B | 1 | 26 | 5 |  |  | 0.02
LGS | B | 1 | 26 | 6 |  |  | 0.02
LGS | B | 1 | 26 | 7 |  |  | 0.02
LGS | B | 1 | 26 | 8 |  |  | 0.02
LGS | B | 1 | 26 | 9 |  |  | 0.02
LGS | B | 1 | 26 | 10 |  |  | 0.02
LGS | B | 1 | 26 | 11 |  |  | 0.02
LGS | B | 1 | 26 | 12 |  |  | 0.02
LGS | B | 1 | 26 | 13 |  |  | 0.02
LGS | B | 1 | 26 | 14 |  |  | 0.02
LGS | B | 1 | 26 | 15 |  |  | 0.02
LGS | B | 1 | 26 | 16 |  |  | 0.02
LGS | B | 1 | 32 | 1 |  |  | 0.02
LGS | B | 1 | 32 | 2 |  |  | 0.02
LGS | B | 1 | 32 | 3 |  |  | 0.02
LGS | B | 1 | 32 | 4 |  |  | 0.02
LGS | B | 1 | 32 | 5 |  |  | 0.02
LGS | B | 1 | 32 | 6 |  |  | 0.02
LGS | B | 1 | 32 | 7 |  |  | 0.02
LGS | B | 1 | 32 | 8 |  |  | 0.02
LGS | B | 1 | 32 | 9 |  |  | 0.02
LGS | B | 1 | 32 | 10 |  |  | 0.02
LGS | B | 1 | 32 | 11 |  |  | 0.02
LGS | B | 1 | 32 | 12 |  |  | 0.02
LGS | B | 1 | 32 | 13 |  |  | 0.02
LGS | B | 1 | 32 | 14 |  |  | 0.02
LGS | B | 1 | 32 | 15 |  |  | 0.02
LGS | B | 1 | 32 | 16 |  |  | 0.02
LGS | B | 1 | 34 | 1 | 0.009752 | 0.010166 | 0.02
LGS | B | 1 | 34 | 2 | 0.0097725 | 0.010166 | 0.02
LGS | B | 1 | 34 | 3 | 0.0098473 | 0.010244 | 0.02
LGS | B | 1 | 34 | 4 | 0.0098148 | 0.010212 | 0.02
LGS | B | 1 | 34 | 5 | 0.0097579 | 0.01015 | 0.02
LGS | B | 1 | 34 | 6 | 0.0097846 | 0.010196 | 0.02
LGS | B | 1 | 34 | 7 | 0.0097996 | 0.010196 | 0.02
LGS | B | 1 | 34 | 8 | 0.0097419 | 0.010117 | 0.02
LGS | B | 1 | 34 | 9 | 0.0096621 | 0.0099689 | 0.02
LGS | B | 1 | 34 | 10 | 0.0097032 | 0.010001 | 0.02
LGS | B | 1 | 34 | 11 | 0.0097288 | 0.010032 | 0.02
LGS | B | 1 | 34 | 12 | 0.00971 | 0.010017 | 0.02
LGS | B | 1 | 34 | 13 | 0.0097682 | 0.01008 | 0.02
LGS | B | 1 | 34 | 14 | 0.0097838 | 0.010095 | 0.02
LGS | B | 1 | 34 | 15 | 0.0097687 | 0.010076 | 0.02
LGS | B | 1 | 34 | 16 | 0.0098276 | 0.010157 | 0.02
MGS | A | 2 | 1 | 1 | 5.8873e-05 | 5.9586e-05 | ~
MGS | A | 2 | 1 | 2 | 5.9342e-05 | 5.991e-05 | ~
MGS | A | 2 | 1 | 3 | 5.9397e-05 | 6.0073e-05 | ~
MGS | A | 2 | 1 | 4 | 5.9426e-05 | 6.0073e-05 | ~
MGS | A | 2 | 1 | 5 | 5.9422e-05 | 6.0079e-05 | ~
MGS | A | 2 | 1 | 6 | 5.9301e-05 | 5.9918e-05 | ~
MGS | A | 2 | 1 | 7 | 5.8833e-05 | 5.9432e-05 | ~
MGS | A | 2 | 1 | 8 | 5.8774e-05 | 5.9432e-05 | ~
MGS | A | 2 | 1 | 9 | 5.5401e-05 | 5.6393e-05 | ~
MGS | A | 2 | 1 | 10 | 5.5326e-05 | 5.6232e-05 | ~
MGS | A | 2 | 1 | 11 | 5.5575e-05 | 5.6558e-05 | ~
MGS | A | 2 | 1 | 12 | 5.5671e-05 | 5.6558e-05 | ~
MGS | A | 2 | 1 | 13 | 5.5677e-05 | 5.6554e-05 | ~
MGS | A | 2 | 1 | 14 | 5.5615e-05 | 5.6552e-05 | ~
MGS | A | 2 | 1 | 15 | 5.5564e-05 | 5.6551e-05 | ~
MGS | A | 2 | 1 | 16 | 5.5477e-05 | 5.6388e-05 | ~
MGS | A | 2 | 7 | 1 | 9.0649e-05 | 9.2493e-05 | ~
MGS | A | 2 | 7 | 2 | 9.0673e-05 | 9.2492e-05 | ~
MGS | A | 2 | 7 | 3 | 9.0844e-05 | 9.2653e-05 | ~
MGS | A | 2 | 7 | 4 | 9.0696e-05 | 9.2482e-05 | ~
MGS | A | 2 | 7 | 5 | 9.0722e-05 | 9.2482e-05 | ~
MGS | A | 2 | 7 | 6 | 9.0359e-05 | 9.2161e-05 | ~
MGS | A | 2 | 7 | 7 | 8.986e-05 | 9.1514e-05 | ~
MGS | A | 2 | 7 | 8 | 8.9732e-05 | 9.1514e-05 | ~
MGS | A | 2 | 7 | 9 | 8.4619e-05 | 8.6736e-05 | ~
MGS | A | 2 | 7 | 10 | 8.463e-05 | 8.6736e-05 | ~
MGS | A | 2 | 7 | 11 | 8.4798e-05 | 8.6897e-05 | ~
MGS | A | 2 | 7 | 12 | 8.5064e-05 | 8.722e-05 | ~
MGS | A | 2 | 7 | 13 | 8.5071e-05 | 8.7203e-05 | ~
MGS | A | 2 | 7 | 14 | 8.5128e-05 | 8.7376e-05 | ~
MGS | A | 2 | 7 | 15 | 8.4961e-05 | 8.7053e-05 | ~
MGS | A | 2 | 7 | 16 | 8.4998e-05 | 8.7214e-05 | ~
MGS | A | 2 | 11 | 1 | 0.00013315 | 0.00013652 | ~
MGS | A | 2 | 11 | 2 | 0.00013296 | 0.0001362 | ~
MGS | A | 2 | 11 | 3 | 0.00013316 | 0.00013652 | ~
MGS | A | 2 | 11 | 4 | 0.00013318 | 0.00013652 | ~
MGS | A | 2 | 11 | 5 | 0.00013271 | 0.00013604 | ~
MGS | A | 2 | 11 | 6 | 0.00013219 | 0.00013539 | ~
MGS | A | 2 | 11 | 7 | 0.00013177 | 0.00013507 | ~
MGS | A | 2 | 11 | 8 | 0.00013144 | 0.00013474 | ~
MGS | A | 2 | 11 | 9 | 0.00012391 | 0.00012771 | ~
MGS | A | 2 | 11 | 10 | 0.00012405 | 0.00012787 | ~
MGS | A | 2 | 11 | 11 | 0.00012403 | 0.00012787 | ~
MGS | A | 2 | 11 | 12 | 0.00012428 | 0.0001282 | ~
MGS | A | 2 | 11 | 13 | 0.00012466 | 0.00012852 | ~
MGS | A | 2 | 11 | 14 | 0.00012468 | 0.00012852 | ~
MGS | A | 2 | 11 | 15 | 0.00012436 | 0.0001282 | ~
MGS | A | 2 | 11 | 16 | 0.00012455 | 0.00012836 | ~
MGS | A | 2 | 15 | 1 | 0.00020636 | 0.00021119 | ~
MGS | A | 2 | 15 | 2 | 0.00020604 | 0.00021119 | ~
MGS | A | 2 | 15 | 3 | 0.00020573 | 0.00021118 | ~
MGS | A | 2 | 15 | 4 | 0.00020598 | 0.00021118 | ~
MGS | A | 2 | 15 | 5 | 0.00020554 | 0.00021118 | ~
MGS | A | 2 | 15 | 6 | 0.00020466 | 0.00020956 | ~
MGS | A | 2 | 15 | 7 | 0.0002038 | 0.00020795 | ~
MGS | A | 2 | 15 | 8 | 0.00020345 | 0.00020795 | ~
MGS | A | 2 | 15 | 9 | 0.0001918 | 0.00019836 | ~
MGS | A | 2 | 15 | 10 | 0.00019199 | 0.00019837 | ~
MGS | A | 2 | 15 | 11 | 0.00019223 | 0.00019837 | ~
MGS | A | 2 | 15 | 12 | 0.0001926 | 0.00019837 | ~
MGS | A | 2 | 15 | 13 | 0.00019351 | 0.00019838 | ~
MGS | A | 2 | 15 | 14 | 0.00019266 | 0.00019838 | ~
MGS | A | 2 | 15 | 15 | 0.00019352 | 0.00019839 | ~
MGS | A | 2 | 15 | 16 | 0.00019353 | 0.00019839 | ~
MGS | A | 2 | 20 | 1 | 0.00030805 | 0.00031599 | ~
MGS | A | 2 | 20 | 2 | 0.00030753 | 0.00031439 | ~
MGS | A | 2 | 20 | 3 | 0.00030737 | 0.00031439 | ~
MGS | A | 2 | 20 | 4 | 0.00030627 | 0.0003144 | ~
MGS | A | 2 | 20 | 5 | 0.00030627 | 0.00031279 | ~
MGS | A | 2 | 20 | 6 | 0.00030446 | 0.00031117 | ~
MGS | A | 2 | 20 | 7 | 0.00030348 | 0.00031117 | ~
MGS | A | 2 | 20 | 8 | 0.00030415 | 0.00031117 | ~
MGS | A | 2 | 20 | 9 | 0.00028681 | 0.0002952 | ~
MGS | A | 2 | 20 | 10 | 0.00028604 | 0.0002952 | ~
MGS | A | 2 | 20 | 11 | 0.00028629 | 0.0002952 | ~
MGS | A | 2 | 20 | 12 | 0.00028705 | 0.0002952 | ~
MGS | A | 2 | 20 | 13 | 0.0002876 | 0.00029682 | ~
MGS | A | 2 | 20 | 14 | 0.00028788 | 0.00029681 | ~
MGS | A | 2 | 20 | 15 | 0.00028807 | 0.00029682 | ~
MGS | A | 2 | 20 | 16 | 0.00028825 | 0.0002968 | ~
MGS | A | 2 | 26 | 1 | 0.0004887 | 0.00050466 | ~
MGS | A | 2 | 26 | 2 | 0.00048809 | 0.00050296 | ~
MGS | A | 2 | 26 | 3 | 0.00048761 | 0.00050298 | ~
MGS | A | 2 | 26 | 4 | 0.00048651 | 0.00050137 | ~
MGS | A | 2 | 26 | 5 | 0.00048526 | 0.00050137 | ~
MGS | A | 2 | 26 | 6 | 0.00048353 | 0.00049815 | ~
MGS | A | 2 | 26 | 7 | 0.00048165 | 0.00049653 | ~
MGS | A | 2 | 26 | 8 | 0.00048266 | 0.00049813 | ~
MGS | A | 2 | 26 | 9 | 0.00045573 | 0.00047259 | ~
MGS | A | 2 | 26 | 10 | 0.00045484 | 0.00047098 | ~
MGS | A | 2 | 26 | 11 | 0.00045509 | 0.00047098 | ~
MGS | A | 2 | 26 | 12 | 0.00045592 | 0.0004726 | ~
MGS | A | 2 | 26 | 13 | 0.00045623 | 0.0004726 | ~
MGS | A | 2 | 26 | 14 | 0.00045681 | 0.0004726 | ~
MGS | A | 2 | 26 | 15 | 0.00045701 | 0.00047259 | ~
MGS | A | 2 | 26 | 16 | 0.00045605 | 0.00047422 | ~
MGS | A | 2 | 32 | 1 | 0.00077008 | 0.00079402 | ~
MGS | A | 2 | 32 | 2 | 0.00076888 | 0.00079272 | ~
MGS | A | 2 | 32 | 3 | 0.00077047 | 0.00079569 | ~
MGS | A | 2 | 32 | 4 | 0.00076691 | 0.00079135 | ~
MGS | A | 2 | 32 | 5 | 0.00076189 | 0.00078548 | ~
MGS | A | 2 | 32 | 6 | 0.00075934 | 0.00078386 | ~
MGS | A | 2 | 32 | 7 | 0.00075754 | 0.0007819 | ~
MGS | A | 2 | 32 | 8 | 0.00076108 | 0.00078514 | ~
MGS | A | 2 | 32 | 9 | 0.00071607 | 0.00074276 | ~
MGS | A | 2 | 32 | 10 | 0.00071395 | 0.00074181 | ~
MGS | A | 2 | 32 | 11 | 0.00071442 | 0.00074221 | ~
MGS | A | 2 | 32 | 12 | 0.00071282 | 0.0007402 | ~
MGS | A | 2 | 32 | 13 | 0.0007167 | 0.00074378 | ~
MGS | A | 2 | 32 | 14 | 0.00071963 | 0.00074717 | ~
MGS | A | 2 | 32 | 15 | 0.0007168 | 0.00074339 | ~
MGS | A | 2 | 32 | 16 | 0.00071705 | 0.00074452 | ~
MGS | A | 2 | 34 | 1 | 3.0811e-05 | 3.1152e-05 | ~
MGS | A | 2 | 34 | 2 | 3.1084e-05 | 3.1489e-05 | ~
MGS | A | 2 | 34 | 3 | 3.1086e-05 | 3.1479e-05 | ~
MGS | A | 2 | 34 | 4 | 3.1115e-05 | 3.1496e-05 | ~
MGS | A | 2 | 34 | 5 | 3.1092e-05 | 3.1473e-05 | ~
MGS | A | 2 | 34 | 6 | 3.1014e-05 | 3.1294e-05 | ~
MGS | A | 2 | 34 | 7 | 3.074e-05 | 3.1119e-05 | ~
MGS | A | 2 | 34 | 8 | 3.0744e-05 | 3.0989e-05 | ~
MGS | A | 2 | 34 | 9 | 2.895e-05 | 2.9523e-05 | ~
MGS | A | 2 | 34 | 10 | 2.8913e-05 | 2.9362e-05 | ~
MGS | A | 2 | 34 | 11 | 2.9046e-05 | 2.9541e-05 | ~
MGS | A | 2 | 34 | 12 | 2.9092e-05 | 2.9525e-05 | ~
MGS | A | 2 | 34 | 13 | 2.9091e-05 | 2.9531e-05 | ~
MGS | A | 2 | 34 | 14 | 2.9053e-05 | 2.9536e-05 | ~
MGS | A | 2 | 34 | 15 | 2.9026e-05 | 2.9528e-05 | ~
MGS | A | 2 | 34 | 16 | 2.8979e-05 | 2.9529e-05 | ~
MGS | B | 2 | 1 | 1 | 5.8769e-05 | 5.9424e-05 | ~
MGS | B | 2 | 1 | 2 | 5.9228e-05 | 5.9909e-05 | ~
MGS | B | 2 | 1 | 3 | 5.9271e-05 | 5.9911e-05 | ~
MGS | B | 2 | 1 | 4 | 5.9307e-05 | 5.9913e-05 | ~
MGS | B | 2 | 1 | 5 | 5.9311e-05 | 5.9918e-05 | ~
MGS | B | 2 | 1 | 6 | 5.9189e-05 | 5.9918e-05 | ~
MGS | B | 2 | 1 | 7 | 5.8716e-05 | 5.9272e-05 | ~
MGS | B | 2 | 1 | 8 | 5.8664e-05 | 5.9271e-05 | ~
MGS | B | 2 | 1 | 9 | 5.5293e-05 | 5.6233e-05 | ~
MGS | B | 2 | 1 | 10 | 5.5223e-05 | 5.6232e-05 | ~
MGS | B | 2 | 1 | 11 | 5.547e-05 | 5.6396e-05 | ~
MGS | B | 2 | 1 | 12 | 5.556e-05 | 5.6558e-05 | ~
MGS | B | 2 | 1 | 13 | 5.5565e-05 | 5.6554e-05 | ~
MGS | B | 2 | 1 | 14 | 5.5503e-05 | 5.6553e-05 | ~
MGS | B | 2 | 1 | 15 | 5.5451e-05 | 5.639e-05 | ~
MGS | B | 2 | 1 | 16 | 5.5359e-05 | 5.6388e-05 | ~
MGS | B | 2 | 7 | 1 | 9.0466e-05 | 9.233e-05 | ~
MGS | B | 2 | 7 | 2 | 9.0476e-05 | 9.2331e-05 | ~
MGS | B | 2 | 7 | 3 | 9.066e-05 | 9.2494e-05 | ~
MGS | B | 2 | 7 | 4 | 9.0515e-05 | 9.2319e-05 | ~
MGS | B | 2 | 7 | 5 | 9.0536e-05 | 9.2321e-05 | ~
MGS | B | 2 | 7 | 6 | 9.0176e-05 | 9.1999e-05 | ~
MGS | B | 2 | 7 | 7 | 8.967e-05 | 9.1353e-05 | ~
MGS | B | 2 | 7 | 8 | 8.9562e-05 | 9.1351e-05 | ~
MGS | B | 2 | 7 | 9 | 8.4447e-05 | 8.6575e-05 | ~
MGS | B | 2 | 7 | 10 | 8.445e-05 | 8.6572e-05 | ~
MGS | B | 2 | 7 | 11 | 8.4626e-05 | 8.6737e-05 | ~
MGS | B | 2 | 7 | 12 | 8.4878e-05 | 8.7059e-05 | ~
MGS | B | 2 | 7 | 13 | 8.489e-05 | 8.7043e-05 | ~
MGS | B | 2 | 7 | 14 | 8.4949e-05 | 8.7215e-05 | ~
MGS | B | 2 | 7 | 15 | 8.4788e-05 | 8.6892e-05 | ~
MGS | B | 2 | 7 | 16 | 8.482e-05 | 8.7052e-05 | ~
MGS | B | 2 | 11 | 1 | 0.00013288 | 0.0001362 | ~
MGS | B | 2 | 11 | 2 | 0.00013268 | 0.00013604 | ~
MGS | B | 2 | 11 | 3 | 0.00013288 | 0.0001362 | ~
MGS | B | 2 | 11 | 4 | 0.00013289 | 0.0001362 | ~
MGS | B | 2 | 11 | 5 | 0.00013243 | 0.00013571 | ~
MGS | B | 2 | 11 | 6 | 0.00013191 | 0.00013523 | ~
MGS | B | 2 | 11 | 7 | 0.00013149 | 0.00013474 | ~
MGS | B | 2 | 11 | 8 | 0.00013117 | 0.00013442 | ~
MGS | B | 2 | 11 | 9 | 0.00012365 | 0.00012755 | ~
MGS | B | 2 | 11 | 10 | 0.00012379 | 0.00012755 | ~
MGS | B | 2 | 11 | 11 | 0.00012376 | 0.00012755 | ~
MGS | B | 2 | 11 | 12 | 0.00012403 | 0.00012788 | ~
MGS | B | 2 | 11 | 13 | 0.00012439 | 0.0001282 | ~
MGS | B | 2 | 11 | 14 | 0.0001244 | 0.00012836 | ~
MGS | B | 2 | 11 | 15 | 0.00012409 | 0.00012804 | ~
MGS | B | 2 | 11 | 16 | 0.00012428 | 0.0001282 | ~
MGS | B | 2 | 15 | 1 | 0.00020594 | 0.00021119 | ~
MGS | B | 2 | 15 | 2 | 0.00020559 | 0.00021119 | ~
MGS | B | 2 | 15 | 3 | 0.00020533 | 0.00020957 | ~
MGS | B | 2 | 15 | 4 | 0.00020554 | 0.00021118 | ~
MGS | B | 2 | 15 | 5 | 0.0002051 | 0.00020956 | ~
MGS | B | 2 | 15 | 6 | 0.00020425 | 0.00020956 | ~
MGS | B | 2 | 15 | 7 | 0.00020337 | 0.00020795 | ~
MGS | B | 2 | 15 | 8 | 0.00020301 | 0.00020794 | ~
MGS | B | 2 | 15 | 9 | 0.0001914 | 0.00019675 | ~
MGS | B | 2 | 15 | 10 | 0.00019158 | 0.00019675 | ~
MGS | B | 2 | 15 | 11 | 0.00019183 | 0.00019837 | ~
MGS | B | 2 | 15 | 12 | 0.00019218 | 0.00019838 | ~
MGS | B | 2 | 15 | 13 | 0.00019243 | 0.00019838 | ~
MGS | B | 2 | 15 | 14 | 0.00019224 | 0.00019838 | ~
MGS | B | 2 | 15 | 15 | 0.00019246 | 0.00019839 | ~
MGS | B | 2 | 15 | 16 | 0.00019264 | 0.00019839 | ~
MGS | B | 2 | 20 | 1 | 0.00030738 | 0.00031438 | ~
MGS | B | 2 | 20 | 2 | 0.00030626 | 0.00031439 | ~
MGS | B | 2 | 20 | 3 | 0.00030626 | 0.00031439 | ~
MGS | B | 2 | 20 | 4 | 0.00030627 | 0.0003144 | ~
MGS | B | 2 | 20 | 5 | 0.00030627 | 0.00031279 | ~
MGS | B | 2 | 20 | 6 | 0.00030382 | 0.00031117 | ~
MGS | B | 2 | 20 | 7 | 0.00030282 | 0.00030955 | ~
MGS | B | 2 | 20 | 8 | 0.00030347 | 0.00031117 | ~
MGS | B | 2 | 20 | 9 | 0.00028616 | 0.0002952 | ~
MGS | B | 2 | 20 | 10 | 0.00028544 | 0.00029359 | ~
MGS | B | 2 | 20 | 11 | 0.00028572 | 0.0002952 | ~
MGS | B | 2 | 20 | 12 | 0.00028643 | 0.0002952 | ~
MGS | B | 2 | 20 | 13 | 0.00028697 | 0.0002952 | ~
MGS | B | 2 | 20 | 14 | 0.00028723 | 0.00029681 | ~
MGS | B | 2 | 20 | 15 | 0.00028745 | 0.00029682 | ~
MGS | B | 2 | 20 | 16 | 0.00028762 | 0.0002968 | ~
MGS | B | 2 | 26 | 1 | 0.00048764 | 0.00050304 | ~
MGS | B | 2 | 26 | 2 | 0.00048701 | 0.00050297 | ~
MGS | B | 2 | 26 | 3 | 0.00048653 | 0.00050298 | ~
MGS | B | 2 | 26 | 4 | 0.00048547 | 0.00050137 | ~
MGS | B | 2 | 26 | 5 | 0.00048354 | 0.00049976 | ~
MGS | B | 2 | 26 | 6 | 0.00048199 | 0.00049814 | ~
MGS | B | 2 | 26 | 7 | 0.00048061 | 0.00049653 | ~
MGS | B | 2 | 26 | 8 | 0.00048161 | 0.00049652 | ~
MGS | B | 2 | 26 | 9 | 0.00045474 | 0.00047097 | ~
MGS | B | 2 | 26 | 10 | 0.00045388 | 0.00047098 | ~
MGS | B | 2 | 26 | 11 | 0.00045413 | 0.00047098 | ~
MGS | B | 2 | 26 | 12 | 0.00045496 | 0.00047098 | ~
MGS | B | 2 | 26 | 13 | 0.00045523 | 0.00047099 | ~
MGS | B | 2 | 26 | 14 | 0.00045588 | 0.00047099 | ~
MGS | B | 2 | 26 | 15 | 0.00045606 | 0.00047259 | ~
MGS | B | 2 | 26 | 16 | 0.0004568 | 0.0004726 | ~
MGS | B | 2 | 32 | 1 | 0.00076833 | 0.0007924 | ~
MGS | B | 2 | 32 | 2 | 0.00076718 | 0.00079111 | ~
MGS | B | 2 | 32 | 3 | 0.00076873 | 0.00079407 | ~
MGS | B | 2 | 32 | 4 | 0.00076521 | 0.00078974 | ~
MGS | B | 2 | 32 | 5 | 0.00076024 | 0.00078386 | ~
MGS | B | 2 | 32 | 6 | 0.00075786 | 0.00078223 | ~
MGS | B | 2 | 32 | 7 | 0.00075581 | 0.00078029 | ~
MGS | B | 2 | 32 | 8 | 0.00075941 | 0.00078353 | ~
MGS | B | 2 | 32 | 9 | 0.00071446 | 0.00074114 | ~
MGS | B | 2 | 32 | 10 | 0.00071234 | 0.0007402 | ~
MGS | B | 2 | 32 | 11 | 0.00071283 | 0.0007406 | ~
MGS | B | 2 | 32 | 12 | 0.00071127 | 0.00073858 | ~
MGS | B | 2 | 32 | 13 | 0.00071507 | 0.00074217 | ~
MGS | B | 2 | 32 | 14 | 0.00071806 | 0.00074557 | ~
MGS | B | 2 | 32 | 15 | 0.00071518 | 0.00074178 | ~
MGS | B | 2 | 32 | 16 | 0.00071543 | 0.00074291 | ~
MGS | B | 2 | 34 | 1 | 3.0765e-05 | 3.1152e-05 | ~
MGS | B | 2 | 34 | 2 | 3.1018e-05 | 3.1327e-05 | ~
MGS | B | 2 | 34 | 3 | 3.1034e-05 | 3.1317e-05 | ~
MGS | B | 2 | 34 | 4 | 3.1066e-05 | 3.1335e-05 | ~
MGS | B | 2 | 34 | 5 | 3.1037e-05 | 3.1313e-05 | ~
MGS | B | 2 | 34 | 6 | 3.0955e-05 | 3.1294e-05 | ~
MGS | B | 2 | 34 | 7 | 3.0688e-05 | 3.0958e-05 | ~
MGS | B | 2 | 34 | 8 | 3.0686e-05 | 3.0989e-05 | ~
MGS | B | 2 | 34 | 9 | 2.8902e-05 | 2.9362e-05 | ~
MGS | B | 2 | 34 | 10 | 2.8857e-05 | 2.9363e-05 | ~
MGS | B | 2 | 34 | 11 | 2.9003e-05 | 2.9542e-05 | ~
MGS | B | 2 | 34 | 12 | 2.9031e-05 | 2.9526e-05 | ~
MGS | B | 2 | 34 | 13 | 2.9044e-05 | 2.9531e-05 | ~
MGS | B | 2 | 34 | 14 | 2.9007e-05 | 2.9536e-05 | ~
MGS | B | 2 | 34 | 15 | 2.8978e-05 | 2.9527e-05 | ~
MGS | B | 2 | 34 | 16 | 2.8927e-05 | 2.9368e-05 | ~
HGA | A | 3 | 1 | 1 | 1.1127e-07 | 9.9292e-08 | ~
HGA | A | 3 | 1 | 2 | 1.1248e-07 | 9.9104e-08 | ~
HGA | A | 3 | 1 | 3 | 1.1313e-07 | 9.9199e-08 | ~
HGA | A | 3 | 1 | 4 | 1.1134e-07 | 9.9564e-08 | ~
HGA | A | 3 | 1 | 5 | 1.1312e-07 | 9.943e-08 | ~
HGA | A | 3 | 1 | 6 | 1.1581e-07 | 1.0002e-07 | ~
HGA | A | 3 | 1 | 7 | 1.1355e-07 | 1.0046e-07 | ~
HGA | A | 3 | 1 | 8 | 1.1293e-07 | 1.0022e-07 | ~
HGA | A | 3 | 1 | 9 | 1.1314e-07 | 1.0087e-07 | ~
HGA | A | 3 | 1 | 10 | 1.1416e-07 | 1.0081e-07 | ~
HGA | A | 3 | 1 | 11 | 1.1619e-07 | 1.0092e-07 | ~
HGA | A | 3 | 1 | 12 | 1.1292e-07 | 1.005e-07 | ~
HGA | A | 3 | 1 | 13 | 1.1192e-07 | 1.008e-07 | ~
HGA | A | 3 | 1 | 14 | 1.1352e-07 | 1.0133e-07 | ~
HGA | A | 3 | 1 | 15 | 1.1257e-07 | 1.0113e-07 | ~
HGA | A | 3 | 1 | 16 | 1.1185e-07 | 1.0061e-07 | ~
HGA | A | 3 | 7 | 1 | 1.1024e-07 | 1.0462e-07 | ~
HGA | A | 3 | 7 | 2 | 1.1247e-07 | 1.0456e-07 | ~
HGA | A | 3 | 7 | 3 | 1.134e-07 | 1.0477e-07 | ~
HGA | A | 3 | 7 | 4 | 1.0877e-07 | 1.0484e-07 | ~
HGA | A | 3 | 7 | 5 | 1.1315e-07 | 1.0487e-07 | ~
HGA | A | 3 | 7 | 6 | 1.159e-07 | 1.0533e-07 | ~
HGA | A | 3 | 7 | 7 | 1.1358e-07 | 1.0544e-07 | ~
HGA | A | 3 | 7 | 8 | 1.1292e-07 | 1.0538e-07 | ~
HGA | A | 3 | 7 | 9 | 1.1315e-07 | 1.0573e-07 | ~
HGA | A | 3 | 7 | 10 | 1.1416e-07 | 1.0575e-07 | ~
HGA | A | 3 | 7 | 11 | 1.1619e-07 | 1.0576e-07 | ~
HGA | A | 3 | 7 | 12 | 1.129e-07 | 1.0579e-07 | ~
HGA | A | 3 | 7 | 13 | 1.0916e-07 | 1.0562e-07 | ~
HGA | A | 3 | 7 | 14 | 1.1354e-07 | 1.0596e-07 | ~
HGA | A | 3 | 7 | 15 | 1.1256e-07 | 1.0603e-07 | ~
HGA | A | 3 | 7 | 16 | 1.1039e-07 | 1.0612e-07 | ~
HGA | A | 3 | 11 | 1 | 1.1023e-07 | 1.077e-07 | ~
HGA | A | 3 | 11 | 2 | 1.1248e-07 | 1.0792e-07 | ~
HGA | A | 3 | 11 | 3 | 1.134e-07 | 1.0789e-07 | ~
HGA | A | 3 | 11 | 4 | 1.0879e-07 | 1.0784e-07 | ~
HGA | A | 3 | 11 | 5 | 1.1315e-07 | 1.0808e-07 | ~
HGA | A | 3 | 11 | 6 | 1.1589e-07 | 1.0829e-07 | ~
HGA | A | 3 | 11 | 7 | 1.1357e-07 | 1.0837e-07 | ~
HGA | A | 3 | 11 | 8 | 1.1291e-07 | 1.0825e-07 | ~
HGA | A | 3 | 11 | 9 | 1.1314e-07 | 1.0869e-07 | ~
HGA | A | 3 | 11 | 10 | 1.1415e-07 | 1.0876e-07 | ~
HGA | A | 3 | 11 | 11 | 1.1619e-07 | 1.0866e-07 | ~
HGA | A | 3 | 11 | 12 | 1.1291e-07 | 1.0878e-07 | ~
HGA | A | 3 | 11 | 13 | 1.0917e-07 | 1.0852e-07 | ~
HGA | A | 3 | 11 | 14 | 1.1353e-07 | 1.0857e-07 | ~
HGA | A | 3 | 11 | 15 | 1.1257e-07 | 1.0889e-07 | ~
HGA | A | 3 | 11 | 16 | 1.104e-07 | 1.0885e-07 | ~
HGA | A | 3 | 15 | 1 | 1.1024e-07 | 1.1029e-07 | ~
HGA | A | 3 | 15 | 2 | 1.1247e-07 | 1.1036e-07 | ~
HGA | A | 3 | 15 | 3 | 1.1339e-07 | 1.1034e-07 | ~
HGA | A | 3 | 15 | 4 | 1.1003e-07 | 1.1037e-07 | ~
HGA | A | 3 | 15 | 5 | 1.1316e-07 | 1.1052e-07 | ~
HGA | A | 3 | 15 | 6 | 1.1591e-07 | 1.106e-07 | ~
HGA | A | 3 | 15 | 7 | 1.1356e-07 | 1.1056e-07 | ~
HGA | A | 3 | 15 | 8 | 1.1293e-07 | 1.1059e-07 | ~
HGA | A | 3 | 15 | 9 | 1.1314e-07 | 1.1074e-07 | ~
HGA | A | 3 | 15 | 10 | 1.1417e-07 | 1.1072e-07 | ~
HGA | A | 3 | 15 | 11 | 1.1618e-07 | 1.1069e-07 | ~
HGA | A | 3 | 15 | 12 | 1.129e-07 | 1.1067e-07 | ~
HGA | A | 3 | 15 | 13 | 1.1062e-07 | 1.1064e-07 | ~
HGA | A | 3 | 15 | 14 | 1.1353e-07 | 1.1053e-07 | ~
HGA | A | 3 | 15 | 15 | 1.1256e-07 | 1.1063e-07 | ~
HGA | A | 3 | 15 | 16 | 1.1081e-07 | 1.1078e-07 | ~
HGA | A | 3 | 20 | 1 | 1.1148e-07 | 1.118e-07 | ~
HGA | A | 3 | 20 | 2 | 1.125e-07 | 1.1181e-07 | ~
HGA | A | 3 | 20 | 3 | 1.1342e-07 | 1.1179e-07 | ~
HGA | A | 3 | 20 | 4 | 1.1154e-07 | 1.1187e-07 | ~
HGA | A | 3 | 20 | 5 | 1.1317e-07 | 1.1198e-07 | ~
HGA | A | 3 | 20 | 6 | 1.1591e-07 | 1.1201e-07 | ~
HGA | A | 3 | 20 | 7 | 1.1358e-07 | 1.1196e-07 | ~
HGA | A | 3 | 20 | 8 | 1.1293e-07 | 1.1196e-07 | ~
HGA | A | 3 | 20 | 9 | 1.1315e-07 | 1.1211e-07 | ~
HGA | A | 3 | 20 | 10 | 1.1418e-07 | 1.1211e-07 | ~
HGA | A | 3 | 20 | 11 | 1.162e-07 | 1.1212e-07 | ~
HGA | A | 3 | 20 | 12 | 1.1292e-07 | 1.121e-07 | ~
HGA | A | 3 | 20 | 13 | 1.1204e-07 | 1.1212e-07 | ~
HGA | A | 3 | 20 | 14 | 1.1355e-07 | 1.1207e-07 | ~
HGA | A | 3 | 20 | 15 | 1.1258e-07 | 1.1195e-07 | ~
HGA | A | 3 | 20 | 16 | 1.1206e-07 | 1.1202e-07 | ~
HGA | A | 3 | 26 | 1 | 1.1266e-07 | 1.1285e-07 | ~
HGA | A | 3 | 26 | 2 | 1.1253e-07 | 1.1285e-07 | ~
HGA | A | 3 | 26 | 3 | 1.134e-07 | 1.1288e-07 | ~
HGA | A | 3 | 26 | 4 | 1.1267e-07 | 1.1293e-07 | ~
HGA | A | 3 | 26 | 5 | 1.1316e-07 | 1.1296e-07 | ~
HGA | A | 3 | 26 | 6 | 1.159e-07 | 1.1297e-07 | ~
HGA | A | 3 | 26 | 7 | 1.1359e-07 | 1.1295e-07 | ~
HGA | A | 3 | 26 | 8 | 1.1296e-07 | 1.1297e-07 | ~
HGA | A | 3 | 26 | 9 | 1.1317e-07 | 1.13e-07 | ~
HGA | A | 3 | 26 | 10 | 1.1418e-07 | 1.1298e-07 | ~
HGA | A | 3 | 26 | 11 | 1.1622e-07 | 1.13e-07 | ~
HGA | A | 3 | 26 | 12 | 1.1313e-07 | 1.1299e-07 | ~
HGA | A | 3 | 26 | 13 | 1.1314e-07 | 1.1301e-07 | ~
HGA | A | 3 | 26 | 14 | 1.1355e-07 | 1.1301e-07 | ~
HGA | A | 3 | 26 | 15 | 1.131e-07 | 1.1296e-07 | ~
HGA | A | 3 | 26 | 16 | 1.1308e-07 | 1.1295e-07 | ~
HGA | A | 3 | 32 | 1 | 1.1027e-07 | 1.1004e-07 | ~
HGA | A | 3 | 32 | 2 | 1.125e-07 | 1.1271e-07 | ~
HGA | A | 3 | 32 | 3 | 1.1341e-07 | 1.1368e-07 | ~
HGA | A | 3 | 32 | 4 | 1.088e-07 | 1.0867e-07 | ~
HGA | A | 3 | 32 | 5 | 1.132e-07 | 1.1332e-07 | ~
HGA | A | 3 | 32 | 6 | 1.1594e-07 | 1.161e-07 | ~
HGA | A | 3 | 32 | 7 | 1.1359e-07 | 1.137e-07 | ~
HGA | A | 3 | 32 | 8 | 1.1294e-07 | 1.1211e-07 | ~
HGA | A | 3 | 32 | 9 | 1.1315e-07 | 1.1189e-07 | ~
HGA | A | 3 | 32 | 10 | 1.1419e-07 | 1.138e-07 | ~
HGA | A | 3 | 32 | 11 | 1.1622e-07 | 1.1605e-07 | ~
HGA | A | 3 | 32 | 12 | 1.1294e-07 | 1.1269e-07 | ~
HGA | A | 3 | 32 | 13 | 1.0918e-07 | 1.0857e-07 | ~
HGA | A | 3 | 32 | 14 | 1.1367e-07 | 1.1352e-07 | ~
HGA | A | 3 | 32 | 15 | 1.1258e-07 | 1.1228e-07 | ~
HGA | A | 3 | 32 | 16 | 1.1043e-07 | 1.0965e-07 | ~
HGA | A | 3 | 34 | 1 | 1.1548e-07 | 1.1594e-07 | ~
HGA | A | 3 | 34 | 2 | 1.125e-07 | 1.1281e-07 | ~
HGA | A | 3 | 34 | 3 | 1.1328e-07 | 1.1373e-07 | ~
HGA | A | 3 | 34 | 4 | 1.1298e-07 | 1.134e-07 | ~
HGA | A | 3 | 34 | 5 | 1.1314e-07 | 1.1338e-07 | ~
HGA | A | 3 | 34 | 6 | 1.1585e-07 | 1.1617e-07 | ~
HGA | A | 3 | 34 | 7 | 1.1582e-07 | 1.1622e-07 | ~
HGA | A | 3 | 34 | 8 | 1.1347e-07 | 1.1391e-07 | ~
HGA | A | 3 | 34 | 9 | 1.1374e-07 | 1.1411e-07 | ~
HGA | A | 3 | 34 | 10 | 1.1583e-07 | 1.1627e-07 | ~
HGA | A | 3 | 34 | 11 | 1.1621e-07 | 1.1611e-07 | ~
HGA | A | 3 | 34 | 12 | 1.1293e-07 | 1.1274e-07 | ~
HGA | A | 3 | 34 | 13 | 1.1295e-07 | 1.1336e-07 | ~
HGA | A | 3 | 34 | 14 | 1.1354e-07 | 1.1357e-07 | ~
HGA | A | 3 | 34 | 15 | 1.1259e-07 | 1.1234e-07 | ~
HGA | A | 3 | 34 | 16 | 1.1588e-07 | 1.1631e-07 | ~
HGA | B | 3 | 1 | 1 | 1.1126e-07 | 9.928e-08 | ~
HGA | B | 3 | 1 | 2 | 1.1248e-07 | 9.9124e-08 | ~
HGA | B | 3 | 1 | 3 | 1.1312e-07 | 9.9186e-08 | ~
HGA | B | 3 | 1 | 4 | 1.1136e-07 | 9.9531e-08 | ~
HGA | B | 3 | 1 | 5 | 1.1311e-07 | 9.9455e-08 | ~
HGA | B | 3 | 1 | 6 | 1.1583e-07 | 9.9974e-08 | ~
HGA | B | 3 | 1 | 7 | 1.1355e-07 | 1.0053e-07 | ~
HGA | B | 3 | 1 | 8 | 1.1293e-07 | 1.0025e-07 | ~
HGA | B | 3 | 1 | 9 | 1.1315e-07 | 1.0087e-07 | ~
HGA | B | 3 | 1 | 10 | 1.1413e-07 | 1.0083e-07 | ~
HGA | B | 3 | 1 | 11 | 1.1618e-07 | 1.0088e-07 | ~
HGA | B | 3 | 1 | 12 | 1.1291e-07 | 1.0049e-07 | ~
HGA | B | 3 | 1 | 13 | 1.1192e-07 | 1.0084e-07 | ~
HGA | B | 3 | 1 | 14 | 1.1354e-07 | 1.013e-07 | ~
HGA | B | 3 | 1 | 15 | 1.1258e-07 | 1.0111e-07 | ~
HGA | B | 3 | 1 | 16 | 1.1185e-07 | 1.0047e-07 | ~
HGA | B | 3 | 7 | 1 | 1.1023e-07 | 1.0463e-07 | ~
HGA | B | 3 | 7 | 2 | 1.1247e-07 | 1.0459e-07 | ~
HGA | B | 3 | 7 | 3 | 1.1338e-07 | 1.0481e-07 | ~
HGA | B | 3 | 7 | 4 | 1.0878e-07 | 1.0485e-07 | ~
HGA | B | 3 | 7 | 5 | 1.1315e-07 | 1.0488e-07 | ~
HGA | B | 3 | 7 | 6 | 1.1591e-07 | 1.0534e-07 | ~
HGA | B | 3 | 7 | 7 | 1.1356e-07 | 1.0543e-07 | ~
HGA | B | 3 | 7 | 8 | 1.1293e-07 | 1.0529e-07 | ~
HGA | B | 3 | 7 | 9 | 1.1313e-07 | 1.0574e-07 | ~
HGA | B | 3 | 7 | 10 | 1.1415e-07 | 1.0579e-07 | ~
HGA | B | 3 | 7 | 11 | 1.1619e-07 | 1.0575e-07 | ~
HGA | B | 3 | 7 | 12 | 1.1293e-07 | 1.0582e-07 | ~
HGA | B | 3 | 7 | 13 | 1.0916e-07 | 1.0565e-07 | ~
HGA | B | 3 | 7 | 14 | 1.1354e-07 | 1.0593e-07 | ~
HGA | B | 3 | 7 | 15 | 1.1256e-07 | 1.0605e-07 | ~
HGA | B | 3 | 7 | 16 | 1.104e-07 | 1.0611e-07 | ~
HGA | B | 3 | 11 | 1 | 1.1022e-07 | 1.0766e-07 | ~
HGA | B | 3 | 11 | 2 | 1.1246e-07 | 1.079e-07 | ~
HGA | B | 3 | 11 | 3 | 1.1339e-07 | 1.0789e-07 | ~
HGA | B | 3 | 11 | 4 | 1.0879e-07 | 1.0784e-07 | ~
HGA | B | 3 | 11 | 5 | 1.1314e-07 | 1.0809e-07 | ~
HGA | B | 3 | 11 | 6 | 1.1591e-07 | 1.083e-07 | ~
HGA | B | 3 | 11 | 7 | 1.1357e-07 | 1.0837e-07 | ~
HGA | B | 3 | 11 | 8 | 1.1292e-07 | 1.0827e-07 | ~
HGA | B | 3 | 11 | 9 | 1.1314e-07 | 1.0871e-07 | ~
HGA | B | 3 | 11 | 10 | 1.1415e-07 | 1.0875e-07 | ~
HGA | B | 3 | 11 | 11 | 1.1619e-07 | 1.0865e-07 | ~
HGA | B | 3 | 11 | 12 | 1.129e-07 | 1.0871e-07 | ~
HGA | B | 3 | 11 | 13 | 1.0914e-07 | 1.0855e-07 | ~
HGA | B | 3 | 11 | 14 | 1.1354e-07 | 1.0863e-07 | ~
HGA | B | 3 | 11 | 15 | 1.1257e-07 | 1.089e-07 | ~
HGA | B | 3 | 11 | 16 | 1.1041e-07 | 1.0889e-07 | ~
HGA | B | 3 | 15 | 1 | 1.1024e-07 | 1.1033e-07 | ~
HGA | B | 3 | 15 | 2 | 1.1248e-07 | 1.1037e-07 | ~
HGA | B | 3 | 15 | 3 | 1.1338e-07 | 1.1032e-07 | ~
HGA | B | 3 | 15 | 4 | 1.1004e-07 | 1.1036e-07 | ~
HGA | B | 3 | 15 | 5 | 1.1317e-07 | 1.105e-07 | ~
HGA | B | 3 | 15 | 6 | 1.1591e-07 | 1.1059e-07 | ~
HGA | B | 3 | 15 | 7 | 1.1359e-07 | 1.1057e-07 | ~
HGA | B | 3 | 15 | 8 | 1.1292e-07 | 1.106e-07 | ~
HGA | B | 3 | 15 | 9 | 1.1315e-07 | 1.1073e-07 | ~
HGA | B | 3 | 15 | 10 | 1.1415e-07 | 1.1073e-07 | ~
HGA | B | 3 | 15 | 11 | 1.1621e-07 | 1.107e-07 | ~
HGA | B | 3 | 15 | 12 | 1.1291e-07 | 1.1065e-07 | ~
HGA | B | 3 | 15 | 13 | 1.1063e-07 | 1.1064e-07 | ~
HGA | B | 3 | 15 | 14 | 1.1354e-07 | 1.1054e-07 | ~
HGA | B | 3 | 15 | 15 | 1.1257e-07 | 1.1064e-07 | ~
HGA | B | 3 | 15 | 16 | 1.108e-07 | 1.1079e-07 | ~
HGA | B | 3 | 20 | 1 | 1.1151e-07 | 1.118e-07 | ~
HGA | B | 3 | 20 | 2 | 1.1249e-07 | 1.118e-07 | ~
HGA | B | 3 | 20 | 3 | 1.134e-07 | 1.1181e-07 | ~
HGA | B | 3 | 20 | 4 | 1.1152e-07 | 1.1186e-07 | ~
HGA | B | 3 | 20 | 5 | 1.1317e-07 | 1.1196e-07 | ~
HGA | B | 3 | 20 | 6 | 1.159e-07 | 1.1202e-07 | ~
HGA | B | 3 | 20 | 7 | 1.1358e-07 | 1.1196e-07 | ~
HGA | B | 3 | 20 | 8 | 1.1294e-07 | 1.1196e-07 | ~
HGA | B | 3 | 20 | 9 | 1.1315e-07 | 1.1211e-07 | ~
HGA | B | 3 | 20 | 10 | 1.1417e-07 | 1.1212e-07 | ~
HGA | B | 3 | 20 | 11 | 1.1622e-07 | 1.1212e-07 | ~
HGA | B | 3 | 20 | 12 | 1.1291e-07 | 1.121e-07 | ~
HGA | B | 3 | 20 | 13 | 1.1203e-07 | 1.1213e-07 | ~
HGA | B | 3 | 20 | 14 | 1.1355e-07 | 1.1207e-07 | ~
HGA | B | 3 | 20 | 15 | 1.1258e-07 | 1.1195e-07 | ~
HGA | B | 3 | 20 | 16 | 1.1205e-07 | 1.1201e-07 | ~
HGA | B | 3 | 26 | 1 | 1.127e-07 | 1.1285e-07 | ~
HGA | B | 3 | 26 | 2 | 1.1253e-07 | 1.1285e-07 | ~
HGA | B | 3 | 26 | 3 | 1.1341e-07 | 1.1289e-07 | ~
HGA | B | 3 | 26 | 4 | 1.1268e-07 | 1.1294e-07 | ~
HGA | B | 3 | 26 | 5 | 1.1318e-07 | 1.1297e-07 | ~
HGA | B | 3 | 26 | 6 | 1.1592e-07 | 1.1297e-07 | ~
HGA | B | 3 | 26 | 7 | 1.136e-07 | 1.1294e-07 | ~
HGA | B | 3 | 26 | 8 | 1.1295e-07 | 1.1297e-07 | ~
HGA | B | 3 | 26 | 9 | 1.1314e-07 | 1.1301e-07 | ~
HGA | B | 3 | 26 | 10 | 1.1417e-07 | 1.1298e-07 | ~
HGA | B | 3 | 26 | 11 | 1.162e-07 | 1.1299e-07 | ~
HGA | B | 3 | 26 | 12 | 1.1313e-07 | 1.1298e-07 | ~
HGA | B | 3 | 26 | 13 | 1.1317e-07 | 1.1301e-07 | ~
HGA | B | 3 | 26 | 14 | 1.1355e-07 | 1.1301e-07 | ~
HGA | B | 3 | 26 | 15 | 1.1311e-07 | 1.1295e-07 | ~
HGA | B | 3 | 26 | 16 | 1.1308e-07 | 1.1294e-07 | ~
HGA | B | 3 | 32 | 1 | 1.1026e-07 | 1.1006e-07 | ~
HGA | B | 3 | 32 | 2 | 1.1249e-07 | 1.1271e-07 | ~
HGA | B | 3 | 32 | 3 | 1.1341e-07 | 1.1367e-07 | ~
HGA | B | 3 | 32 | 4 | 1.088e-07 | 1.0869e-07 | ~
HGA | B | 3 | 32 | 5 | 1.1318e-07 | 1.1332e-07 | ~
HGA | B | 3 | 32 | 6 | 1.1593e-07 | 1.161e-07 | ~
HGA | B | 3 | 32 | 7 | 1.1362e-07 | 1.1371e-07 | ~
HGA | B | 3 | 32 | 8 | 1.1295e-07 | 1.1211e-07 | ~
HGA | B | 3 | 32 | 9 | 1.1317e-07 | 1.1189e-07 | ~
HGA | B | 3 | 32 | 10 | 1.1417e-07 | 1.138e-07 | ~
HGA | B | 3 | 32 | 11 | 1.1622e-07 | 1.1604e-07 | ~
HGA | B | 3 | 32 | 12 | 1.1294e-07 | 1.1269e-07 | ~
HGA | B | 3 | 32 | 13 | 1.0917e-07 | 1.0857e-07 | ~
HGA | B | 3 | 32 | 14 | 1.1367e-07 | 1.1352e-07 | ~
HGA | B | 3 | 32 | 15 | 1.1258e-07 | 1.1228e-07 | ~
HGA | B | 3 | 32 | 16 | 1.1042e-07 | 1.0964e-07 | ~
HGA | B | 3 | 34 | 1 | 1.1549e-07 | 1.1595e-07 | ~
HGA | B | 3 | 34 | 2 | 1.1249e-07 | 1.1281e-07 | ~
HGA | B | 3 | 34 | 3 | 1.1328e-07 | 1.1373e-07 | ~
HGA | B | 3 | 34 | 4 | 1.1296e-07 | 1.1343e-07 | ~
HGA | B | 3 | 34 | 5 | 1.1312e-07 | 1.1338e-07 | ~
HGA | B | 3 | 34 | 6 | 1.1585e-07 | 1.1617e-07 | ~
HGA | B | 3 | 34 | 7 | 1.1582e-07 | 1.1624e-07 | ~
HGA | B | 3 | 34 | 8 | 1.1348e-07 | 1.1388e-07 | ~
HGA | B | 3 | 34 | 9 | 1.1374e-07 | 1.1413e-07 | ~
HGA | B | 3 | 34 | 10 | 1.1583e-07 | 1.1627e-07 | ~
HGA | B | 3 | 34 | 11 | 1.1621e-07 | 1.1611e-07 | ~
HGA | B | 3 | 34 | 12 | 1.1294e-07 | 1.1274e-07 | ~
HGA | B | 3 | 34 | 13 | 1.1294e-07 | 1.1339e-07 | ~
HGA | B | 3 | 34 | 14 | 1.1356e-07 | 1.1357e-07 | ~
HGA | B | 3 | 34 | 15 | 1.1259e-07 | 1.1234e-07 | ~
HGA | B | 3 | 34 | 16 | 1.1588e-07 | 1.1629e-07 | ~
HGB | A | 3 | 1 | 1 | 1.1151e-07 | 1.0017e-07 | ~
HGB | A | 3 | 1 | 2 | 1.1247e-07 | 1.0006e-07 | ~
HGB | A | 3 | 1 | 3 | 1.1336e-07 | 1.0023e-07 | ~
HGB | A | 3 | 1 | 4 | 1.1168e-07 | 1.0033e-07 | ~
HGB | A | 3 | 1 | 5 | 1.131e-07 | 1.0048e-07 | ~
HGB | A | 3 | 1 | 6 | 1.1603e-07 | 1.0057e-07 | ~
HGB | A | 3 | 1 | 7 | 1.1402e-07 | 1.0074e-07 | ~
HGB | A | 3 | 1 | 8 | 1.1311e-07 | 1.0078e-07 | ~
HGB | A | 3 | 1 | 9 | 1.131e-07 | 1.0044e-07 | ~
HGB | A | 3 | 1 | 10 | 1.1401e-07 | 1.0005e-07 | ~
HGB | A | 3 | 1 | 11 | 1.1618e-07 | 9.9491e-08 | ~
HGB | A | 3 | 1 | 12 | 1.1267e-07 | 9.9793e-08 | ~
HGB | A | 3 | 1 | 13 | 1.1152e-07 | 1.0028e-07 | ~
HGB | A | 3 | 1 | 14 | 1.1324e-07 | 1.0047e-07 | ~
HGB | A | 3 | 1 | 15 | 1.1226e-07 | 1.0043e-07 | ~
HGB | A | 3 | 1 | 16 | 1.1167e-07 | 1.0055e-07 | ~
HGB | A | 3 | 7 | 1 | 1.1043e-07 | 1.0537e-07 | ~
HGB | A | 3 | 7 | 2 | 1.1249e-07 | 1.0528e-07 | ~
HGB | A | 3 | 7 | 3 | 1.1345e-07 | 1.0528e-07 | ~
HGB | A | 3 | 7 | 4 | 1.0876e-07 | 1.0537e-07 | ~
HGB | A | 3 | 7 | 5 | 1.1309e-07 | 1.0554e-07 | ~
HGB | A | 3 | 7 | 6 | 1.1602e-07 | 1.0556e-07 | ~
HGB | A | 3 | 7 | 7 | 1.1403e-07 | 1.0563e-07 | ~
HGB | A | 3 | 7 | 8 | 1.1312e-07 | 1.0563e-07 | ~
HGB | A | 3 | 7 | 9 | 1.1311e-07 | 1.055e-07 | ~
HGB | A | 3 | 7 | 10 | 1.14e-07 | 1.0528e-07 | ~
HGB | A | 3 | 7 | 11 | 1.1619e-07 | 1.0496e-07 | ~
HGB | A | 3 | 7 | 12 | 1.1269e-07 | 1.0497e-07 | ~
HGB | A | 3 | 7 | 13 | 1.0851e-07 | 1.0521e-07 | ~
HGB | A | 3 | 7 | 14 | 1.1334e-07 | 1.0544e-07 | ~
HGB | A | 3 | 7 | 15 | 1.1228e-07 | 1.0552e-07 | ~
HGB | A | 3 | 7 | 16 | 1.1014e-07 | 1.0545e-07 | ~
HGB | A | 3 | 11 | 1 | 1.1045e-07 | 1.0818e-07 | ~
HGB | A | 3 | 11 | 2 | 1.1241e-07 | 1.0823e-07 | ~
HGB | A | 3 | 11 | 3 | 1.1347e-07 | 1.0823e-07 | ~
HGB | A | 3 | 11 | 4 | 1.0875e-07 | 1.083e-07 | ~
HGB | A | 3 | 11 | 5 | 1.1306e-07 | 1.0839e-07 | ~
HGB | A | 3 | 11 | 6 | 1.1604e-07 | 1.0844e-07 | ~
HGB | A | 3 | 11 | 7 | 1.1387e-07 | 1.0845e-07 | ~
HGB | A | 3 | 11 | 8 | 1.1326e-07 | 1.0852e-07 | ~
HGB | A | 3 | 11 | 9 | 1.1309e-07 | 1.0846e-07 | ~
HGB | A | 3 | 11 | 10 | 1.1402e-07 | 1.0829e-07 | ~
HGB | A | 3 | 11 | 11 | 1.1619e-07 | 1.0813e-07 | ~
HGB | A | 3 | 11 | 12 | 1.1267e-07 | 1.0803e-07 | ~
HGB | A | 3 | 11 | 13 | 1.0853e-07 | 1.0816e-07 | ~
HGB | A | 3 | 11 | 14 | 1.1336e-07 | 1.083e-07 | ~
HGB | A | 3 | 11 | 15 | 1.1227e-07 | 1.0845e-07 | ~
HGB | A | 3 | 11 | 16 | 1.1017e-07 | 1.0847e-07 | ~
HGB | A | 3 | 15 | 1 | 1.1046e-07 | 1.1054e-07 | ~
HGB | A | 3 | 15 | 2 | 1.1248e-07 | 1.1068e-07 | ~
HGB | A | 3 | 15 | 3 | 1.1346e-07 | 1.1062e-07 | ~
HGB | A | 3 | 15 | 4 | 1.1039e-07 | 1.1065e-07 | ~
HGB | A | 3 | 15 | 5 | 1.1308e-07 | 1.1068e-07 | ~
HGB | A | 3 | 15 | 6 | 1.1604e-07 | 1.107e-07 | ~
HGB | A | 3 | 15 | 7 | 1.1404e-07 | 1.1072e-07 | ~
HGB | A | 3 | 15 | 8 | 1.1302e-07 | 1.1073e-07 | ~
HGB | A | 3 | 15 | 9 | 1.1303e-07 | 1.1078e-07 | ~
HGB | A | 3 | 15 | 10 | 1.1402e-07 | 1.1069e-07 | ~
HGB | A | 3 | 15 | 11 | 1.1619e-07 | 1.1055e-07 | ~
HGB | A | 3 | 15 | 12 | 1.127e-07 | 1.1049e-07 | ~
HGB | A | 3 | 15 | 13 | 1.1022e-07 | 1.105e-07 | ~
HGB | A | 3 | 15 | 14 | 1.1335e-07 | 1.1057e-07 | ~
HGB | A | 3 | 15 | 15 | 1.1227e-07 | 1.1068e-07 | ~
HGB | A | 3 | 15 | 16 | 1.1048e-07 | 1.1076e-07 | ~
HGB | A | 3 | 20 | 1 | 1.1169e-07 | 1.1199e-07 | ~
HGB | A | 3 | 20 | 2 | 1.125e-07 | 1.1208e-07 | ~
HGB | A | 3 | 20 | 3 | 1.1348e-07 | 1.1202e-07 | ~
HGB | A | 3 | 20 | 4 | 1.1172e-07 | 1.1205e-07 | ~
HGB | A | 3 | 20 | 5 | 1.1311e-07 | 1.1209e-07 | ~
HGB | A | 3 | 20 | 6 | 1.1604e-07 | 1.1208e-07 | ~
HGB | A | 3 | 20 | 7 | 1.1402e-07 | 1.1207e-07 | ~
HGB | A | 3 | 20 | 8 | 1.1308e-07 | 1.1208e-07 | ~
HGB | A | 3 | 20 | 9 | 1.1305e-07 | 1.1216e-07 | ~
HGB | A | 3 | 20 | 10 | 1.1407e-07 | 1.1211e-07 | ~
HGB | A | 3 | 20 | 11 | 1.162e-07 | 1.1209e-07 | ~
HGB | A | 3 | 20 | 12 | 1.1269e-07 | 1.1198e-07 | ~
HGB | A | 3 | 20 | 13 | 1.1168e-07 | 1.1196e-07 | ~
HGB | A | 3 | 20 | 14 | 1.1336e-07 | 1.1203e-07 | ~
HGB | A | 3 | 20 | 15 | 1.123e-07 | 1.1201e-07 | ~
HGB | A | 3 | 20 | 16 | 1.1189e-07 | 1.1213e-07 | ~
HGB | A | 3 | 26 | 1 | 1.1284e-07 | 1.1295e-07 | ~
HGB | A | 3 | 26 | 2 | 1.1285e-07 | 1.1301e-07 | ~
HGB | A | 3 | 26 | 3 | 1.1348e-07 | 1.1302e-07 | ~
HGB | A | 3 | 26 | 4 | 1.129e-07 | 1.1303e-07 | ~
HGB | A | 3 | 26 | 5 | 1.1314e-07 | 1.1304e-07 | ~
HGB | A | 3 | 26 | 6 | 1.1607e-07 | 1.1305e-07 | ~
HGB | A | 3 | 26 | 7 | 1.1408e-07 | 1.1306e-07 | ~
HGB | A | 3 | 26 | 8 | 1.1309e-07 | 1.1309e-07 | ~
HGB | A | 3 | 26 | 9 | 1.131e-07 | 1.1317e-07 | ~
HGB | A | 3 | 26 | 10 | 1.1401e-07 | 1.1311e-07 | ~
HGB | A | 3 | 26 | 11 | 1.1622e-07 | 1.1311e-07 | ~
HGB | A | 3 | 26 | 12 | 1.1278e-07 | 1.1304e-07 | ~
HGB | A | 3 | 26 | 13 | 1.1279e-07 | 1.1305e-07 | ~
HGB | A | 3 | 26 | 14 | 1.134e-07 | 1.1303e-07 | ~
HGB | A | 3 | 26 | 15 | 1.1283e-07 | 1.1308e-07 | ~
HGB | A | 3 | 26 | 16 | 1.1283e-07 | 1.1305e-07 | ~
HGB | A | 3 | 32 | 1 | 1.105e-07 | 1.1019e-07 | ~
HGB | A | 3 | 32 | 2 | 1.1252e-07 | 1.1274e-07 | ~
HGB | A | 3 | 32 | 3 | 1.1343e-07 | 1.1373e-07 | ~
HGB | A | 3 | 32 | 4 | 1.0886e-07 | 1.0877e-07 | ~
HGB | A | 3 | 32 | 5 | 1.131e-07 | 1.1334e-07 | ~
HGB | A | 3 | 32 | 6 | 1.1606e-07 | 1.1626e-07 | ~
HGB | A | 3 | 32 | 7 | 1.1405e-07 | 1.1414e-07 | ~
HGB | A | 3 | 32 | 8 | 1.1303e-07 | 1.1244e-07 | ~
HGB | A | 3 | 32 | 9 | 1.1305e-07 | 1.1206e-07 | ~
HGB | A | 3 | 32 | 10 | 1.1402e-07 | 1.1403e-07 | ~
HGB | A | 3 | 32 | 11 | 1.162e-07 | 1.1636e-07 | ~
HGB | A | 3 | 32 | 12 | 1.1271e-07 | 1.1266e-07 | ~
HGB | A | 3 | 32 | 13 | 1.0858e-07 | 1.0822e-07 | ~
HGB | A | 3 | 32 | 14 | 1.1328e-07 | 1.1348e-07 | ~
HGB | A | 3 | 32 | 15 | 1.1231e-07 | 1.1228e-07 | ~
HGB | A | 3 | 32 | 16 | 1.1023e-07 | 1.0983e-07 | ~
HGB | A | 3 | 34 | 1 | 1.1578e-07 | 1.1618e-07 | ~
HGB | A | 3 | 34 | 2 | 1.1249e-07 | 1.1282e-07 | ~
HGB | A | 3 | 34 | 3 | 1.1339e-07 | 1.1382e-07 | ~
HGB | A | 3 | 34 | 4 | 1.1302e-07 | 1.1344e-07 | ~
HGB | A | 3 | 34 | 5 | 1.1311e-07 | 1.1339e-07 | ~
HGB | A | 3 | 34 | 6 | 1.1606e-07 | 1.1636e-07 | ~
HGB | A | 3 | 34 | 7 | 1.158e-07 | 1.1618e-07 | ~
HGB | A | 3 | 34 | 8 | 1.1379e-07 | 1.1417e-07 | ~
HGB | A | 3 | 34 | 9 | 1.1392e-07 | 1.1431e-07 | ~
HGB | A | 3 | 34 | 10 | 1.1614e-07 | 1.1657e-07 | ~
HGB | A | 3 | 34 | 11 | 1.162e-07 | 1.1645e-07 | ~
HGB | A | 3 | 34 | 12 | 1.1269e-07 | 1.127e-07 | ~
HGB | A | 3 | 34 | 13 | 1.1304e-07 | 1.1345e-07 | ~
HGB | A | 3 | 34 | 14 | 1.1325e-07 | 1.1355e-07 | ~
HGB | A | 3 | 34 | 15 | 1.1229e-07 | 1.1235e-07 | ~
HGB | A | 3 | 34 | 16 | 1.1594e-07 | 1.1636e-07 | ~
HGB | B | 3 | 1 | 1 | 1.115e-07 | 1.0018e-07 | ~
HGB | B | 3 | 1 | 2 | 1.1249e-07 | 1.0013e-07 | ~
HGB | B | 3 | 1 | 3 | 1.1335e-07 | 1.0021e-07 | ~
HGB | B | 3 | 1 | 4 | 1.1169e-07 | 1.003e-07 | ~
HGB | B | 3 | 1 | 5 | 1.1309e-07 | 1.0047e-07 | ~
HGB | B | 3 | 1 | 6 | 1.1603e-07 | 1.0057e-07 | ~
HGB | B | 3 | 1 | 7 | 1.1401e-07 | 1.0081e-07 | ~
HGB | B | 3 | 1 | 8 | 1.131e-07 | 1.0078e-07 | ~
HGB | B | 3 | 1 | 9 | 1.1309e-07 | 1.0042e-07 | ~
HGB | B | 3 | 1 | 10 | 1.1401e-07 | 1.0007e-07 | ~
HGB | B | 3 | 1 | 11 | 1.1617e-07 | 9.9484e-08 | ~
HGB | B | 3 | 1 | 12 | 1.1267e-07 | 9.9713e-08 | ~
HGB | B | 3 | 1 | 13 | 1.115e-07 | 1.0026e-07 | ~
HGB | B | 3 | 1 | 14 | 1.1325e-07 | 1.0047e-07 | ~
HGB | B | 3 | 1 | 15 | 1.1226e-07 | 1.0043e-07 | ~
HGB | B | 3 | 1 | 16 | 1.1165e-07 | 1.0052e-07 | ~
HGB | B | 3 | 7 | 1 | 1.1041e-07 | 1.0541e-07 | ~
HGB | B | 3 | 7 | 2 | 1.1248e-07 | 1.0534e-07 | ~
HGB | B | 3 | 7 | 3 | 1.1347e-07 | 1.0527e-07 | ~
HGB | B | 3 | 7 | 4 | 1.0875e-07 | 1.0541e-07 | ~
HGB | B | 3 | 7 | 5 | 1.1309e-07 | 1.055e-07 | ~
HGB | B | 3 | 7 | 6 | 1.1605e-07 | 1.0556e-07 | ~
HGB | B | 3 | 7 | 7 | 1.1401e-07 | 1.0565e-07 | ~
HGB | B | 3 | 7 | 8 | 1.1314e-07 | 1.0563e-07 | ~
HGB | B | 3 | 7 | 9 | 1.1311e-07 | 1.0551e-07 | ~
HGB | B | 3 | 7 | 10 | 1.1403e-07 | 1.0532e-07 | ~
HGB | B | 3 | 7 | 11 | 1.1619e-07 | 1.05e-07 | ~
HGB | B | 3 | 7 | 12 | 1.1269e-07 | 1.0495e-07 | ~
HGB | B | 3 | 7 | 13 | 1.0852e-07 | 1.0516e-07 | ~
HGB | B | 3 | 7 | 14 | 1.1335e-07 | 1.0543e-07 | ~
HGB | B | 3 | 7 | 15 | 1.1226e-07 | 1.0551e-07 | ~
HGB | B | 3 | 7 | 16 | 1.1012e-07 | 1.0546e-07 | ~
HGB | B | 3 | 11 | 1 | 1.1047e-07 | 1.0818e-07 | ~
HGB | B | 3 | 11 | 2 | 1.124e-07 | 1.0819e-07 | ~
HGB | B | 3 | 11 | 3 | 1.1346e-07 | 1.0825e-07 | ~
HGB | B | 3 | 11 | 4 | 1.0876e-07 | 1.0831e-07 | ~
HGB | B | 3 | 11 | 5 | 1.1307e-07 | 1.0843e-07 | ~
HGB | B | 3 | 11 | 6 | 1.1602e-07 | 1.0843e-07 | ~
HGB | B | 3 | 11 | 7 | 1.1388e-07 | 1.0847e-07 | ~
HGB | B | 3 | 11 | 8 | 1.1326e-07 | 1.0846e-07 | ~
HGB | B | 3 | 11 | 9 | 1.1308e-07 | 1.0841e-07 | ~
HGB | B | 3 | 11 | 10 | 1.1403e-07 | 1.0828e-07 | ~
HGB | B | 3 | 11 | 11 | 1.1618e-07 | 1.0811e-07 | ~
HGB | B | 3 | 11 | 12 | 1.127e-07 | 1.0802e-07 | ~
HGB | B | 3 | 11 | 13 | 1.0854e-07 | 1.0815e-07 | ~
HGB | B | 3 | 11 | 14 | 1.1336e-07 | 1.083e-07 | ~
HGB | B | 3 | 11 | 15 | 1.1228e-07 | 1.0844e-07 | ~
HGB | B | 3 | 11 | 16 | 1.1017e-07 | 1.0851e-07 | ~
HGB | B | 3 | 15 | 1 | 1.1045e-07 | 1.1055e-07 | ~
HGB | B | 3 | 15 | 2 | 1.1248e-07 | 1.107e-07 | ~
HGB | B | 3 | 15 | 3 | 1.1344e-07 | 1.1063e-07 | ~
HGB | B | 3 | 15 | 4 | 1.1043e-07 | 1.1064e-07 | ~
HGB | B | 3 | 15 | 5 | 1.1309e-07 | 1.1068e-07 | ~
HGB | B | 3 | 15 | 6 | 1.1603e-07 | 1.107e-07 | ~
HGB | B | 3 | 15 | 7 | 1.1403e-07 | 1.1073e-07 | ~
HGB | B | 3 | 15 | 8 | 1.13e-07 | 1.1075e-07 | ~
HGB | B | 3 | 15 | 9 | 1.1301e-07 | 1.1075e-07 | ~
HGB | B | 3 | 15 | 10 | 1.1403e-07 | 1.1068e-07 | ~
HGB | B | 3 | 15 | 11 | 1.162e-07 | 1.1057e-07 | ~
HGB | B | 3 | 15 | 12 | 1.1268e-07 | 1.1052e-07 | ~
HGB | B | 3 | 15 | 13 | 1.1024e-07 | 1.1049e-07 | ~
HGB | B | 3 | 15 | 14 | 1.1335e-07 | 1.1059e-07 | ~
HGB | B | 3 | 15 | 15 | 1.1227e-07 | 1.1069e-07 | ~
HGB | B | 3 | 15 | 16 | 1.1048e-07 | 1.1077e-07 | ~
HGB | B | 3 | 20 | 1 | 1.117e-07 | 1.1197e-07 | ~
HGB | B | 3 | 20 | 2 | 1.125e-07 | 1.1211e-07 | ~
HGB | B | 3 | 20 | 3 | 1.1347e-07 | 1.1203e-07 | ~
HGB | B | 3 | 20 | 4 | 1.1171e-07 | 1.1204e-07 | ~
HGB | B | 3 | 20 | 5 | 1.1311e-07 | 1.1208e-07 | ~
HGB | B | 3 | 20 | 6 | 1.1605e-07 | 1.1209e-07 | ~
HGB | B | 3 | 20 | 7 | 1.1405e-07 | 1.1208e-07 | ~
HGB | B | 3 | 20 | 8 | 1.1309e-07 | 1.1209e-07 | ~
HGB | B | 3 | 20 | 9 | 1.1308e-07 | 1.1215e-07 | ~
HGB | B | 3 | 20 | 10 | 1.1405e-07 | 1.1213e-07 | ~
HGB | B | 3 | 20 | 11 | 1.162e-07 | 1.121e-07 | ~
HGB | B | 3 | 20 | 12 | 1.1269e-07 | 1.12e-07 | ~
HGB | B | 3 | 20 | 13 | 1.1168e-07 | 1.1195e-07 | ~
HGB | B | 3 | 20 | 14 | 1.1336e-07 | 1.1204e-07 | ~
HGB | B | 3 | 20 | 15 | 1.1228e-07 | 1.1199e-07 | ~
HGB | B | 3 | 20 | 16 | 1.1191e-07 | 1.1213e-07 | ~
HGB | B | 3 | 26 | 1 | 1.1286e-07 | 1.1296e-07 | ~
HGB | B | 3 | 26 | 2 | 1.1286e-07 | 1.1301e-07 | ~
HGB | B | 3 | 26 | 3 | 1.1348e-07 | 1.1302e-07 | ~
HGB | B | 3 | 26 | 4 | 1.1287e-07 | 1.1304e-07 | ~
HGB | B | 3 | 26 | 5 | 1.1316e-07 | 1.1304e-07 | ~
HGB | B | 3 | 26 | 6 | 1.1608e-07 | 1.1305e-07 | ~
HGB | B | 3 | 26 | 7 | 1.1407e-07 | 1.1304e-07 | ~
HGB | B | 3 | 26 | 8 | 1.131e-07 | 1.1308e-07 | ~
HGB | B | 3 | 26 | 9 | 1.1309e-07 | 1.1317e-07 | ~
HGB | B | 3 | 26 | 10 | 1.1401e-07 | 1.1313e-07 | ~
HGB | B | 3 | 26 | 11 | 1.1622e-07 | 1.131e-07 | ~
HGB | B | 3 | 26 | 12 | 1.1277e-07 | 1.1305e-07 | ~
HGB | B | 3 | 26 | 13 | 1.1279e-07 | 1.1304e-07 | ~
HGB | B | 3 | 26 | 14 | 1.1338e-07 | 1.1302e-07 | ~
HGB | B | 3 | 26 | 15 | 1.1283e-07 | 1.1309e-07 | ~
HGB | B | 3 | 26 | 16 | 1.1282e-07 | 1.1305e-07 | ~
HGB | B | 3 | 32 | 1 | 1.1048e-07 | 1.102e-07 | ~
HGB | B | 3 | 32 | 2 | 1.125e-07 | 1.1273e-07 | ~
HGB | B | 3 | 32 | 3 | 1.1343e-07 | 1.1372e-07 | ~
HGB | B | 3 | 32 | 4 | 1.0884e-07 | 1.0878e-07 | ~
HGB | B | 3 | 32 | 5 | 1.131e-07 | 1.1334e-07 | ~
HGB | B | 3 | 32 | 6 | 1.1603e-07 | 1.1625e-07 | ~
HGB | B | 3 | 32 | 7 | 1.1405e-07 | 1.1414e-07 | ~
HGB | B | 3 | 32 | 8 | 1.1304e-07 | 1.1244e-07 | ~
HGB | B | 3 | 32 | 9 | 1.1303e-07 | 1.1207e-07 | ~
HGB | B | 3 | 32 | 10 | 1.1404e-07 | 1.1403e-07 | ~
HGB | B | 3 | 32 | 11 | 1.1621e-07 | 1.1636e-07 | ~
HGB | B | 3 | 32 | 12 | 1.1271e-07 | 1.1266e-07 | ~
HGB | B | 3 | 32 | 13 | 1.0857e-07 | 1.0823e-07 | ~
HGB | B | 3 | 32 | 14 | 1.1329e-07 | 1.1348e-07 | ~
HGB | B | 3 | 32 | 15 | 1.1231e-07 | 1.1229e-07 | ~
HGB | B | 3 | 32 | 16 | 1.1022e-07 | 1.0982e-07 | ~
HGB | B | 3 | 34 | 1 | 1.1577e-07 | 1.1621e-07 | ~
HGB | B | 3 | 34 | 2 | 1.1251e-07 | 1.1282e-07 | ~
HGB | B | 3 | 34 | 3 | 1.1339e-07 | 1.1382e-07 | ~
HGB | B | 3 | 34 | 4 | 1.1304e-07 | 1.1343e-07 | ~
HGB | B | 3 | 34 | 5 | 1.131e-07 | 1.1339e-07 | ~
HGB | B | 3 | 34 | 6 | 1.1605e-07 | 1.1636e-07 | ~
HGB | B | 3 | 34 | 7 | 1.158e-07 | 1.1619e-07 | ~
HGB | B | 3 | 34 | 8 | 1.138e-07 | 1.1417e-07 | ~
HGB | B | 3 | 34 | 9 | 1.1393e-07 | 1.1432e-07 | ~
HGB | B | 3 | 34 | 10 | 1.1614e-07 | 1.1659e-07 | ~
HGB | B | 3 | 34 | 11 | 1.162e-07 | 1.1645e-07 | ~
HGB | B | 3 | 34 | 12 | 1.127e-07 | 1.127e-07 | ~
HGB | B | 3 | 34 | 13 | 1.1305e-07 | 1.1341e-07 | ~
HGB | B | 3 | 34 | 14 | 1.1325e-07 | 1.1355e-07 | ~
HGB | B | 3 | 34 | 15 | 1.1229e-07 | 1.1235e-07 | ~
HGB | B | 3 | 34 | 16 | 1.1595e-07 | 1.1635e-07 | ~

## TEB gain transition
Band | HAM | Band Median |  | Band Average |  | Tmax | T Upper Bount
|  | Eside A | Eside B | Eside A | Eside B
|  | Ttrans | Ttrans | Ttrans | Ttrans
M13 | A | 343.1427 | 342.7818 | 343.2433 | 342.8629 | 343 | 348
M13 | B | 343.145 | 342.8012 | 343.2416 | 342.8804 | 343 | 348
Band | HAM | Detector | Detector Dependent |  | Tmax | T Upper Bount |  | Notes:
|  |  | Eside A | Eside B |  |  |  | 1. This table provides VIIRS F2 TEB gain transitions.
|  |  | Ttrans | Ttrans |  |  |  | 2. The results were derived from RC-1 using the LABB source (p4 for Aside and p1 for Bside).
M13 | A | 1 | 342.1124 | 341.6877 | 343 | 348 |  | 3. Transition was determined from partial views of the LABB source.
M13 | A | 2 | 343.7796 | 343.4197 | 343 | 348 |  | 4. Ttrans, high gain Tmax, and specified T upper bound values are in K.
M13 | A | 3 | 341.7884 | 341.4116 | 343 | 348 |  | 5. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
M13 | A | 4 | 343.7454 | 343.4052 | 343 | 348 |  | 6. The test was set to operational day mode.
M13 | A | 5 | 342.0891 | 341.7228 | 343 | 348
M13 | A | 6 | 344.3215 | 343.9666 | 343 | 348
M13 | A | 7 | 342.4862 | 342.1006 | 343 | 348
M13 | A | 8 | 344.4466 | 344.0849 | 343 | 348
M13 | A | 9 | 342.4182 | 342.0285 | 343 | 348
M13 | A | 10 | 344.0556 | 343.6983 | 343 | 348
M13 | A | 11 | 342.54 | 342.1584 | 343 | 348
M13 | A | 12 | 344.3741 | 344.0222 | 343 | 348
M13 | A | 13 | 342.2699 | 341.8732 | 343 | 348
M13 | A | 14 | 344.1002 | 343.717 | 343 | 348
M13 | A | 15 | 342.2151 | 341.7717 | 343 | 348
M13 | A | 16 | 345.1508 | 344.738 | 343 | 348
M13 | B | 1 | 342.0963 | 341.7084 | 343 | 348
M13 | B | 2 | 343.7875 | 343.4389 | 343 | 348
M13 | B | 3 | 341.7847 | 341.4264 | 343 | 348
M13 | B | 4 | 343.756 | 343.4327 | 343 | 348
M13 | B | 5 | 342.0877 | 341.7473 | 343 | 348
M13 | B | 6 | 344.3334 | 343.9848 | 343 | 348
M13 | B | 7 | 342.4699 | 342.1169 | 343 | 348
M13 | B | 8 | 344.4613 | 344.1097 | 343 | 348
M13 | B | 9 | 342.3942 | 342.0366 | 343 | 348
M13 | B | 10 | 344.0592 | 343.7052 | 343 | 348
M13 | B | 11 | 342.5339 | 342.1698 | 343 | 348
M13 | B | 12 | 344.3839 | 344.0414 | 343 | 348
M13 | B | 13 | 342.2634 | 341.8776 | 343 | 348
M13 | B | 14 | 344.1064 | 343.7539 | 343 | 348
M13 | B | 15 | 342.2039 | 341.7977 | 343 | 348
M13 | B | 16 | 345.1442 | 344.7386 | 343 | 348

## RSB gain transition
Band | HAM | Band Median |  | Band Average |  | Lmax | L Upper Bount
|  | Eside A | Eside B | Eside A | Eside B
|  | Ltrans | Ltrans | Ltrans | Ltrans
M1 | A | 161.9001 | #NUM! | 161.867 | #DIV/0! | 135 | 202.5
M1 | B | 161.6982 | #NUM! | 161.6948 | #DIV/0! | 135 | 202.5
M2 | A | 148.1526 | #NUM! | 148.0261 | #DIV/0! | 127 | 190.5
M2 | B | 148.0087 | #NUM! | 147.9042 | #DIV/0! | 127 | 190.5
M3 | A | 116.3217 | #NUM! | 116.448 | #DIV/0! | 107 | 128.4
M3 | B | 116.4373 | #NUM! | 116.4336 | #DIV/0! | 107 | 128.4
M4 | A | 90.2449 | #NUM! | 90.2502 | #DIV/0! | 78 | 93.6
M4 | B | 90.2876 | #NUM! | 90.2505 | #DIV/0! | 78 | 93.6
M5 | A | 61.9415 | #NUM! | 61.9827 | #DIV/0! | 59 | 70.8
M5 | B | 62.0058 | #NUM! | 61.9962 | #DIV/0! | 59 | 70.8
M7 | A | 30.7376 | #NUM! | 30.749 | #DIV/0! | 29 | 34.8
M7 | B | 30.7274 | #NUM! | 30.7331 | #DIV/0! | 29 | 34.8
Band | HAM | Detector | Detector Dependent |  | Lmax | L Upper Bount |  | Notes:
|  |  | Eside A | Eside B |  |  |  | 1. This table provides VIIRS F2 RSB gain transitions.
|  |  | Ltrans | Ltrans |  |  |  | 2. The results were derived from RC-1 using the SIS-100 source (p4 second run for Aside, p2 for Bside).
M1 | A | 1 | 162.1884 |  | 135 | 202.5 |  | 3. Transition was determined from partial views of the SIS-100 source.
M1 | A | 2 | 163.1463 |  | 135 | 202.5 |  | 4. Ltrans, high gain Lmax, and specified L upper bound values are in radiance.
M1 | A | 3 | 161.477 |  | 135 | 202.5 |  | 5. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
M1 | A | 4 | 161.1462 |  | 135 | 202.5 |  | 6. The test was set to operational day mode.
M1 | A | 5 | 162.1088 |  | 135 | 202.5
M1 | A | 6 | 162.1604 |  | 135 | 202.5
M1 | A | 7 | 161.5731 |  | 135 | 202.5
M1 | A | 8 | 162.0006 |  | 135 | 202.5
M1 | A | 9 | 161.1756 |  | 135 | 202.5
M1 | A | 10 | 161.7996 |  | 135 | 202.5
M1 | A | 11 | 161.2972 |  | 135 | 202.5
M1 | A | 12 | 163.2067 |  | 135 | 202.5
M1 | A | 13 | 162.6874 |  | 135 | 202.5
M1 | A | 14 | 162.3274 |  | 135 | 202.5
M1 | A | 15 | 160.6135 |  | 135 | 202.5
M1 | A | 16 | 160.9637 |  | 135 | 202.5
M1 | B | 1 | 161.7052 |  | 135 | 202.5
M1 | B | 2 | 162.7907 |  | 135 | 202.5
M1 | B | 3 | 162.1006 |  | 135 | 202.5
M1 | B | 4 | 161.4267 |  | 135 | 202.5
M1 | B | 5 | 161.6913 |  | 135 | 202.5
M1 | B | 6 | 161.8289 |  | 135 | 202.5
M1 | B | 7 | 161.1235 |  | 135 | 202.5
M1 | B | 8 | 161.5851 |  | 135 | 202.5
M1 | B | 9 | 161.8169 |  | 135 | 202.5
M1 | B | 10 | 161.4 |  | 135 | 202.5
M1 | B | 11 | 161.0717 |  | 135 | 202.5
M1 | B | 12 | 162.7599 |  | 135 | 202.5
M1 | B | 13 | 162.5315 |  | 135 | 202.5
M1 | B | 14 | 162.1112 |  | 135 | 202.5
M1 | B | 15 | 160.5839 |  | 135 | 202.5
M1 | B | 16 | 160.589 |  | 135 | 202.5
M2 | A | 1 | 148.2553 |  | 127 | 190.5
M2 | A | 2 | 148.2691 |  | 127 | 190.5
M2 | A | 3 | 148.1418 |  | 127 | 190.5
M2 | A | 4 | 146.8931 |  | 127 | 190.5
M2 | A | 5 | 147.1653 |  | 127 | 190.5
M2 | A | 6 | 147.7837 |  | 127 | 190.5
M2 | A | 7 | 148.905 |  | 127 | 190.5
M2 | A | 8 | 148.1634 |  | 127 | 190.5
M2 | A | 9 | 148.7451 |  | 127 | 190.5
M2 | A | 10 | 148.7287 |  | 127 | 190.5
M2 | A | 11 | 147.7726 |  | 127 | 190.5
M2 | A | 12 | 148.8393 |  | 127 | 190.5
M2 | A | 13 | 148.6837 |  | 127 | 190.5
M2 | A | 14 | 147.4838 |  | 127 | 190.5
M2 | A | 15 | 147.1971 |  | 127 | 190.5
M2 | A | 16 | 147.3903 |  | 127 | 190.5
M2 | B | 1 | 148.1654 |  | 127 | 190.5
M2 | B | 2 | 148.1206 |  | 127 | 190.5
M2 | B | 3 | 148.3546 |  | 127 | 190.5
M2 | B | 4 | 147.079 |  | 127 | 190.5
M2 | B | 5 | 147.0879 |  | 127 | 190.5
M2 | B | 6 | 147.5269 |  | 127 | 190.5
M2 | B | 7 | 148.2093 |  | 127 | 190.5
M2 | B | 8 | 147.9036 |  | 127 | 190.5
M2 | B | 9 | 148.0719 |  | 127 | 190.5
M2 | B | 10 | 148.5111 |  | 127 | 190.5
M2 | B | 11 | 147.9556 |  | 127 | 190.5
M2 | B | 12 | 148.2401 |  | 127 | 190.5
M2 | B | 13 | 148.0619 |  | 127 | 190.5
M2 | B | 14 | 147.8386 |  | 127 | 190.5
M2 | B | 15 | 147.6409 |  | 127 | 190.5
M2 | B | 16 | 147.7 |  | 127 | 190.5
M3 | A | 1 | 117.3463 |  | 107 | 128.4
M3 | A | 2 | 116.9484 |  | 107 | 128.4
M3 | A | 3 | 116.7858 |  | 107 | 128.4
M3 | A | 4 | 116.1258 |  | 107 | 128.4
M3 | A | 5 | 116.0764 |  | 107 | 128.4
M3 | A | 6 | 115.7677 |  | 107 | 128.4
M3 | A | 7 | 116.2958 |  | 107 | 128.4
M3 | A | 8 | 116.3475 |  | 107 | 128.4
M3 | A | 9 | 116.8401 |  | 107 | 128.4
M3 | A | 10 | 116.746 |  | 107 | 128.4
M3 | A | 11 | 116.6478 |  | 107 | 128.4
M3 | A | 12 | 116.1844 |  | 107 | 128.4
M3 | A | 13 | 117.4522 |  | 107 | 128.4
M3 | A | 14 | 115.9873 |  | 107 | 128.4
M3 | A | 15 | 115.5261 |  | 107 | 128.4
M3 | A | 16 | 116.0904 |  | 107 | 128.4
M3 | B | 1 | 116.9433 |  | 107 | 128.4
M3 | B | 2 | 116.5715 |  | 107 | 128.4
M3 | B | 3 | 116.6048 |  | 107 | 128.4
M3 | B | 4 | 115.8061 |  | 107 | 128.4
M3 | B | 5 | 115.7411 |  | 107 | 128.4
M3 | B | 6 | 116.2737 |  | 107 | 128.4
M3 | B | 7 | 116.179 |  | 107 | 128.4
M3 | B | 8 | 116.0264 |  | 107 | 128.4
M3 | B | 9 | 116.5618 |  | 107 | 128.4
M3 | B | 10 | 116.4129 |  | 107 | 128.4
M3 | B | 11 | 116.319 |  | 107 | 128.4
M3 | B | 12 | 116.5551 |  | 107 | 128.4
M3 | B | 13 | 117.791 |  | 107 | 128.4
M3 | B | 14 | 116.7367 |  | 107 | 128.4
M3 | B | 15 | 115.9536 |  | 107 | 128.4
M3 | B | 16 | 116.4618 |  | 107 | 128.4
M4 | A | 1 | 90.3048 |  | 78 | 93.6
M4 | A | 2 | 90.5024 |  | 78 | 93.6
M4 | A | 3 | 89.9471 |  | 78 | 93.6
M4 | A | 4 | 90.2637 |  | 78 | 93.6
M4 | A | 5 | 89.9991 |  | 78 | 93.6
M4 | A | 6 | 89.9719 |  | 78 | 93.6
M4 | A | 7 | 90.0795 |  | 78 | 93.6
M4 | A | 8 | 89.9944 |  | 78 | 93.6
M4 | A | 9 | 90.1407 |  | 78 | 93.6
M4 | A | 10 | 90.5138 |  | 78 | 93.6
M4 | A | 11 | 90.3463 |  | 78 | 93.6
M4 | A | 12 | 90.6855 |  | 78 | 93.6
M4 | A | 13 | 90.6475 |  | 78 | 93.6
M4 | A | 14 | 90.2262 |  | 78 | 93.6
M4 | A | 15 | 90.0349 |  | 78 | 93.6
M4 | A | 16 | 90.3446 |  | 78 | 93.6
M4 | B | 1 | 90.4005 |  | 78 | 93.6
M4 | B | 2 | 90.5759 |  | 78 | 93.6
M4 | B | 3 | 90.0907 |  | 78 | 93.6
M4 | B | 4 | 90.3047 |  | 78 | 93.6
M4 | B | 5 | 89.9597 |  | 78 | 93.6
M4 | B | 6 | 90.014 |  | 78 | 93.6
M4 | B | 7 | 90.1091 |  | 78 | 93.6
M4 | B | 8 | 90.0901 |  | 78 | 93.6
M4 | B | 9 | 90.2705 |  | 78 | 93.6
M4 | B | 10 | 90.3394 |  | 78 | 93.6
M4 | B | 11 | 90.4824 |  | 78 | 93.6
M4 | B | 12 | 90.4768 |  | 78 | 93.6
M4 | B | 13 | 90.6758 |  | 78 | 93.6
M4 | B | 14 | 90.123 |  | 78 | 93.6
M4 | B | 15 | 89.7852 |  | 78 | 93.6
M4 | B | 16 | 90.3099 |  | 78 | 93.6
M5 | A | 1 | 61.9898 |  | 59 | 70.8
M5 | A | 2 | 61.8889 |  | 59 | 70.8
M5 | A | 3 | 62.1242 |  | 59 | 70.8
M5 | A | 4 | 61.8153 |  | 59 | 70.8
M5 | A | 5 | 61.9363 |  | 59 | 70.8
M5 | A | 6 | 62.0164 |  | 59 | 70.8
M5 | A | 7 | 62.2793 |  | 59 | 70.8
M5 | A | 8 | 62.336 |  | 59 | 70.8
M5 | A | 9 | 61.8775 |  | 59 | 70.8
M5 | A | 10 | 61.9467 |  | 59 | 70.8
M5 | A | 11 | 61.7906 |  | 59 | 70.8
M5 | A | 12 | 61.8844 |  | 59 | 70.8
M5 | A | 13 | 62.155 |  | 59 | 70.8
M5 | A | 14 | 62.191 |  | 59 | 70.8
M5 | A | 15 | 61.8865 |  | 59 | 70.8
M5 | A | 16 | 61.6059 |  | 59 | 70.8
M5 | B | 1 | 61.9634 |  | 59 | 70.8
M5 | B | 2 | 61.755 |  | 59 | 70.8
M5 | B | 3 | 62.1373 |  | 59 | 70.8
M5 | B | 4 | 61.984 |  | 59 | 70.8
M5 | B | 5 | 62.1748 |  | 59 | 70.8
M5 | B | 6 | 62.2934 |  | 59 | 70.8
M5 | B | 7 | 62.2205 |  | 59 | 70.8
M5 | B | 8 | 62.5203 |  | 59 | 70.8
M5 | B | 9 | 62.0743 |  | 59 | 70.8
M5 | B | 10 | 62.0276 |  | 59 | 70.8
M5 | B | 11 | 61.624 |  | 59 | 70.8
M5 | B | 12 | 61.5544 |  | 59 | 70.8
M5 | B | 13 | 61.863 |  | 59 | 70.8
M5 | B | 14 | 62.2453 |  | 59 | 70.8
M5 | B | 15 | 61.936 |  | 59 | 70.8
M5 | B | 16 | 61.5659 |  | 59 | 70.8
M7 | A | 1 | 30.8233 |  | 29 | 34.8
M7 | A | 2 | 30.7322 |  | 29 | 34.8
M7 | A | 3 | 30.8212 |  | 29 | 34.8
M7 | A | 4 | 30.7874 |  | 29 | 34.8
M7 | A | 5 | 30.6649 |  | 29 | 34.8
M7 | A | 6 | 30.7324 |  | 29 | 34.8
M7 | A | 7 | 30.8293 |  | 29 | 34.8
M7 | A | 8 | 30.9035 |  | 29 | 34.8
M7 | A | 9 | 30.6752 |  | 29 | 34.8
M7 | A | 10 | 30.7429 |  | 29 | 34.8
M7 | A | 11 | 30.6945 |  | 29 | 34.8
M7 | A | 12 | 30.5759 |  | 29 | 34.8
M7 | A | 13 | 30.6931 |  | 29 | 34.8
M7 | A | 14 | 30.8602 |  | 29 | 34.8
M7 | A | 15 | 30.6631 |  | 29 | 34.8
M7 | A | 16 | 30.7846 |  | 29 | 34.8
M7 | B | 1 | 30.7979 |  | 29 | 34.8
M7 | B | 2 | 30.7116 |  | 29 | 34.8
M7 | B | 3 | 30.7985 |  | 29 | 34.8
M7 | B | 4 | 30.7789 |  | 29 | 34.8
M7 | B | 5 | 30.6517 |  | 29 | 34.8
M7 | B | 6 | 30.7144 |  | 29 | 34.8
M7 | B | 7 | 30.8046 |  | 29 | 34.8
M7 | B | 8 | 30.903 |  | 29 | 34.8
M7 | B | 9 | 30.6706 |  | 29 | 34.8
M7 | B | 10 | 30.7403 |  | 29 | 34.8
M7 | B | 11 | 30.6668 |  | 29 | 34.8
M7 | B | 12 | 30.5538 |  | 29 | 34.8
M7 | B | 13 | 30.6869 |  | 29 | 34.8
M7 | B | 14 | 30.8325 |  | 29 | 34.8
M7 | B | 15 | 30.654 |  | 29 | 34.8
M7 | B | 16 | 30.7641 |  | 29 | 34.8

## DNB gain transition
Band | HAM | Transition | Agg Zone | Band Median |  | Band Average
|  |  |  | Eside A | Eside B | Eside A | Eside B
|  |  |  | Ltrans | Ltrans | Ltrans | Ltrans
DNB | A | MGS->LGS | 1 | 2.6208e-05 | #NUM! | 2.6276e-05 | #DIV/0!
DNB | A | MGS->LGS | 7 | 4.0256e-05 | #NUM! | 4.0467e-05 | #DIV/0!
DNB | A | MGS->LGS | 11 | 5.9353e-05 | #NUM! | 5.9572e-05 | #DIV/0!
DNB | A | MGS->LGS | 15 | 6.2722e-05 | #NUM! | 6.2907e-05 | #DIV/0!
DNB | A | MGS->LGS | 20 | 0.00013104 | #NUM! | 0.00011737 | #DIV/0!
DNB | A | MGS->LGS | 26 | 0.00020444 | #NUM! | 0.00020491 | #DIV/0!
DNB | A | MGS->LGS | 32 | 0.00033092 | #NUM! | 0.0003107 | #DIV/0!
DNB | A | MGS->LGS | 34 | 1.355e-05 | #NUM! | 1.3886e-05 | #DIV/0!
DNB | B | MGS->LGS | 1 | 2.6006e-05 | #NUM! | 2.6053e-05 | #DIV/0!
DNB | B | MGS->LGS | 7 | 4.0229e-05 | #NUM! | 4.0386e-05 | #DIV/0!
DNB | B | MGS->LGS | 11 | 5.9293e-05 | #NUM! | 5.943e-05 | #DIV/0!
DNB | B | MGS->LGS | 15 | 6.2521e-05 | #NUM! | 6.6342e-05 | #DIV/0!
DNB | B | MGS->LGS | 20 | 0.00012656 | #NUM! | 0.0001134 | #DIV/0!
DNB | B | MGS->LGS | 26 | 0.00020548 | #NUM! | 0.00020547 | #DIV/0!
DNB | B | MGS->LGS | 32 | 0.00033036 | #NUM! | 0.00031281 | #DIV/0!
DNB | B | MGS->LGS | 34 | 1.3506e-05 | #NUM! | 1.3857e-05 | #DIV/0!
DNB | A | HGS->MGS | 1 | 1.8271e-07 | #NUM! | 1.826e-07 | #DIV/0!
DNB | A | HGS->MGS | 7 | 2.9661e-07 | #NUM! | 2.9686e-07 | #DIV/0!
DNB | A | HGS->MGS | 11 | 4.4178e-07 | #NUM! | 4.4151e-07 | #DIV/0!
DNB | A | HGS->MGS | 15 | 7.0161e-07 | #NUM! | 7.0194e-07 | #DIV/0!
DNB | A | HGS->MGS | 20 | 1.0594e-06 | #NUM! | 1.0599e-06 | #DIV/0!
DNB | A | HGS->MGS | 26 | 1.5656e-06 | #NUM! | 1.5726e-06 | #DIV/0!
DNB | A | HGS->MGS | 32 | 2.5024e-06 | #NUM! | 2.5095e-06 | #DIV/0!
DNB | A | HGS->MGS | 34 | 9.9851e-08 | #NUM! | 9.9903e-08 | #DIV/0!
DNB | B | HGS->MGS | 1 | 1.8254e-07 | #NUM! | 1.8248e-07 | #DIV/0!
DNB | B | HGS->MGS | 7 | 2.9645e-07 | #NUM! | 2.9652e-07 | #DIV/0!
DNB | B | HGS->MGS | 11 | 4.4111e-07 | #NUM! | 4.4115e-07 | #DIV/0!
DNB | B | HGS->MGS | 15 | 7.0075e-07 | #NUM! | 7.0122e-07 | #DIV/0!
DNB | B | HGS->MGS | 20 | 1.0573e-06 | #NUM! | 1.0583e-06 | #DIV/0!
DNB | B | HGS->MGS | 26 | 1.5693e-06 | #NUM! | 1.5748e-06 | #DIV/0!
DNB | B | HGS->MGS | 32 | 2.4851e-06 | #NUM! | 2.441e-06 | #DIV/0!
DNB | B | HGS->MGS | 34 | 9.978e-08 | #NUM! | 9.9828e-08 | #DIV/0!
Band | HAM | Transition | Agg Zone | Detector | Detector Dependent |  |  | Notes:
[5 empty] | Eside A | Eside B |  | 1. This table provides VIIRS F2 DNB gain transitions.
[5 empty] | Ltrans | Ltrans |  | 2. The results were derived from RC-1 using the SIS-100 source.
DNB | A | MGS->LGS | 1 | 1 | 2.7417e-05 |  |  | 3. Transition was determined from partial views of the SIS-100 source.
DNB | A | MGS->LGS | 1 | 2 | 2.7221e-05 |  |  | 4. Ltrans values are in radiance, W cm-2 sr-1
DNB | A | MGS->LGS | 1 | 3 | 2.6853e-05 |  |  | 5. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
DNB | A | MGS->LGS | 1 | 4 | 2.6536e-05 |  |  | 6. The test was set to diagnostic mode.
DNB | A | MGS->LGS | 1 | 5 | 2.765e-05 |  |  | 7. Operational mode offset tables were used during B-side collection. This resulted in some saturated pixels being reported and the wrong gain state reporting, so ambient transition points could not be calculated
DNB | A | MGS->LGS | 1 | 6 | 2.7354e-05
DNB | A | MGS->LGS | 1 | 7 | 2.7118e-05
DNB | A | MGS->LGS | 1 | 8 | 2.691e-05
DNB | A | MGS->LGS | 1 | 9 | 2.5811e-05
DNB | A | MGS->LGS | 1 | 10 | 2.4745e-05
DNB | A | MGS->LGS | 1 | 11 | 2.588e-05
DNB | A | MGS->LGS | 1 | 12 | 2.4938e-05
DNB | A | MGS->LGS | 1 | 13 | 2.5554e-05
DNB | A | MGS->LGS | 1 | 14 | 2.4872e-05
DNB | A | MGS->LGS | 1 | 15 | 2.5744e-05
DNB | A | MGS->LGS | 1 | 16 | 2.5816e-05
DNB | A | MGS->LGS | 1 | 1 | 4.2202e-05
DNB | A | MGS->LGS | 1 | 2 | 4.1655e-05
DNB | A | MGS->LGS | 1 | 3 | 4.2223e-05
DNB | A | MGS->LGS | 1 | 4 | 4.2014e-05
DNB | A | MGS->LGS | 1 | 5 | 4.1527e-05
DNB | A | MGS->LGS | 1 | 6 | 4.2098e-05
DNB | A | MGS->LGS | 1 | 7 | 4.092e-05
DNB | A | MGS->LGS | 1 | 8 | 4.1403e-05
DNB | A | MGS->LGS | 1 | 9 | 3.9114e-05
DNB | A | MGS->LGS | 1 | 10 | 3.8718e-05
DNB | A | MGS->LGS | 1 | 11 | 3.9453e-05
DNB | A | MGS->LGS | 1 | 12 | 3.898e-05
DNB | A | MGS->LGS | 1 | 13 | 3.9592e-05
DNB | A | MGS->LGS | 1 | 14 | 3.9139e-05
DNB | A | MGS->LGS | 1 | 15 | 3.9504e-05
DNB | A | MGS->LGS | 1 | 16 | 3.8928e-05
DNB | A | MGS->LGS | 7 | 1 | 6.1334e-05
DNB | A | MGS->LGS | 7 | 2 | 6.1755e-05
DNB | A | MGS->LGS | 7 | 3 | 6.1643e-05
DNB | A | MGS->LGS | 7 | 4 | 6.1444e-05
DNB | A | MGS->LGS | 7 | 5 | 6.1844e-05
DNB | A | MGS->LGS | 7 | 6 | 6.155e-05
DNB | A | MGS->LGS | 7 | 7 | 6.1358e-05
DNB | A | MGS->LGS | 7 | 8 | 6.0651e-05
DNB | A | MGS->LGS | 7 | 9 | 5.7592e-05
DNB | A | MGS->LGS | 7 | 10 | 5.7268e-05
DNB | A | MGS->LGS | 7 | 11 | 5.7769e-05
DNB | A | MGS->LGS | 7 | 12 | 5.7413e-05
DNB | A | MGS->LGS | 7 | 13 | 5.8042e-05
DNB | A | MGS->LGS | 7 | 14 | 5.759e-05
DNB | A | MGS->LGS | 7 | 15 | 5.7844e-05
DNB | A | MGS->LGS | 7 | 16 | 5.8055e-05
DNB | A | MGS->LGS | 7 | 1 | 5.8644e-05
DNB | A | MGS->LGS | 7 | 2 | 5.9195e-05
DNB | A | MGS->LGS | 7 | 3 | 5.7606e-05
DNB | A | MGS->LGS | 7 | 4 | 5.8327e-05
DNB | A | MGS->LGS | 7 | 5 | 5.9881e-05
DNB | A | MGS->LGS | 7 | 6 | 5.9671e-05
DNB | A | MGS->LGS | 7 | 7 | 6.0348e-05
DNB | A | MGS->LGS | 7 | 8 | 6.173e-05
DNB | A | MGS->LGS | 7 | 9 | 6.3713e-05
DNB | A | MGS->LGS | 7 | 10 | 6.4116e-05
DNB | A | MGS->LGS | 7 | 11 | 6.3775e-05
DNB | A | MGS->LGS | 7 | 12 | 6.5127e-05
DNB | A | MGS->LGS | 7 | 13 | 6.7475e-05
DNB | A | MGS->LGS | 7 | 14 | 6.8028e-05
DNB | A | MGS->LGS | 7 | 15 | 6.9827e-05
DNB | A | MGS->LGS | 7 | 16 | 6.9046e-05
DNB | A | MGS->LGS | 11 | 1 | 0.00013512
DNB | A | MGS->LGS | 11 | 2 | 0.00013527
DNB | A | MGS->LGS | 11 | 3 | 0.00013639
DNB | A | MGS->LGS | 11 | 4 | 0.00013736
DNB | A | MGS->LGS | 11 | 5 | 0.00013734
DNB | A | MGS->LGS | 11 | 6 | 0.00013736
DNB | A | MGS->LGS | 11 | 7 | 0.00013789
DNB | A | MGS->LGS | 11 | 8 | 0.00013911
DNB | A | MGS->LGS | 11 | 9 | 8.6768e-05
DNB | A | MGS->LGS | 11 | 10 | 8.7858e-05
DNB | A | MGS->LGS | 11 | 11 | 8.63e-05
DNB | A | MGS->LGS | 11 | 12 | 8.6956e-05
DNB | A | MGS->LGS | 11 | 13 | 8.8995e-05
DNB | A | MGS->LGS | 11 | 14 | 9.2225e-05
DNB | A | MGS->LGS | 11 | 15 | 0.00012695
DNB | A | MGS->LGS | 11 | 16 | 0.00012596
DNB | A | MGS->LGS | 11 | 1 | 0.00021754
DNB | A | MGS->LGS | 11 | 2 | 0.00021644
DNB | A | MGS->LGS | 11 | 3 | 0.00021902
DNB | A | MGS->LGS | 11 | 4 | 0.00021967
DNB | A | MGS->LGS | 11 | 5 | 0.00021946
DNB | A | MGS->LGS | 11 | 6 | 0.00022036
DNB | A | MGS->LGS | 11 | 7 | 0.00022172
DNB | A | MGS->LGS | 11 | 8 | 0.00022338
DNB | A | MGS->LGS | 11 | 9 | 0.00019244
DNB | A | MGS->LGS | 11 | 10 | 0.00019161
DNB | A | MGS->LGS | 11 | 11 | 0.00019233
DNB | A | MGS->LGS | 11 | 12 | 0.0001915
DNB | A | MGS->LGS | 11 | 13 | 0.00019095
DNB | A | MGS->LGS | 11 | 14 | 0.00018863
DNB | A | MGS->LGS | 11 | 15 | 0.00018572
DNB | A | MGS->LGS | 11 | 16 | 0.00018782
DNB | A | MGS->LGS | 15 | 1 | 0.00032368
DNB | A | MGS->LGS | 15 | 2 | 0.00032675
DNB | A | MGS->LGS | 15 | 3 | 0.00032895
DNB | A | MGS->LGS | 15 | 4 | 3.3219e-05
DNB | A | MGS->LGS | 15 | 5 | 0.00033041
DNB | A | MGS->LGS | 15 | 6 | 0.00033553
DNB | A | MGS->LGS | 15 | 7 | 0.00033691
DNB | A | MGS->LGS | 15 | 8 | 0.0003399
DNB | A | MGS->LGS | 15 | 9 | 0.00033143
DNB | A | MGS->LGS | 15 | 10 | 0.00029663
DNB | A | MGS->LGS | 15 | 11 | 0.0003324
DNB | A | MGS->LGS | 15 | 12 | 0.00032777
DNB | A | MGS->LGS | 15 | 13 | 0.00032861
DNB | A | MGS->LGS | 15 | 14 | 0.00033172
DNB | A | MGS->LGS | 15 | 15 | 0.00033329
DNB | A | MGS->LGS | 15 | 16 | 0.00033405
DNB | A | MGS->LGS | 15 | 1 | 1.4333e-05
DNB | A | MGS->LGS | 15 | 2 | 1.4509e-05
DNB | A | MGS->LGS | 15 | 3 | 1.4487e-05
DNB | A | MGS->LGS | 15 | 4 | 1.3555e-05
DNB | A | MGS->LGS | 15 | 5 | 1.4445e-05
DNB | A | MGS->LGS | 15 | 6 | 1.4394e-05
DNB | A | MGS->LGS | 15 | 7 | 1.431e-05
DNB | A | MGS->LGS | 15 | 8 | 1.4202e-05
DNB | A | MGS->LGS | 15 | 9 | 1.3467e-05
DNB | A | MGS->LGS | 15 | 10 | 1.3477e-05
DNB | A | MGS->LGS | 15 | 11 | 1.3545e-05
DNB | A | MGS->LGS | 15 | 12 | 1.3488e-05
DNB | A | MGS->LGS | 15 | 13 | 1.3497e-05
DNB | A | MGS->LGS | 15 | 14 | 1.3508e-05
DNB | A | MGS->LGS | 15 | 15 | 1.3476e-05
DNB | A | MGS->LGS | 15 | 16 | 1.3482e-05
DNB | B | MGS->LGS | 20 | 1 | 2.6271e-05
DNB | B | MGS->LGS | 20 | 2 | 2.7199e-05
DNB | B | MGS->LGS | 20 | 3 | 2.6829e-05
DNB | B | MGS->LGS | 20 | 4 | 2.6542e-05
DNB | B | MGS->LGS | 20 | 5 | 2.6395e-05
DNB | B | MGS->LGS | 20 | 6 | 2.7359e-05
DNB | B | MGS->LGS | 20 | 7 | 2.7104e-05
DNB | B | MGS->LGS | 20 | 8 | 2.6543e-05
DNB | B | MGS->LGS | 20 | 9 | 2.5742e-05
DNB | B | MGS->LGS | 20 | 10 | 2.4744e-05
DNB | B | MGS->LGS | 20 | 11 | 2.5508e-05
DNB | B | MGS->LGS | 20 | 12 | 2.4918e-05
DNB | B | MGS->LGS | 20 | 13 | 2.5547e-05
DNB | B | MGS->LGS | 20 | 14 | 2.5674e-05
DNB | B | MGS->LGS | 20 | 15 | 2.5547e-05
DNB | B | MGS->LGS | 20 | 16 | 2.4928e-05
DNB | B | MGS->LGS | 20 | 1 | 4.1401e-05
DNB | B | MGS->LGS | 20 | 2 | 4.1604e-05
DNB | B | MGS->LGS | 20 | 3 | 4.2227e-05
DNB | B | MGS->LGS | 20 | 4 | 4.1971e-05
DNB | B | MGS->LGS | 20 | 5 | 4.1498e-05
DNB | B | MGS->LGS | 20 | 6 | 4.2006e-05
DNB | B | MGS->LGS | 20 | 7 | 4.0908e-05
DNB | B | MGS->LGS | 20 | 8 | 4.1341e-05
DNB | B | MGS->LGS | 20 | 9 | 3.9108e-05
DNB | B | MGS->LGS | 20 | 10 | 3.8712e-05
DNB | B | MGS->LGS | 20 | 11 | 3.938e-05
DNB | B | MGS->LGS | 20 | 12 | 3.8976e-05
DNB | B | MGS->LGS | 20 | 13 | 3.9549e-05
DNB | B | MGS->LGS | 20 | 14 | 3.9162e-05
DNB | B | MGS->LGS | 20 | 15 | 3.9413e-05
DNB | B | MGS->LGS | 20 | 16 | 3.8919e-05
DNB | B | MGS->LGS | 26 | 1 | 6.1245e-05
DNB | B | MGS->LGS | 26 | 2 | 6.1726e-05
DNB | B | MGS->LGS | 26 | 3 | 6.1568e-05
DNB | B | MGS->LGS | 26 | 4 | 6.146e-05
DNB | B | MGS->LGS | 26 | 5 | 6.1718e-05
DNB | B | MGS->LGS | 26 | 6 | 6.1308e-05
DNB | B | MGS->LGS | 26 | 7 | 6.1286e-05
DNB | B | MGS->LGS | 26 | 8 | 6.0712e-05
DNB | B | MGS->LGS | 26 | 9 | 5.7534e-05
DNB | B | MGS->LGS | 26 | 10 | 5.7255e-05
DNB | B | MGS->LGS | 26 | 11 | 5.7035e-05
DNB | B | MGS->LGS | 26 | 12 | 5.7485e-05
DNB | B | MGS->LGS | 26 | 13 | 5.7875e-05
DNB | B | MGS->LGS | 26 | 14 | 5.762e-05
DNB | B | MGS->LGS | 26 | 15 | 5.7761e-05
DNB | B | MGS->LGS | 26 | 16 | 5.7295e-05
DNB | B | MGS->LGS | 26 | 1 | 5.8871e-05
DNB | B | MGS->LGS | 26 | 2 | 5.9472e-05
DNB | B | MGS->LGS | 26 | 3 | 5.7595e-05
DNB | B | MGS->LGS | 26 | 4 | 5.8369e-05
DNB | B | MGS->LGS | 26 | 5 | 5.8808e-05
DNB | B | MGS->LGS | 26 | 6 | 5.8573e-05
DNB | B | MGS->LGS | 26 | 7 | 5.9224e-05
DNB | B | MGS->LGS | 26 | 8 | 6.0737e-05
DNB | B | MGS->LGS | 26 | 9 | 6.2521e-05
DNB | B | MGS->LGS | 26 | 10 | 6.2862e-05
DNB | B | MGS->LGS | 26 | 11 | 6.2577e-05
DNB | B | MGS->LGS | 26 | 12 | 6.4039e-05
DNB | B | MGS->LGS | 26 | 13 | 6.5997e-05
DNB | B | MGS->LGS | 26 | 14 | 6.6825e-05
DNB | B | MGS->LGS | 26 | 15 | 6.8438e-05
DNB | B | MGS->LGS | 26 | 16 | 6.7721e-05
DNB | B | MGS->LGS | 32 | 1 | 0.00013518
DNB | B | MGS->LGS | 32 | 2 | 0.00013436
DNB | B | MGS->LGS | 32 | 3 | 0.00013591
DNB | B | MGS->LGS | 32 | 4 | 0.00013734
DNB | B | MGS->LGS | 32 | 5 | 0.00013737
DNB | B | MGS->LGS | 32 | 6 | 0.00013622
DNB | B | MGS->LGS | 32 | 7 | 0.00013774
DNB | B | MGS->LGS | 32 | 8 | 0.00013898
DNB | B | MGS->LGS | 32 | 9 | 8.5935e-05
DNB | B | MGS->LGS | 32 | 10 | 8.7015e-05
DNB | B | MGS->LGS | 32 | 11 | 8.5457e-05
DNB | B | MGS->LGS | 32 | 12 | 8.6234e-05
DNB | B | MGS->LGS | 32 | 13 | 8.8354e-05
DNB | B | MGS->LGS | 32 | 14 | 9.1423e-05
DNB | B | MGS->LGS | 32 | 15 | 9.2065e-05
DNB | B | MGS->LGS | 32 | 16 | 0.00012656
DNB | B | MGS->LGS | 32 | 1 | 0.00021742
DNB | B | MGS->LGS | 32 | 2 | 0.00021684
DNB | B | MGS->LGS | 32 | 3 | 0.00021857
DNB | B | MGS->LGS | 32 | 4 | 0.00021998
DNB | B | MGS->LGS | 32 | 5 | 0.00021912
DNB | B | MGS->LGS | 32 | 6 | 0.00022016
DNB | B | MGS->LGS | 32 | 7 | 0.00021983
DNB | B | MGS->LGS | 32 | 8 | 0.00022264
DNB | B | MGS->LGS | 32 | 9 | 0.00019405
DNB | B | MGS->LGS | 32 | 10 | 0.00019408
DNB | B | MGS->LGS | 32 | 11 | 0.00019412
DNB | B | MGS->LGS | 32 | 12 | 0.00019198
DNB | B | MGS->LGS | 32 | 13 | 0.00019139
DNB | B | MGS->LGS | 32 | 14 | 0.00019094
DNB | B | MGS->LGS | 32 | 15 | 0.00018791
DNB | B | MGS->LGS | 32 | 16 | 0.00018854
DNB | B | MGS->LGS | 34 | 1 | 0.00032295
DNB | B | MGS->LGS | 34 | 2 | 0.00032665
DNB | B | MGS->LGS | 34 | 3 | 0.00032824
DNB | B | MGS->LGS | 34 | 4 | 4.3672e-05
DNB | B | MGS->LGS | 34 | 5 | 0.00032997
DNB | B | MGS->LGS | 34 | 6 | 0.00033461
DNB | B | MGS->LGS | 34 | 7 | 0.00033603
DNB | B | MGS->LGS | 34 | 8 | 0.00033872
DNB | B | MGS->LGS | 34 | 9 | 0.00032974
DNB | B | MGS->LGS | 34 | 10 | 0.0003318
DNB | B | MGS->LGS | 34 | 11 | 0.00033074
DNB | B | MGS->LGS | 34 | 12 | 0.00032702
DNB | B | MGS->LGS | 34 | 13 | 0.00032826
DNB | B | MGS->LGS | 34 | 14 | 0.00033101
DNB | B | MGS->LGS | 34 | 15 | 0.00033226
DNB | B | MGS->LGS | 34 | 16 | 0.00033334
DNB | B | MGS->LGS | 34 | 1 | 1.4312e-05
DNB | B | MGS->LGS | 34 | 2 | 1.4478e-05
DNB | B | MGS->LGS | 34 | 3 | 1.4463e-05
DNB | B | MGS->LGS | 34 | 4 | 1.3463e-05
DNB | B | MGS->LGS | 34 | 5 | 1.4419e-05
DNB | B | MGS->LGS | 34 | 6 | 1.4367e-05
DNB | B | MGS->LGS | 34 | 7 | 1.4286e-05
DNB | B | MGS->LGS | 34 | 8 | 1.4176e-05
DNB | B | MGS->LGS | 34 | 9 | 1.3445e-05
DNB | B | MGS->LGS | 34 | 10 | 1.3451e-05
DNB | B | MGS->LGS | 34 | 11 | 1.3525e-05
DNB | B | MGS->LGS | 34 | 12 | 1.3459e-05
DNB | B | MGS->LGS | 34 | 13 | 1.3475e-05
DNB | B | MGS->LGS | 34 | 14 | 1.3486e-05
DNB | B | MGS->LGS | 34 | 15 | 1.3452e-05
DNB | B | MGS->LGS | 34 | 16 | 1.3457e-05
DNB | A | HGS->MGS | 1 | 1 | 1.8192e-07
DNB | A | HGS->MGS | 1 | 2 | 1.8423e-07
DNB | A | HGS->MGS | 1 | 3 | 1.8336e-07
DNB | A | HGS->MGS | 1 | 4 | 1.8369e-07
DNB | A | HGS->MGS | 1 | 5 | 1.8385e-07
DNB | A | HGS->MGS | 1 | 6 | 1.8329e-07
DNB | A | HGS->MGS | 1 | 7 | 1.8271e-07
DNB | A | HGS->MGS | 1 | 8 | 1.8272e-07
DNB | A | HGS->MGS | 1 | 9 | 1.8261e-07
DNB | A | HGS->MGS | 1 | 10 | 1.8093e-07
DNB | A | HGS->MGS | 1 | 11 | 1.8118e-07
DNB | A | HGS->MGS | 1 | 12 | 1.8232e-07
DNB | A | HGS->MGS | 1 | 13 | 1.8136e-07
DNB | A | HGS->MGS | 1 | 14 | 1.8303e-07
DNB | A | HGS->MGS | 1 | 15 | 1.8275e-07
DNB | A | HGS->MGS | 1 | 16 | 1.8159e-07
DNB | A | HGS->MGS | 1 | 1 | 2.9943e-07
DNB | A | HGS->MGS | 1 | 2 | 2.9813e-07
DNB | A | HGS->MGS | 1 | 3 | 2.993e-07
DNB | A | HGS->MGS | 1 | 4 | 2.9807e-07
DNB | A | HGS->MGS | 1 | 5 | 2.9763e-07
DNB | A | HGS->MGS | 1 | 6 | 2.9645e-07
DNB | A | HGS->MGS | 1 | 7 | 2.9716e-07
DNB | A | HGS->MGS | 1 | 8 | 2.9623e-07
DNB | A | HGS->MGS | 1 | 9 | 2.9571e-07
DNB | A | HGS->MGS | 1 | 10 | 2.9516e-07
DNB | A | HGS->MGS | 1 | 11 | 2.9523e-07
DNB | A | HGS->MGS | 1 | 12 | 2.9541e-07
DNB | A | HGS->MGS | 1 | 13 | 2.9481e-07
DNB | A | HGS->MGS | 1 | 14 | 2.967e-07
DNB | A | HGS->MGS | 1 | 15 | 2.9651e-07
DNB | A | HGS->MGS | 1 | 16 | 2.978e-07
DNB | A | HGS->MGS | 7 | 1 | 4.4858e-07
DNB | A | HGS->MGS | 7 | 2 | 4.452e-07
DNB | A | HGS->MGS | 7 | 3 | 4.4426e-07
DNB | A | HGS->MGS | 7 | 4 | 4.4361e-07
DNB | A | HGS->MGS | 7 | 5 | 4.4275e-07
DNB | A | HGS->MGS | 7 | 6 | 4.388e-07
DNB | A | HGS->MGS | 7 | 7 | 4.4327e-07
DNB | A | HGS->MGS | 7 | 8 | 4.3823e-07
DNB | A | HGS->MGS | 7 | 9 | 4.3989e-07
DNB | A | HGS->MGS | 7 | 10 | 4.4192e-07
DNB | A | HGS->MGS | 7 | 11 | 4.3634e-07
DNB | A | HGS->MGS | 7 | 12 | 4.3971e-07
DNB | A | HGS->MGS | 7 | 13 | 4.4253e-07
DNB | A | HGS->MGS | 7 | 14 | 4.4077e-07
DNB | A | HGS->MGS | 7 | 15 | 4.3658e-07
DNB | A | HGS->MGS | 7 | 16 | 4.4165e-07
DNB | A | HGS->MGS | 7 | 1 | 7.0972e-07
DNB | A | HGS->MGS | 7 | 2 | 7.0621e-07
DNB | A | HGS->MGS | 7 | 3 | 7.0429e-07
DNB | A | HGS->MGS | 7 | 4 | 7.0638e-07
DNB | A | HGS->MGS | 7 | 5 | 7.0498e-07
DNB | A | HGS->MGS | 7 | 6 | 7.0337e-07
DNB | A | HGS->MGS | 7 | 7 | 7.0327e-07
DNB | A | HGS->MGS | 7 | 8 | 6.9893e-07
DNB | A | HGS->MGS | 7 | 9 | 6.9597e-07
DNB | A | HGS->MGS | 7 | 10 | 7.0072e-07
DNB | A | HGS->MGS | 7 | 11 | 7.0044e-07
DNB | A | HGS->MGS | 7 | 12 | 7.0025e-07
DNB | A | HGS->MGS | 7 | 13 | 7.0062e-07
DNB | A | HGS->MGS | 7 | 14 | 6.9612e-07
DNB | A | HGS->MGS | 7 | 15 | 6.9729e-07
DNB | A | HGS->MGS | 7 | 16 | 7.025e-07
DNB | A | HGS->MGS | 11 | 1 | 1.054e-06
DNB | A | HGS->MGS | 11 | 2 | 1.0698e-06
DNB | A | HGS->MGS | 11 | 3 | 1.0686e-06
DNB | A | HGS->MGS | 11 | 4 | 1.0661e-06
DNB | A | HGS->MGS | 11 | 5 | 1.062e-06
DNB | A | HGS->MGS | 11 | 6 | 1.0611e-06
DNB | A | HGS->MGS | 11 | 7 | 1.0578e-06
DNB | A | HGS->MGS | 11 | 8 | 1.0596e-06
DNB | A | HGS->MGS | 11 | 9 | 1.0556e-06
DNB | A | HGS->MGS | 11 | 10 | 1.0542e-06
DNB | A | HGS->MGS | 11 | 11 | 1.0561e-06
DNB | A | HGS->MGS | 11 | 12 | 1.0565e-06
DNB | A | HGS->MGS | 11 | 13 | 1.0598e-06
DNB | A | HGS->MGS | 11 | 14 | 1.0594e-06
DNB | A | HGS->MGS | 11 | 15 | 1.0594e-06
DNB | A | HGS->MGS | 11 | 16 | 1.0592e-06
DNB | A | HGS->MGS | 11 | 1 | 1.6852e-06
DNB | A | HGS->MGS | 11 | 2 | 1.5402e-06
DNB | A | HGS->MGS | 11 | 3 | 1.5929e-06
DNB | A | HGS->MGS | 11 | 4 | 1.5978e-06
DNB | A | HGS->MGS | 11 | 5 | 1.5843e-06
DNB | A | HGS->MGS | 11 | 6 | 1.5798e-06
DNB | A | HGS->MGS | 11 | 7 | 1.5828e-06
DNB | A | HGS->MGS | 11 | 8 | 1.5711e-06
DNB | A | HGS->MGS | 11 | 9 | 1.5581e-06
DNB | A | HGS->MGS | 11 | 10 | 1.5446e-06
DNB | A | HGS->MGS | 11 | 11 | 1.5394e-06
DNB | A | HGS->MGS | 11 | 12 | 1.5452e-06
DNB | A | HGS->MGS | 11 | 13 | 1.5511e-06
DNB | A | HGS->MGS | 11 | 14 | 1.558e-06
DNB | A | HGS->MGS | 11 | 15 | 1.566e-06
DNB | A | HGS->MGS | 11 | 16 | 1.5651e-06
DNB | A | HGS->MGS | 15 | 1 | 2.3953e-06
DNB | A | HGS->MGS | 15 | 2 | 2.4128e-06
DNB | A | HGS->MGS | 15 | 3 | 2.6504e-06
DNB | A | HGS->MGS | 15 | 4 | 2.4988e-06
DNB | A | HGS->MGS | 15 | 5 | 2.4505e-06
DNB | A | HGS->MGS | 15 | 6 | 2.7514e-06
DNB | A | HGS->MGS | 15 | 7 | 2.4643e-06
DNB | A | HGS->MGS | 15 | 8 | 2.4759e-06
DNB | A | HGS->MGS | 15 | 9 | 2.5146e-06
DNB | A | HGS->MGS | 15 | 10 | 2.5061e-06
DNB | A | HGS->MGS | 15 | 11 | 2.5174e-06
DNB | A | HGS->MGS | 15 | 12 | 2.581e-06
DNB | A | HGS->MGS | 15 | 13 | 2.3556e-06
DNB | A | HGS->MGS | 15 | 14 | 2.5688e-06
DNB | A | HGS->MGS | 15 | 15 | 2.6478e-06
DNB | A | HGS->MGS | 15 | 16 | 2.3614e-06
DNB | A | HGS->MGS | 15 | 1 | 9.535e-08
DNB | A | HGS->MGS | 15 | 2 | 1.0879e-07
DNB | A | HGS->MGS | 15 | 3 | 1.0149e-07
DNB | A | HGS->MGS | 15 | 4 | 9.5375e-08
DNB | A | HGS->MGS | 15 | 5 | 9.9551e-08
DNB | A | HGS->MGS | 15 | 6 | 1.0397e-07
DNB | A | HGS->MGS | 15 | 7 | 9.8686e-08
DNB | A | HGS->MGS | 15 | 8 | 9.9775e-08
DNB | A | HGS->MGS | 15 | 9 | 1.0617e-07
DNB | A | HGS->MGS | 15 | 10 | 9.9926e-08
DNB | A | HGS->MGS | 15 | 11 | 1.0186e-07
DNB | A | HGS->MGS | 15 | 12 | 9.8057e-08
DNB | A | HGS->MGS | 15 | 13 | 9.4085e-08
DNB | A | HGS->MGS | 15 | 14 | 1.0064e-07
DNB | A | HGS->MGS | 15 | 15 | 1.0004e-07
DNB | A | HGS->MGS | 15 | 16 | 9.4674e-08
DNB | B | HGS->MGS | 20 | 1 | 1.822e-07
DNB | B | HGS->MGS | 20 | 2 | 1.8408e-07
DNB | B | HGS->MGS | 20 | 3 | 1.8273e-07
DNB | B | HGS->MGS | 20 | 4 | 1.8237e-07
DNB | B | HGS->MGS | 20 | 5 | 1.8399e-07
DNB | B | HGS->MGS | 20 | 6 | 1.8318e-07
DNB | B | HGS->MGS | 20 | 7 | 1.8253e-07
DNB | B | HGS->MGS | 20 | 8 | 1.8325e-07
DNB | B | HGS->MGS | 20 | 9 | 1.8255e-07
DNB | B | HGS->MGS | 20 | 10 | 1.809e-07
DNB | B | HGS->MGS | 20 | 11 | 1.8185e-07
DNB | B | HGS->MGS | 20 | 12 | 1.8212e-07
DNB | B | HGS->MGS | 20 | 13 | 1.812e-07
DNB | B | HGS->MGS | 20 | 14 | 1.8257e-07
DNB | B | HGS->MGS | 20 | 15 | 1.8257e-07
DNB | B | HGS->MGS | 20 | 16 | 1.8165e-07
DNB | B | HGS->MGS | 20 | 1 | 2.9741e-07
DNB | B | HGS->MGS | 20 | 2 | 2.9747e-07
DNB | B | HGS->MGS | 20 | 3 | 2.9946e-07
DNB | B | HGS->MGS | 20 | 4 | 2.9807e-07
DNB | B | HGS->MGS | 20 | 5 | 2.978e-07
DNB | B | HGS->MGS | 20 | 6 | 2.9662e-07
DNB | B | HGS->MGS | 20 | 7 | 2.9633e-07
DNB | B | HGS->MGS | 20 | 8 | 2.9638e-07
DNB | B | HGS->MGS | 20 | 9 | 2.9596e-07
DNB | B | HGS->MGS | 20 | 10 | 2.9484e-07
DNB | B | HGS->MGS | 20 | 11 | 2.9468e-07
DNB | B | HGS->MGS | 20 | 12 | 2.954e-07
DNB | B | HGS->MGS | 20 | 13 | 2.9463e-07
DNB | B | HGS->MGS | 20 | 14 | 2.9653e-07
DNB | B | HGS->MGS | 20 | 15 | 2.9538e-07
DNB | B | HGS->MGS | 20 | 16 | 2.9734e-07
DNB | B | HGS->MGS | 26 | 1 | 4.4686e-07
DNB | B | HGS->MGS | 26 | 2 | 4.435e-07
DNB | B | HGS->MGS | 26 | 3 | 4.4489e-07
DNB | B | HGS->MGS | 26 | 4 | 4.4361e-07
DNB | B | HGS->MGS | 26 | 5 | 4.4309e-07
DNB | B | HGS->MGS | 26 | 6 | 4.4025e-07
DNB | B | HGS->MGS | 26 | 7 | 4.3662e-07
DNB | B | HGS->MGS | 26 | 8 | 4.411e-07
DNB | B | HGS->MGS | 26 | 9 | 4.3872e-07
DNB | B | HGS->MGS | 26 | 10 | 4.4257e-07
DNB | B | HGS->MGS | 26 | 11 | 4.3739e-07
DNB | B | HGS->MGS | 26 | 12 | 4.3933e-07
DNB | B | HGS->MGS | 26 | 13 | 4.4236e-07
DNB | B | HGS->MGS | 26 | 14 | 4.4046e-07
DNB | B | HGS->MGS | 26 | 15 | 4.3645e-07
DNB | B | HGS->MGS | 26 | 16 | 4.4113e-07
DNB | B | HGS->MGS | 26 | 1 | 7.089e-07
DNB | B | HGS->MGS | 26 | 2 | 7.0465e-07
DNB | B | HGS->MGS | 26 | 3 | 7.0322e-07
DNB | B | HGS->MGS | 26 | 4 | 7.0669e-07
DNB | B | HGS->MGS | 26 | 5 | 7.0448e-07
DNB | B | HGS->MGS | 26 | 6 | 7.0303e-07
DNB | B | HGS->MGS | 26 | 7 | 7.0228e-07
DNB | B | HGS->MGS | 26 | 8 | 6.9849e-07
DNB | B | HGS->MGS | 26 | 9 | 6.9615e-07
DNB | B | HGS->MGS | 26 | 10 | 6.9856e-07
DNB | B | HGS->MGS | 26 | 11 | 6.9986e-07
DNB | B | HGS->MGS | 26 | 12 | 6.9939e-07
DNB | B | HGS->MGS | 26 | 13 | 7.0044e-07
DNB | B | HGS->MGS | 26 | 14 | 6.953e-07
DNB | B | HGS->MGS | 26 | 15 | 6.9696e-07
DNB | B | HGS->MGS | 26 | 16 | 7.0106e-07
DNB | B | HGS->MGS | 32 | 1 | 1.0551e-06
DNB | B | HGS->MGS | 32 | 2 | 1.0695e-06
DNB | B | HGS->MGS | 32 | 3 | 1.0674e-06
DNB | B | HGS->MGS | 32 | 4 | 1.0641e-06
DNB | B | HGS->MGS | 32 | 5 | 1.0621e-06
DNB | B | HGS->MGS | 32 | 6 | 1.0587e-06
DNB | B | HGS->MGS | 32 | 7 | 1.059e-06
DNB | B | HGS->MGS | 32 | 8 | 1.0497e-06
DNB | B | HGS->MGS | 32 | 9 | 1.0539e-06
DNB | B | HGS->MGS | 32 | 10 | 1.053e-06
DNB | B | HGS->MGS | 32 | 11 | 1.054e-06
DNB | B | HGS->MGS | 32 | 12 | 1.0558e-06
DNB | B | HGS->MGS | 32 | 13 | 1.0575e-06
DNB | B | HGS->MGS | 32 | 14 | 1.059e-06
DNB | B | HGS->MGS | 32 | 15 | 1.0571e-06
DNB | B | HGS->MGS | 32 | 16 | 1.0565e-06
DNB | B | HGS->MGS | 32 | 1 | 1.6657e-06
DNB | B | HGS->MGS | 32 | 2 | 1.6099e-06
DNB | B | HGS->MGS | 32 | 3 | 1.5371e-06
DNB | B | HGS->MGS | 32 | 4 | 1.5939e-06
DNB | B | HGS->MGS | 32 | 5 | 1.5895e-06
DNB | B | HGS->MGS | 32 | 6 | 1.5888e-06
DNB | B | HGS->MGS | 32 | 7 | 1.5769e-06
DNB | B | HGS->MGS | 32 | 8 | 1.5726e-06
DNB | B | HGS->MGS | 32 | 9 | 1.5571e-06
DNB | B | HGS->MGS | 32 | 10 | 1.551e-06
DNB | B | HGS->MGS | 32 | 11 | 1.5412e-06
DNB | B | HGS->MGS | 32 | 12 | 1.551e-06
DNB | B | HGS->MGS | 32 | 13 | 1.5515e-06
DNB | B | HGS->MGS | 32 | 14 | 1.5652e-06
DNB | B | HGS->MGS | 32 | 15 | 1.566e-06
DNB | B | HGS->MGS | 32 | 16 | 1.5791e-06
DNB | B | HGS->MGS | 34 | 1 | 2.4093e-06
DNB | B | HGS->MGS | 34 | 2 | 2.4096e-06
DNB | B | HGS->MGS | 34 | 3 | 2.7281e-06
DNB | B | HGS->MGS | 34 | 4 | 2.4795e-06
DNB | B | HGS->MGS | 34 | 5 | 2.4514e-06
DNB | B | HGS->MGS | 34 | 6 | 2.4619e-06
DNB | B | HGS->MGS | 34 | 7 | 2.4651e-06
DNB | B | HGS->MGS | 34 | 8 | 2.4907e-06
DNB | B | HGS->MGS | 34 | 9 | 2.5363e-06
DNB | B | HGS->MGS | 34 | 10 | 2.5179e-06
DNB | B | HGS->MGS | 34 | 11 | 2.5097e-06
DNB | B | HGS->MGS | 34 | 12 | 2.5622e-06
DNB | B | HGS->MGS | 34 | 13 | 2.3538e-06
DNB | B | HGS->MGS | 34 | 14 | 2.5353e-06
DNB | B | HGS->MGS | 34 | 15 | 2.6454e-06
DNB | B | HGS->MGS | 34 | 16 | 2.3622e-06
DNB | B | HGS->MGS | 34 | 1 | 9.526e-08
DNB | B | HGS->MGS | 34 | 2 | 1.0866e-07
DNB | B | HGS->MGS | 34 | 3 | 1.0149e-07
DNB | B | HGS->MGS | 34 | 4 | 9.5312e-08
DNB | B | HGS->MGS | 34 | 5 | 9.9526e-08
DNB | B | HGS->MGS | 34 | 6 | 1.0387e-07
DNB | B | HGS->MGS | 34 | 7 | 9.8643e-08
DNB | B | HGS->MGS | 34 | 8 | 9.9705e-08
DNB | B | HGS->MGS | 34 | 9 | 1.061e-07
DNB | B | HGS->MGS | 34 | 10 | 9.9855e-08
DNB | B | HGS->MGS | 34 | 11 | 1.0179e-07
DNB | B | HGS->MGS | 34 | 12 | 9.7964e-08
DNB | B | HGS->MGS | 34 | 13 | 9.4037e-08
DNB | B | HGS->MGS | 34 | 14 | 1.0052e-07
DNB | B | HGS->MGS | 34 | 15 | 9.9922e-08
DNB | B | HGS->MGS | 34 | 16 | 9.459e-08

## TEB RVS
Band | HAM | Band Median |  |  |  | Band Average
|  | a0 | a1 | a2 | Uncertainty | a0 | a1 | a2 | Uncertainty
I4 | A | 1.0061 | -2.6749e-06 | -1.5766e-06 | 0.052455 | 1.0063 | -1.7511e-05 | -1.4425e-06 | 0.053508
I4 | B | 1.005 | 4.608e-05 | -2.1845e-06 | 0.055739 | 1.0055 | 2.5397e-05 | -1.934e-06 | 0.05808
I5 | A | 1.042 | 0.00062218 | -2.2051e-05 | 0.079962 | 1.0426 | 0.00060981 | -2.1883e-05 | 0.085399
I5 | B | 1.046 | 0.00067612 | -2.3922e-05 | 0.084147 | 1.046 | 0.00066414 | -2.3743e-05 | 0.085505
M12 | A | 1.0057 | 1.6582e-05 | -1.9437e-06 | 0.052429 | 1.0058 | 2.2924e-05 | -1.9883e-06 | 0.051263
M12 | B | 1.0057 | 1.9966e-05 | -2.0066e-06 | 0.051833 | 1.006 | 2.2019e-05 | -2.0334e-06 | 0.049414
M13 | A | 1.0058 | 1.0604e-05 | -1.7763e-06 | 0.060572 | 1.006 | 3.4044e-07 | -1.6596e-06 | 0.059135
M13 | B | 1.0059 | 3.1996e-05 | -2.1254e-06 | 0.060843 | 1.0059 | 1.056e-05 | -1.7984e-06 | 0.058583
M14 | A | 1.0958 | 0.0012637 | -4.7313e-05 | 0.093693 | 1.0956 | 0.0012552 | -4.7256e-05 | 0.092116
M14 | B | 1.0947 | 0.0012681 | -4.723e-05 | 0.093692 | 1.0948 | 0.0012687 | -4.7259e-05 | 0.094158
M15 | A | 1.0575 | 0.00093777 | -3.1598e-05 | 0.069706 | 1.0573 | 0.00097008 | -3.1929e-05 | 0.071884
M15 | B | 1.0613 | 0.0009932 | -3.362e-05 | 0.072314 | 1.061 | 0.0010254 | -3.389e-05 | 0.074612
M16A | A | 1.028 | 0.00048509 | -1.5729e-05 | 0.057336 | 1.0283 | 0.00047248 | -1.5672e-05 | 0.059928
M16A | B | 1.0313 | 0.00054753 | -1.7732e-05 | 0.064438 | 1.0315 | 0.00052708 | -1.7445e-05 | 0.063948
M16B | A | 1.0275 | 0.00045643 | -1.5158e-05 | 0.061831 | 1.0276 | 0.00045377 | -1.515e-05 | 0.062489
M16B | B | 1.0302 | 0.00055 | -1.7298e-05 | 0.065576 | 1.0302 | 0.00052911 | -1.712e-05 | 0.064718
Band | HAM | Detector | Detector Dependent |  |  |  |  | Notes:
|  |  | a0 | a1 | a2 | Uncertainty |  | 1. This table provides VIIRS F2 TEB RVS.
2. The results were derived from FP-10 using the LABB source.
I4 | A | 1 | 1.0051 | 6.551e-05 | -2.4943e-06 | 0.062047 |  | 3. Quadratic fitting was performed on RVS vs HAM AOI.
I4 | A | 2 | 1.0039 | 0.0001366 | -3.3426e-06 | 0.050563 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I4 | A | 3 | 1.0055 | 4.0161e-05 | -2.1926e-06 | 0.040535 |  | 5. The test was set to diagnostic mode.
I4 | A | 4 | 1.0031 | 0.00014781 | -3.3259e-06 | 0.055385
I4 | A | 5 | 1.0031 | 0.00012548 | -2.93e-06 | 0.049877
I4 | A | 6 | 1.0059 | -1.3719e-06 | -1.6079e-06 | 0.0519
I4 | A | 7 | 1.0076 | -6.5217e-05 | -1.0193e-06 | 0.052326
I4 | A | 8 | 1.0056 | 2.3021e-05 | -1.921e-06 | 0.047623
I4 | A | 9 | 1.0063 | -3.978e-06 | -1.669e-06 | 0.045985
I4 | A | 10 | 1.0034 | 0.00015224 | -3.459e-06 | 0.053299
I4 | A | 11 | 1.0042 | 7.7762e-05 | -2.4498e-06 | 0.050634
I4 | A | 12 | 1.0046 | 5.1204e-05 | -2.1109e-06 | 0.044918
I4 | A | 13 | 1.0056 | 8.5977e-06 | -1.6941e-06 | 0.046104
I4 | A | 14 | 1.0056 | -1.0729e-07 | -1.5432e-06 | 0.047361
I4 | A | 15 | 1.0048 | 4.5983e-05 | -2.0806e-06 | 0.047825
I4 | A | 16 | 1.0057 | 1.1867e-05 | -1.7828e-06 | 0.045274
I4 | A | 17 | 1.006 | 4.4868e-06 | -1.727e-06 | 0.049688
I4 | A | 18 | 1.0058 | 1.1476e-05 | -1.7803e-06 | 0.055216
I4 | A | 19 | 1.0078 | -7.8807e-05 | -8.4727e-07 | 0.041501
I4 | A | 20 | 1.0081 | -0.00010221 | -5.3838e-07 | 0.059823
I4 | A | 21 | 1.0065 | -1.5632e-05 | -1.5453e-06 | 0.061492
I4 | A | 22 | 1.0076 | -9.5805e-05 | -5.1953e-07 | 0.053162
I4 | A | 23 | 1.0063 | -3.3713e-05 | -1.1707e-06 | 0.052387
I4 | A | 24 | 1.0083 | -0.00011906 | -3.1271e-07 | 0.059016
I4 | A | 25 | 1.0066 | -3.4126e-05 | -1.2561e-06 | 0.053736
I4 | A | 26 | 1.0073 | -7.8935e-05 | -7.1456e-07 | 0.053814
I4 | A | 27 | 1.0068 | -7.1654e-05 | -6.9511e-07 | 0.052523
I4 | A | 28 | 1.0093 | -0.00018036 | 4.2559e-07 | 0.057929
I4 | A | 29 | 1.0084 | -0.00014435 | 7.8074e-08 | 0.069472
I4 | A | 30 | 1.0072 | -8.657e-05 | -5.4337e-07 | 0.073651
I4 | A | 31 | 1.0103 | -0.00021701 | 7.7126e-07 | 0.065586
I4 | A | 32 | 1.0086 | -0.00013363 | -1.615e-07 | 0.061592
I4 | B | 1 | 1.0057 | 4.4851e-05 | -2.3194e-06 | 0.056675
I4 | B | 2 | 1.0096 | -0.00020722 | 7.9197e-07 | 0.099134
I4 | B | 3 | 1.005 | 7.1129e-05 | -2.5556e-06 | 0.064045
I4 | B | 4 | 1.0037 | 0.00011594 | -2.9594e-06 | 0.052546
I4 | B | 5 | 1.0019 | 0.00022132 | -4.1999e-06 | 0.067194
I4 | B | 6 | 1.0041 | 0.00010713 | -2.9188e-06 | 0.047382
I4 | B | 7 | 1.0057 | 3.9099e-05 | -2.2137e-06 | 0.04917
I4 | B | 8 | 1.0043 | 9.2196e-05 | -2.7218e-06 | 0.059852
I4 | B | 9 | 1.0032 | 0.00016469 | -3.6066e-06 | 0.048499
I4 | B | 10 | 1.0117 | -0.00035129 | 2.5953e-06 | 0.093374
I4 | B | 11 | 1.0049 | 6.8414e-05 | -2.4973e-06 | 0.052652
I4 | B | 12 | 1.0025 | 0.0001803 | -3.6832e-06 | 0.046763
I4 | B | 13 | 1.0043 | 8.2118e-05 | -2.5547e-06 | 0.038105
I4 | B | 14 | 1.0052 | 5.9254e-05 | -2.4077e-06 | 0.060035
I4 | B | 15 | 1.0048 | 4.5727e-05 | -2.0767e-06 | 0.046466
I4 | B | 16 | 1.0045 | 8.9246e-05 | -2.7326e-06 | 0.050422
I4 | B | 17 | 1.0034 | 0.00014117 | -3.29e-06 | 0.050254
I4 | B | 18 | 1.0048 | 6.0142e-05 | -2.3139e-06 | 0.057074
I4 | B | 19 | 1.0052 | 3.8603e-05 | -2.0825e-06 | 0.053895
I4 | B | 20 | 1.005 | 4.3705e-05 | -2.1091e-06 | 0.054982
I4 | B | 21 | 1.0041 | 0.00010364 | -2.8456e-06 | 0.060478
I4 | B | 22 | 1.0066 | -2.1248e-05 | -1.4806e-06 | 0.064443
I4 | B | 23 | 1.0049 | 3.7191e-05 | -1.9837e-06 | 0.06031
I4 | B | 24 | 1.005 | 4.6434e-05 | -2.1552e-06 | 0.055039
I4 | B | 25 | 1.0058 | -6.0788e-06 | -1.5e-06 | 0.048912
I4 | B | 26 | 1.0056 | 3.3625e-06 | -1.6061e-06 | 0.053416
I4 | B | 27 | 1.0097 | -0.00018322 | 3.6988e-07 | 0.065818
I4 | B | 28 | 1.0056 | 2.2187e-05 | -1.9011e-06 | 0.055579
I4 | B | 29 | 1.0081 | -0.00011534 | -3.2887e-07 | 0.058846
I4 | B | 30 | 1.0044 | 5.0061e-05 | -2.0494e-06 | 0.071019
I4 | B | 31 | 1.0072 | -7.7331e-05 | -6.9642e-07 | 0.0559
I4 | B | 32 | 1.0087 | -0.00015345 | 1.4539e-07 | 0.060289
I5 | A | 1 | 1.0411 | 0.00071154 | -2.3176e-05 | 0.091223
I5 | A | 2 | 1.0421 | 0.00058781 | -2.1384e-05 | 0.07898
I5 | A | 3 | 1.0388 | 0.00083963 | -2.4676e-05 | 0.10767
I5 | A | 4 | 1.0446 | 0.00047497 | -2.0196e-05 | 0.11745
I5 | A | 5 | 1.0445 | 0.00054868 | -2.14e-05 | 0.095043
I5 | A | 6 | 1.0416 | 0.00062518 | -2.1862e-05 | 0.097793
I5 | A | 7 | 1.0427 | 0.00061918 | -2.2069e-05 | 0.101
I5 | A | 8 | 1.0384 | 0.00076237 | -2.3259e-05 | 0.074164
I5 | A | 9 | 1.0395 | 0.0007711 | -2.3709e-05 | 0.0873
I5 | A | 10 | 1.0393 | 0.00076458 | -2.3558e-05 | 0.092621
I5 | A | 11 | 1.0425 | 0.00064064 | -2.2371e-05 | 0.076437
I5 | A | 12 | 1.0412 | 0.00063538 | -2.1926e-05 | 0.080489
I5 | A | 13 | 1.0404 | 0.00075577 | -2.3724e-05 | 0.065262
I5 | A | 14 | 1.0401 | 0.00070796 | -2.2828e-05 | 0.073469
I5 | A | 15 | 1.0448 | 0.00056281 | -2.1732e-05 | 0.066093
I5 | A | 16 | 1.0514 | 0.00021411 | -1.7756e-05 | 0.11532
I5 | A | 17 | 1.0403 | 0.00074026 | -2.3416e-05 | 0.067944
I5 | A | 18 | 1.0481 | 0.00025152 | -1.7471e-05 | 0.11494
I5 | A | 19 | 1.0412 | 0.00074113 | -2.3679e-05 | 0.059817
I5 | A | 20 | 1.0419 | 0.00064574 | -2.2313e-05 | 0.088507
I5 | A | 21 | 1.0425 | 0.00061891 | -2.2033e-05 | 0.067367
I5 | A | 22 | 1.042 | 0.00057443 | -2.1146e-05 | 0.076182
I5 | A | 23 | 1.0407 | 0.00070217 | -2.2919e-05 | 0.075419
I5 | A | 24 | 1.0419 | 0.00056321 | -2.0933e-05 | 0.070526
I5 | A | 25 | 1.0442 | 0.00060695 | -2.2301e-05 | 0.080888
I5 | A | 26 | 1.0437 | 0.00052705 | -2.0833e-05 | 0.094727
I5 | A | 27 | 1.0446 | 0.0005473 | -2.1422e-05 | 0.071934
I5 | A | 28 | 1.0483 | 0.00031214 | -1.8517e-05 | 0.11245
I5 | A | 29 | 1.0387 | 0.00082519 | -2.4408e-05 | 0.079434
I5 | A | 30 | 1.0391 | 0.00074412 | -2.3149e-05 | 0.072289
I5 | A | 31 | 1.0473 | 0.00040441 | -1.9782e-05 | 0.072136
I5 | A | 32 | 1.0442 | 0.00048777 | -2.0303e-05 | 0.10789
I5 | B | 1 | 1.0462 | 0.00072276 | -2.4757e-05 | 0.091937
I5 | B | 2 | 1.0442 | 0.00069431 | -2.375e-05 | 0.093128
I5 | B | 3 | 1.0491 | 0.00045795 | -2.1175e-05 | 0.10489
I5 | B | 4 | 1.0435 | 0.00071213 | -2.3858e-05 | 0.090934
I5 | B | 5 | 1.0482 | 0.00058346 | -2.2995e-05 | 0.077687
I5 | B | 6 | 1.0435 | 0.00073221 | -2.417e-05 | 0.088138
I5 | B | 7 | 1.0466 | 0.00067606 | -2.4088e-05 | 0.079905
I5 | B | 8 | 1.049 | 0.00046131 | -2.1196e-05 | 0.086321
I5 | B | 9 | 1.0494 | 0.00057775 | -2.3235e-05 | 0.099778
I5 | B | 10 | 1.045 | 0.0006533 | -2.327e-05 | 0.083331
I5 | B | 11 | 1.0431 | 0.00084437 | -2.5942e-05 | 0.073648
I5 | B | 12 | 1.047 | 0.0005554 | -2.2216e-05 | 0.068286
I5 | B | 13 | 1.0471 | 0.0006262 | -2.341e-05 | 0.071175
I5 | B | 14 | 1.0444 | 0.00070493 | -2.3986e-05 | 0.079051
I5 | B | 15 | 1.0463 | 0.00071217 | -2.4611e-05 | 0.065347
I5 | B | 16 | 1.0417 | 0.0009445 | -2.7196e-05 | 0.11815
I5 | B | 17 | 1.054 | 0.00035326 | -2.0792e-05 | 0.086696
I5 | B | 18 | 1.0419 | 0.0008217 | -2.5211e-05 | 0.14573
I5 | B | 19 | 1.0438 | 0.00082735 | -2.5855e-05 | 0.074061
I5 | B | 20 | 1.0412 | 0.00086788 | -2.579e-05 | 0.076851
I5 | B | 21 | 1.043 | 0.0008379 | -2.5784e-05 | 0.066101
I5 | B | 22 | 1.0467 | 0.00060894 | -2.3016e-05 | 0.083357
I5 | B | 23 | 1.0443 | 0.00085686 | -2.6468e-05 | 0.084936
I5 | B | 24 | 1.0459 | 0.0006551 | -2.3566e-05 | 0.069966
I5 | B | 25 | 1.0474 | 0.00066129 | -2.4088e-05 | 0.067445
I5 | B | 26 | 1.0509 | 0.00039307 | -2.0575e-05 | 0.078661
I5 | B | 27 | 1.044 | 0.00073396 | -2.434e-05 | 0.085365
I5 | B | 28 | 1.0519 | 0.00033804 | -1.9957e-05 | 0.095674
I5 | B | 29 | 1.0476 | 0.00061082 | -2.3289e-05 | 0.086909
I5 | B | 30 | 1.0441 | 0.00072625 | -2.4241e-05 | 0.076948
I5 | B | 31 | 1.0463 | 0.00067618 | -2.4022e-05 | 0.093718
I5 | B | 32 | 1.0455 | 0.00062518 | -2.294e-05 | 0.09204
M12 | A | 1 | 1.0054 | 6.2344e-05 | -2.519e-06 | 0.052476
M12 | A | 2 | 1.0044 | 0.00010275 | -2.9178e-06 | 0.05512
M12 | A | 3 | 1.005 | 8.9145e-05 | -2.8582e-06 | 0.052564
M12 | A | 4 | 1.0041 | 0.00011997 | -3.1161e-06 | 0.044507
M12 | A | 5 | 1.0053 | 7.2384e-05 | -2.6798e-06 | 0.032714
M12 | A | 6 | 1.0053 | 3.5437e-05 | -2.0637e-06 | 0.053675
M12 | A | 7 | 1.0049 | 9.6137e-05 | -2.9495e-06 | 0.037132
M12 | A | 8 | 1.006 | 1.0298e-05 | -1.8237e-06 | 0.052382
M12 | A | 9 | 1.0063 | 2.2866e-05 | -2.1196e-06 | 0.04105
M12 | A | 10 | 1.0058 | 2.822e-06 | -1.6453e-06 | 0.052153
M12 | A | 11 | 1.0072 | -3.9983e-05 | -1.3198e-06 | 0.051709
M12 | A | 12 | 1.0072 | -6.9797e-05 | -8.1733e-07 | 0.055232
M12 | A | 13 | 1.007 | -4.7804e-05 | -1.149e-06 | 0.047746
M12 | A | 14 | 1.0065 | -1.415e-05 | -1.5715e-06 | 0.056634
M12 | A | 15 | 1.0072 | -6.503e-05 | -8.9983e-07 | 0.055189
M12 | A | 16 | 1.0056 | -1.0599e-05 | -1.3635e-06 | 0.079919
M12 | B | 1 | 1.0049 | 8.4213e-05 | -2.7561e-06 | 0.050874
M12 | B | 2 | 1.0048 | 7.8007e-05 | -2.6353e-06 | 0.053995
M12 | B | 3 | 1.005 | 8.7418e-05 | -2.8327e-06 | 0.036439
M12 | B | 4 | 1.0047 | 0.00011242 | -3.1589e-06 | 0.045209
M12 | B | 5 | 1.0048 | 0.00010238 | -3.0331e-06 | 0.033177
M12 | B | 6 | 1.0047 | 8.8689e-05 | -2.7733e-06 | 0.049479
M12 | B | 7 | 1.0051 | 8.932e-05 | -2.8801e-06 | 0.032329
M12 | B | 8 | 1.0067 | -1.2777e-05 | -1.6448e-06 | 0.049731
M12 | B | 9 | 1.0067 | 1.2969e-05 | -2.0554e-06 | 0.040006
M12 | B | 10 | 1.0055 | 2.6964e-05 | -1.9577e-06 | 0.054044
M12 | B | 11 | 1.0074 | -3.8114e-05 | -1.4178e-06 | 0.054772
M12 | B | 12 | 1.0072 | -6.9838e-05 | -8.3687e-07 | 0.054552
M12 | B | 13 | 1.0085 | -0.00010451 | -6.1744e-07 | 0.052791
M12 | B | 14 | 1.007 | -3.1146e-05 | -1.4022e-06 | 0.055039
M12 | B | 15 | 1.0077 | -7.3519e-05 | -9.0465e-07 | 0.057189
M12 | B | 16 | 1.0059 | -1.8196e-07 | -1.6286e-06 | 0.070992
M13 | A | 1 | 1.0047 | 6.1273e-05 | -2.3265e-06 | 0.060648
M13 | A | 2 | 1.0035 | 0.00013155 | -3.1612e-06 | 0.077079
M13 | A | 3 | 1.0035 | 0.00014388 | -3.3635e-06 | 0.060804
M13 | A | 4 | 1.0057 | 1.3422e-05 | -1.8085e-06 | 0.060496
M13 | A | 5 | 1.0043 | 6.577e-05 | -2.2756e-06 | 0.046578
M13 | A | 6 | 1.0053 | 3.4862e-05 | -2.052e-06 | 0.059965
M13 | A | 7 | 1.0054 | 1.531e-05 | -1.7442e-06 | 0.059582
M13 | A | 8 | 1.0064 | -1.2887e-05 | -1.5538e-06 | 0.057338
M13 | A | 9 | 1.0061 | 7.7867e-06 | -1.8199e-06 | 0.064272
M13 | A | 10 | 1.0067 | -3.2025e-05 | -1.3312e-06 | 0.068581
M13 | A | 11 | 1.0082 | -0.0001091 | -4.452e-07 | 0.03683
M13 | A | 12 | 1.0084 | -0.0001275 | -1.8725e-07 | 0.03333
M13 | A | 13 | 1.0057 | 4.078e-05 | -2.252e-06 | 0.067524
M13 | A | 14 | 1.0082 | -0.00011659 | -3.2505e-07 | 0.051217
M13 | A | 15 | 1.0059 | -7.4608e-06 | -1.5015e-06 | 0.073933
M13 | A | 16 | 1.0077 | -0.00010361 | -4.0646e-07 | 0.067987
M13 | B | 1 | 1.004 | 0.00010363 | -2.8297e-06 | 0.058735
M13 | B | 2 | 1.0051 | 6.1322e-05 | -2.419e-06 | 0.067559
M13 | B | 3 | 1.0034 | 0.00015332 | -3.4952e-06 | 0.059341
M13 | B | 4 | 1.0054 | 3.9471e-05 | -2.1447e-06 | 0.062304
M13 | B | 5 | 1.0043 | 7.1501e-05 | -2.3827e-06 | 0.047963
M13 | B | 6 | 1.0062 | -5.9353e-06 | -1.6216e-06 | 0.061842
M13 | B | 7 | 1.0044 | 6.631e-05 | -2.3077e-06 | 0.057668
M13 | B | 8 | 1.0063 | -4.1123e-06 | -1.6595e-06 | 0.059844
M13 | B | 9 | 1.0056 | 3.9314e-05 | -2.2009e-06 | 0.065799
M13 | B | 10 | 1.0053 | 4.9419e-05 | -2.2737e-06 | 0.064889
M13 | B | 11 | 1.0065 | -2.5028e-05 | -1.3895e-06 | 0.037657
M13 | B | 12 | 1.0082 | -0.00011197 | -4.0307e-07 | 0.036854
M13 | B | 13 | 1.0061 | 2.4679e-05 | -2.1061e-06 | 0.064667
M13 | B | 14 | 1.0075 | -8.6442e-05 | -6.4629e-07 | 0.04956
M13 | B | 15 | 1.0076 | -8.4329e-05 | -6.8457e-07 | 0.074521
M13 | B | 16 | 1.0081 | -0.0001222 | -2.0958e-07 | 0.068125
M14 | A | 1 | 1.0928 | 0.0014025 | -4.8919e-05 | 0.10113
M14 | A | 2 | 1.0933 | 0.0013042 | -4.7442e-05 | 0.09435
M14 | A | 3 | 1.0963 | 0.0012466 | -4.7297e-05 | 0.10129
M14 | A | 4 | 1.0945 | 0.0012774 | -4.7329e-05 | 0.10028
M14 | A | 5 | 1.0945 | 0.0013404 | -4.8357e-05 | 0.096623
M14 | A | 6 | 1.0966 | 0.0011867 | -4.6406e-05 | 0.10324
M14 | A | 7 | 1.0979 | 0.0011616 | -4.6338e-05 | 0.093099
M14 | A | 8 | 1.0936 | 0.0013212 | -4.7802e-05 | 0.079737
M14 | A | 9 | 1.0957 | 0.0012871 | -4.7799e-05 | 0.085614
M14 | A | 10 | 1.096 | 0.0012017 | -4.6476e-05 | 0.089973
M14 | A | 11 | 1.095 | 0.0013195 | -4.8169e-05 | 0.094288
M14 | A | 12 | 1.0964 | 0.00117 | -4.6051e-05 | 0.087193
M14 | A | 13 | 1.0961 | 0.0012971 | -4.8085e-05 | 0.077481
M14 | A | 14 | 1.0951 | 0.00125 | -4.7031e-05 | 0.079252
M14 | A | 15 | 1.099 | 0.0011527 | -4.6484e-05 | 0.089857
M14 | A | 16 | 1.0969 | 0.001165 | -4.6111e-05 | 0.10044
M14 | B | 1 | 1.0923 | 0.0013988 | -4.8739e-05 | 0.10202
M14 | B | 2 | 1.0931 | 0.0012854 | -4.708e-05 | 0.092332
M14 | B | 3 | 1.0968 | 0.0011957 | -4.6585e-05 | 0.10377
M14 | B | 4 | 1.0917 | 0.0013902 | -4.8416e-05 | 0.10138
M14 | B | 5 | 1.0938 | 0.0013521 | -4.8355e-05 | 0.094302
M14 | B | 6 | 1.0908 | 0.0014993 | -4.9972e-05 | 0.10859
M14 | B | 7 | 1.0968 | 0.0011878 | -4.6458e-05 | 0.09263
M14 | B | 8 | 1.0944 | 0.0012563 | -4.6945e-05 | 0.081118
M14 | B | 9 | 1.0941 | 0.0013292 | -4.806e-05 | 0.093082
M14 | B | 10 | 1.095 | 0.0012261 | -4.6617e-05 | 0.097592
M14 | B | 11 | 1.0953 | 0.00128 | -4.7582e-05 | 0.094772
M14 | B | 12 | 1.0953 | 0.001186 | -4.6029e-05 | 0.089062
M14 | B | 13 | 1.0964 | 0.0012501 | -4.7381e-05 | 0.078017
M14 | B | 14 | 1.0923 | 0.0013587 | -4.8071e-05 | 0.085276
M14 | B | 15 | 1.0987 | 0.001139 | -4.6175e-05 | 0.091635
M14 | B | 16 | 1.1002 | 0.00096406 | -4.368e-05 | 0.10093
M15 | A | 1 | 1.0547 | 0.0011496 | -3.4194e-05 | 0.083573
M15 | A | 2 | 1.0553 | 0.0010332 | -3.2443e-05 | 0.061822
M15 | A | 3 | 1.0561 | 0.0010558 | -3.3039e-05 | 0.087056
M15 | A | 4 | 1.0563 | 0.00098179 | -3.1851e-05 | 0.066513
M15 | A | 5 | 1.0587 | 0.00091849 | -3.1467e-05 | 0.084828
M15 | A | 6 | 1.058 | 0.00090887 | -3.1129e-05 | 0.073276
M15 | A | 7 | 1.0564 | 0.0010339 | -3.2744e-05 | 0.07658
M15 | A | 8 | 1.0574 | 0.00090659 | -3.0913e-05 | 0.060106
M15 | A | 9 | 1.0581 | 0.00094204 | -3.1704e-05 | 0.068248
M15 | A | 10 | 1.0576 | 0.0009273 | -3.1307e-05 | 0.061909
M15 | A | 11 | 1.0578 | 0.00097238 | -3.2128e-05 | 0.075223
M15 | A | 12 | 1.0576 | 0.0009315 | -3.1388e-05 | 0.067894
M15 | A | 13 | 1.0594 | 0.00090815 | -3.1493e-05 | 0.071164
M15 | A | 14 | 1.0592 | 0.00086058 | -3.0657e-05 | 0.066187
M15 | A | 15 | 1.0561 | 0.0010577 | -3.3065e-05 | 0.078264
M15 | A | 16 | 1.0573 | 0.00093351 | -3.1342e-05 | 0.067507
M15 | B | 1 | 1.0595 | 0.0011484 | -3.552e-05 | 0.080904
M15 | B | 2 | 1.059 | 0.0010995 | -3.4566e-05 | 0.067795
M15 | B | 3 | 1.0601 | 0.0010965 | -3.4821e-05 | 0.087004
M15 | B | 4 | 1.06 | 0.0010341 | -3.3749e-05 | 0.069563
M15 | B | 5 | 1.0623 | 0.00098163 | -3.35e-05 | 0.082204
M15 | B | 6 | 1.0626 | 0.00092081 | -3.2592e-05 | 0.070974
M15 | B | 7 | 1.0586 | 0.0011807 | -3.5788e-05 | 0.086706
M15 | B | 8 | 1.0604 | 0.00099033 | -3.313e-05 | 0.059283
M15 | B | 9 | 1.0611 | 0.0010666 | -3.4591e-05 | 0.070597
M15 | B | 10 | 1.0616 | 0.00096309 | -3.301e-05 | 0.066795
M15 | B | 11 | 1.0622 | 0.00099607 | -3.374e-05 | 0.078044
M15 | B | 12 | 1.0619 | 0.00095664 | -3.2981e-05 | 0.073634
M15 | B | 13 | 1.0637 | 0.00093793 | -3.3183e-05 | 0.070994
M15 | B | 14 | 1.0621 | 0.00095168 | -3.2974e-05 | 0.06892
M15 | B | 15 | 1.0599 | 0.0011097 | -3.4967e-05 | 0.083669
M15 | B | 16 | 1.0615 | 0.0009722 | -3.3124e-05 | 0.076703
M16A | A | 1 | 1.0299 | 0.00043284 | -1.5435e-05 | 0.077953
M16A | A | 2 | 1.0275 | 0.00048952 | -1.5715e-05 | 0.055162
M16A | A | 3 | 1.0272 | 0.00053581 | -1.6401e-05 | 0.05602
M16A | A | 4 | 1.0278 | 0.00047289 | -1.5524e-05 | 0.05682
M16A | A | 5 | 1.028 | 0.00052232 | -1.6422e-05 | 0.05895
M16A | A | 6 | 1.0274 | 0.00049238 | -1.5743e-05 | 0.05763
M16A | A | 7 | 1.0295 | 0.00043255 | -1.5336e-05 | 0.055367
M16A | A | 8 | 1.0279 | 0.00049262 | -1.5888e-05 | 0.049488
M16A | A | 9 | 1.0284 | 0.00048065 | -1.5836e-05 | 0.063143
M16A | A | 10 | 1.0296 | 0.00041214 | -1.5013e-05 | 0.057043
M16A | A | 11 | 1.0268 | 0.00055532 | -1.6636e-05 | 0.065543
M16A | A | 12 | 1.0307 | 0.00034313 | -1.4176e-05 | 0.050269
M16A | A | 13 | 1.031 | 0.00036341 | -1.4609e-05 | 0.06535
M16A | A | 14 | 1.0292 | 0.00039963 | -1.4708e-05 | 0.070944
M16A | A | 15 | 1.0274 | 0.00051648 | -1.6156e-05 | 0.055568
M16A | A | 16 | 1.025 | 0.00061793 | -1.7158e-05 | 0.063595
M16A | B | 1 | 1.0309 | 0.00058948 | -1.8335e-05 | 0.072692
M16A | B | 2 | 1.0303 | 0.00058616 | -1.8105e-05 | 0.057236
M16A | B | 3 | 1.0308 | 0.0005704 | -1.7993e-05 | 0.055208
M16A | B | 4 | 1.0303 | 0.00056071 | -1.7694e-05 | 0.068438
M16A | B | 5 | 1.0324 | 0.00053434 | -1.7816e-05 | 0.06799
M16A | B | 6 | 1.0306 | 0.0005616 | -1.7771e-05 | 0.068066
M16A | B | 7 | 1.0315 | 0.00053237 | -1.7533e-05 | 0.058499
M16A | B | 8 | 1.0335 | 0.00041952 | -1.6216e-05 | 0.054706
M16A | B | 9 | 1.0311 | 0.00057144 | -1.808e-05 | 0.071925
M16A | B | 10 | 1.0334 | 0.00041717 | -1.6141e-05 | 0.062413
M16A | B | 11 | 1.0283 | 0.00068397 | -1.9178e-05 | 0.066463
M16A | B | 12 | 1.0329 | 0.00043921 | -1.6385e-05 | 0.059057
M16A | B | 13 | 1.0318 | 0.00051777 | -1.7375e-05 | 0.072342
M16A | B | 14 | 1.0338 | 0.00036971 | -1.5485e-05 | 0.075555
M16A | B | 15 | 1.0298 | 0.00060484 | -1.8274e-05 | 0.050833
M16A | B | 16 | 1.0321 | 0.00047454 | -1.6739e-05 | 0.061741
M16B | A | 1 | 1.0285 | 0.00042319 | -1.4909e-05 | 0.061847
M16B | A | 2 | 1.0284 | 0.00038791 | -1.4298e-05 | 0.075443
M16B | A | 3 | 1.0263 | 0.00052889 | -1.6058e-05 | 0.075582
M16B | A | 4 | 1.0261 | 0.00048883 | -1.5341e-05 | 0.070876
M16B | A | 5 | 1.0281 | 0.00045639 | -1.535e-05 | 0.052859
M16B | A | 6 | 1.0261 | 0.00049185 | -1.5368e-05 | 0.071199
M16B | A | 7 | 1.0269 | 0.00051644 | -1.6007e-05 | 0.059309
M16B | A | 8 | 1.0264 | 0.00046039 | -1.4932e-05 | 0.0585
M16B | A | 9 | 1.0279 | 0.00045435 | -1.5249e-05 | 0.061814
M16B | A | 10 | 1.0299 | 0.00035835 | -1.4213e-05 | 0.055948
M16B | A | 11 | 1.0256 | 0.00058477 | -1.6774e-05 | 0.067501
M16B | A | 12 | 1.0283 | 0.00042907 | -1.4953e-05 | 0.059845
M16B | A | 13 | 1.0285 | 0.00042059 | -1.4862e-05 | 0.065105
M16B | A | 14 | 1.0296 | 0.00031415 | -1.3399e-05 | 0.061008
M16B | A | 15 | 1.0272 | 0.00048864 | -1.5618e-05 | 0.062084
M16B | A | 16 | 1.0271 | 0.00045647 | -1.5067e-05 | 0.040904
M16B | B | 1 | 1.0302 | 0.00053494 | -1.7223e-05 | 0.065675
M16B | B | 2 | 1.0288 | 0.00058712 | -1.7717e-05 | 0.067088
M16B | B | 3 | 1.0302 | 0.00053134 | -1.7181e-05 | 0.065875
M16B | B | 4 | 1.0282 | 0.00058952 | -1.7587e-05 | 0.071866
M16B | B | 5 | 1.0319 | 0.00047492 | -1.6704e-05 | 0.060902
M16B | B | 6 | 1.0264 | 0.00071474 | -1.916e-05 | 0.061912
M16B | B | 7 | 1.0303 | 0.00056736 | -1.7804e-05 | 0.065477
M16B | B | 8 | 1.0289 | 0.00056506 | -1.7374e-05 | 0.066684
M16B | B | 9 | 1.0298 | 0.00056936 | -1.768e-05 | 0.072395
M16B | B | 10 | 1.0336 | 0.00037628 | -1.5521e-05 | 0.064735
M16B | B | 11 | 1.0278 | 0.00065028 | -1.8479e-05 | 0.060916
M16B | B | 12 | 1.0286 | 0.00058222 | -1.7563e-05 | 0.053219
M16B | B | 13 | 1.0305 | 0.00052459 | -1.7144e-05 | 0.074795
M16B | B | 14 | 1.0332 | 0.00033939 | -1.481e-05 | 0.067054
M16B | B | 15 | 1.0332 | 0.00039104 | -1.5668e-05 | 0.063283
M16B | B | 16 | 1.0309 | 0.00046756 | -1.6302e-05 | 0.053619

## RSB RVS
Band | HAM | Band Median |  |  |  | Band Average
|  | a0 | a1 | a2 | Uncertainty | a0 | a1 | a2 | Uncertainty
I1 | A | 1.019 | -0.04635 | 0.02659 | 0.11921 | 1.0187 | -0.045867 | 0.026629 | 0.11954
I1 | B | 1.0165 | -0.045192 | 0.028167 | 0.11542 | 1.0165 | -0.045554 | 0.028352 | 0.11511
I2 | A | 1.0048 | -0.030662 | 0.024765 | 0.11537 | 1.0048 | -0.030866 | 0.024923 | 0.11584
I2 | B | 1.013 | -0.028725 | 0.015623 | 0.11723 | 1.0129 | -0.028673 | 0.015583 | 0.11752
I3 | A | 1.0075 | -0.032728 | 0.025846 | 0.14131 | 1.0088 | -0.049326 | 0.038801 | 0.24375
I3 | B | 1.0033 | -0.031048 | 0.026464 | 0.15365 | 1.006 | -0.045086 | 0.0373 | 0.24156
M1 | A | 1.0131 | 0.011912 | -0.022657 | 0.28867 | 1.0128 | 0.012191 | -0.023042 | 0.29106
M1 | B | 0.98403 | 0.069632 | -0.0512 | 0.30095 | 0.98379 | 0.068358 | -0.050217 | 0.30186
M2 | A | 1.0134 | -0.0087675 | -0.0038374 | 0.21176 | 1.0135 | -0.0090374 | -0.0036004 | 0.212
M2 | B | 1.0001 | 0.024225 | -0.022693 | 0.21869 | 0.99932 | 0.024593 | -0.022693 | 0.21896
M3 | A | 1.0167 | -0.030828 | 0.014026 | 0.13635 | 1.0168 | -0.030345 | 0.0137 | 0.13655
M3 | B | 1.0073 | -0.0034066 | -0.0031754 | 0.142 | 1.0072 | -0.0036796 | -0.0029392 | 0.1407
M4 | A | 1.0183 | -0.045445 | 0.026705 | 0.12976 | 1.0182 | -0.045184 | 0.026431 | 0.13012
M4 | B | 1.0126 | -0.027138 | 0.014284 | 0.12912 | 1.0126 | -0.026849 | 0.014115 | 0.12932
M5 | A | 1.0202 | -0.050694 | 0.029313 | 0.1885 | 1.0201 | -0.049156 | 0.028488 | 0.18825
M5 | B | 1.019 | -0.053873 | 0.034336 | 0.18522 | 1.0191 | -0.054243 | 0.034281 | 0.18495
M6 | A | 1.0165 | -0.035471 | 0.01884 | 0.13931 | 1.0163 | -0.034949 | 0.018494 | 0.13946
M6 | B | 1.0189 | -0.048761 | 0.029371 | 0.13737 | 1.0188 | -0.049143 | 0.029674 | 0.13672
M7 | A | 1.0053 | -0.033754 | 0.027337 | 0.10983 | 1.0053 | -0.033665 | 0.027144 | 0.10979
M7 | B | 1.0133 | -0.032327 | 0.0186 | 0.11131 | 1.0132 | -0.03158 | 0.018077 | 0.1118
M8 | A | 1.002 | -0.034836 | 0.031366 | 0.090305 | 1.0019 | -0.033691 | 0.03024 | 0.089482
M8 | B | 0.99986 | -0.033055 | 0.031355 | 0.092034 | 0.99973 | -0.031922 | 0.030487 | 0.091144
M9 | A | 1.0261 | -0.10159 | 0.072729 | 0.22534 | 1.026 | -0.099913 | 0.071292 | 0.22568
M9 | B | 1.0244 | -0.099669 | 0.072517 | 0.22251 | 1.0241 | -0.098537 | 0.071696 | 0.22228
M10 | A | 1.0051 | -0.033684 | 0.027228 | 0.082891 | 1.0051 | -0.033471 | 0.027134 | 0.081774
M10 | B | 1.0033 | -0.031216 | 0.026904 | 0.082577 | 1.0035 | -0.031868 | 0.027096 | 0.082414
M11 | A | 1.0069 | -0.034688 | 0.026663 | 0.095013 | 1.0067 | -0.033698 | 0.025908 | 0.095781
M11 | B | 1.0057 | -0.032344 | 0.025789 | 0.094413 | 1.0058 | -0.032347 | 0.025476 | 0.094527
Band | HAM | Detector | Detector Dependent |  |  |  |  | Notes:
|  |  | a0 | a1 | a2 | Uncertainty |  | 1. This table provides VIIRS F2 RSB RVS.
2. The results were derived from FP-10 using the SIS-100 source.
I1 | A | 1 | 1.0191 | -0.045607 | 0.026098 | 0.11796 |  | 3. Quadratic fitting was performed on RVS vs HAM AOI.
I1 | A | 2 | 1.0196 | -0.047962 | 0.027879 | 0.12499 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | A | 3 | 1.0197 | -0.04674 | 0.026606 | 0.11607 |  | 5. The test was set to operational day mode.
I1 | A | 4 | 1.0202 | -0.049712 | 0.028951 | 0.11109
I1 | A | 5 | 1.019 | -0.04706 | 0.027534 | 0.11372
I1 | A | 6 | 1.0198 | -0.049769 | 0.029373 | 0.1132
I1 | A | 7 | 1.0205 | -0.050785 | 0.029731 | 0.12265
I1 | A | 8 | 1.0192 | -0.048837 | 0.028994 | 0.11866
I1 | A | 9 | 1.0193 | -0.048584 | 0.02867 | 0.11711
I1 | A | 10 | 1.0204 | -0.051175 | 0.030191 | 0.11798
I1 | A | 11 | 1.019 | -0.046007 | 0.026575 | 0.11713
I1 | A | 12 | 1.0184 | -0.045151 | 0.026218 | 0.11628
I1 | A | 13 | 1.0204 | -0.051165 | 0.030198 | 0.11856
I1 | A | 14 | 1.0188 | -0.044098 | 0.024941 | 0.12202
I1 | A | 15 | 1.0193 | -0.048536 | 0.028641 | 0.11883
I1 | A | 16 | 1.0193 | -0.047167 | 0.02733 | 0.1232
I1 | A | 17 | 1.0166 | -0.037396 | 0.020489 | 0.12087
I1 | A | 18 | 1.0191 | -0.045049 | 0.025547 | 0.12433
I1 | A | 19 | 1.019 | -0.046693 | 0.027208 | 0.12242
I1 | A | 20 | 1.0194 | -0.047657 | 0.02772 | 0.12071
I1 | A | 21 | 1.0192 | -0.047318 | 0.027558 | 0.1174
I1 | A | 22 | 1.018 | -0.042179 | 0.023786 | 0.12592
I1 | A | 23 | 1.0181 | -0.044274 | 0.025668 | 0.12198
I1 | A | 24 | 1.0175 | -0.042122 | 0.024244 | 0.11935
I1 | A | 25 | 1.018 | -0.043937 | 0.025513 | 0.11706
I1 | A | 26 | 1.0177 | -0.043017 | 0.024882 | 0.11907
I1 | A | 27 | 1.0173 | -0.043077 | 0.02525 | 0.11176
I1 | A | 28 | 1.0171 | -0.041241 | 0.023736 | 0.12884
I1 | A | 29 | 1.019 | -0.047684 | 0.028149 | 0.12174
I1 | A | 30 | 1.0174 | -0.042368 | 0.024551 | 0.12153
I1 | A | 31 | 1.0172 | -0.042126 | 0.024482 | 0.12243
I1 | A | 32 | 1.0174 | -0.043263 | 0.025399 | 0.12043
I1 | B | 1 | 1.0163 | -0.044517 | 0.027569 | 0.11028
I1 | B | 2 | 1.0177 | -0.048823 | 0.030344 | 0.11577
I1 | B | 3 | 1.0167 | -0.045072 | 0.027688 | 0.1113
I1 | B | 4 | 1.0165 | -0.045312 | 0.028082 | 0.11064
I1 | B | 5 | 1.017 | -0.04759 | 0.02984 | 0.11038
I1 | B | 6 | 1.0169 | -0.047333 | 0.029709 | 0.11368
I1 | B | 7 | 1.0173 | -0.048676 | 0.030558 | 0.11178
I1 | B | 8 | 1.0171 | -0.047835 | 0.02995 | 0.11683
I1 | B | 9 | 1.0183 | -0.051487 | 0.032335 | 0.11078
I1 | B | 10 | 1.0165 | -0.046045 | 0.028835 | 0.11791
I1 | B | 11 | 1.0169 | -0.046436 | 0.028824 | 0.11373
I1 | B | 12 | 1.0173 | -0.048471 | 0.030384 | 0.12073
I1 | B | 13 | 1.0167 | -0.046391 | 0.028926 | 0.11855
I1 | B | 14 | 1.0163 | -0.043274 | 0.026367 | 0.12266
I1 | B | 15 | 1.0168 | -0.046718 | 0.029211 | 0.11559
I1 | B | 16 | 1.0179 | -0.049295 | 0.030656 | 0.11197
I1 | B | 17 | 1.0164 | -0.042637 | 0.02565 | 0.11555
I1 | B | 18 | 1.0164 | -0.042976 | 0.026022 | 0.11647
I1 | B | 19 | 1.0169 | -0.047246 | 0.029567 | 0.11588
I1 | B | 20 | 1.0163 | -0.044815 | 0.027832 | 0.11832
I1 | B | 21 | 1.0154 | -0.042169 | 0.026159 | 0.12036
I1 | B | 22 | 1.0166 | -0.044399 | 0.027195 | 0.11222
I1 | B | 23 | 1.0152 | -0.042392 | 0.026507 | 0.1161
I1 | B | 24 | 1.0161 | -0.043823 | 0.027067 | 0.11832
I1 | B | 25 | 1.0159 | -0.044327 | 0.027738 | 0.11528
I1 | B | 26 | 1.0165 | -0.045367 | 0.028208 | 0.1145
I1 | B | 27 | 1.0156 | -0.043438 | 0.02719 | 0.11262
I1 | B | 28 | 1.0157 | -0.043883 | 0.027467 | 0.11401
I1 | B | 29 | 1.0156 | -0.044461 | 0.028149 | 0.11249
I1 | B | 30 | 1.0151 | -0.041147 | 0.02546 | 0.11941
I1 | B | 31 | 1.0158 | -0.044756 | 0.028184 | 0.11573
I1 | B | 32 | 1.0163 | -0.046616 | 0.02958 | 0.11374
I2 | A | 1 | 1.0043 | -0.027693 | 0.022345 | 0.11354
I2 | A | 2 | 1.0047 | -0.029277 | 0.023509 | 0.11434
I2 | A | 3 | 1.0047 | -0.02907 | 0.023304 | 0.11336
I2 | A | 4 | 1.0058 | -0.032781 | 0.025844 | 0.11515
I2 | A | 5 | 1.0062 | -0.034343 | 0.026966 | 0.1118
I2 | A | 6 | 1.0035 | -0.025972 | 0.021459 | 0.11772
I2 | A | 7 | 1.0053 | -0.031186 | 0.024779 | 0.1144
I2 | A | 8 | 1.0048 | -0.030168 | 0.024246 | 0.11538
I2 | A | 9 | 1.0035 | -0.02712 | 0.022531 | 0.11536
I2 | A | 10 | 1.0059 | -0.033939 | 0.026868 | 0.11739
I2 | A | 11 | 1.0059 | -0.033917 | 0.026846 | 0.1182
I2 | A | 12 | 1.0053 | -0.032229 | 0.025797 | 0.11701
I2 | A | 13 | 1.0029 | -0.024212 | 0.020336 | 0.1141
I2 | A | 14 | 1.005 | -0.029789 | 0.023727 | 0.11328
I2 | A | 15 | 1.0045 | -0.030396 | 0.024751 | 0.11276
I2 | A | 16 | 1.0045 | -0.029849 | 0.024268 | 0.11489
I2 | A | 17 | 1.0061 | -0.034519 | 0.02723 | 0.11403
I2 | A | 18 | 1.0043 | -0.029563 | 0.024188 | 0.11048
I2 | A | 19 | 1.0039 | -0.028374 | 0.02335 | 0.11352
I2 | A | 20 | 1.0054 | -0.032782 | 0.026235 | 0.11452
I2 | A | 21 | 1.0045 | -0.030119 | 0.02451 | 0.11559
I2 | A | 22 | 1.0053 | -0.032345 | 0.025857 | 0.11559
I2 | A | 23 | 1.0054 | -0.032913 | 0.026375 | 0.11881
I2 | A | 24 | 1.0042 | -0.029969 | 0.024597 | 0.1167
I2 | A | 25 | 1.004 | -0.029168 | 0.024073 | 0.12064
I2 | A | 26 | 1.0055 | -0.033378 | 0.026683 | 0.12201
I2 | A | 27 | 1.0048 | -0.032368 | 0.026332 | 0.11776
I2 | A | 28 | 1.0047 | -0.030927 | 0.025089 | 0.12203
I2 | A | 29 | 1.0048 | -0.032102 | 0.026123 | 0.11672
I2 | A | 30 | 1.0045 | -0.032166 | 0.026419 | 0.11732
I2 | A | 31 | 1.0037 | -0.028838 | 0.024018 | 0.11747
I2 | A | 32 | 1.0061 | -0.036244 | 0.028898 | 0.11498
I2 | B | 1 | 1.013 | -0.028201 | 0.015043 | 0.11562
I2 | B | 2 | 1.012 | -0.025495 | 0.013401 | 0.11666
I2 | B | 3 | 1.0122 | -0.025503 | 0.013186 | 0.11985
I2 | B | 4 | 1.0131 | -0.028578 | 0.015315 | 0.11885
I2 | B | 5 | 1.0124 | -0.026456 | 0.013934 | 0.11669
I2 | B | 6 | 1.0123 | -0.026731 | 0.014274 | 0.114
I2 | B | 7 | 1.0124 | -0.027238 | 0.014659 | 0.11672
I2 | B | 8 | 1.0132 | -0.029523 | 0.016106 | 0.11132
I2 | B | 9 | 1.0129 | -0.02869 | 0.015627 | 0.11701
I2 | B | 10 | 1.0132 | -0.029571 | 0.016146 | 0.11171
I2 | B | 11 | 1.0134 | -0.030051 | 0.016477 | 0.11782
I2 | B | 12 | 1.013 | -0.028759 | 0.015619 | 0.11849
I2 | B | 13 | 1.0127 | -0.026631 | 0.013853 | 0.11537
I2 | B | 14 | 1.0121 | -0.026335 | 0.014071 | 0.11673
I2 | B | 15 | 1.0132 | -0.029491 | 0.016102 | 0.12309
I2 | B | 16 | 1.0137 | -0.030586 | 0.016698 | 0.12116
I2 | B | 17 | 1.0128 | -0.027148 | 0.014275 | 0.11968
I2 | B | 18 | 1.0137 | -0.030778 | 0.016862 | 0.11887
I2 | B | 19 | 1.0139 | -0.031305 | 0.017174 | 0.11604
I2 | B | 20 | 1.0119 | -0.02616 | 0.014129 | 0.11911
I2 | B | 21 | 1.0125 | -0.026863 | 0.014204 | 0.12057
I2 | B | 22 | 1.0128 | -0.02812 | 0.015157 | 0.11682
I2 | B | 23 | 1.0133 | -0.030361 | 0.016824 | 0.11951
I2 | B | 24 | 1.0132 | -0.030222 | 0.016786 | 0.11225
I2 | B | 25 | 1.0131 | -0.029431 | 0.016107 | 0.11744
I2 | B | 26 | 1.0139 | -0.032279 | 0.018109 | 0.11795
I2 | B | 27 | 1.0137 | -0.031204 | 0.017304 | 0.12222
I2 | B | 28 | 1.0118 | -0.026327 | 0.014371 | 0.11959
I2 | B | 29 | 1.0126 | -0.029237 | 0.016426 | 0.11652
I2 | B | 30 | 1.0134 | -0.030624 | 0.016959 | 0.1213
I2 | B | 31 | 1.012 | -0.027403 | 0.015219 | 0.11629
I2 | B | 32 | 1.0137 | -0.03222 | 0.018237 | 0.11532
I3 | A | 1 | 1.0035 | -0.025739 | 0.021252 | 0.15022
I3 | A | 2 | 1.0049 | -0.031109 | 0.025104 | 0.14974
I3 | A | 3 | 1.0041 | -0.02786 | 0.022728 | 0.14916
I3 | A | 4 | 1.1306 | -0.57975 | 0.43208 | 3.0451
I3 | A | 5 | 1.0034 | -0.02606 | 0.021677 | 0.14568
I3 | A | 6 | 1.0044 | -0.02918 | 0.023672 | 0.1496
I3 | A | 7 | 1.0052 | -0.031266 | 0.024915 | 0.15049
I3 | A | 8 | 1.0054 | -0.033486 | 0.026852 | 0.14689
I3 | A | 9 | 1.0034 | -0.026571 | 0.022108 | 0.15177
I3 | A | 10 | 1.0038 | -0.029175 | 0.024269 | 0.15164
I3 | A | 11 | 1.0038 | -0.028129 | 0.023227 | 0.15358
I3 | A | 12 | 1.0073 | -0.038803 | 0.030244 | 0.14687
I3 | A | 13 | 1.0041 | -0.02873 | 0.023503 | 0.14926
I3 | A | 14 | 1.0052 | -0.03297 | 0.02656 | 0.15196
I3 | A | 15 | 1.0053 | -0.031837 | 0.025376 | 0.1511
I3 | A | 16 | 1.0056 | -0.035754 | 0.028821 | 0.14575
I3 | A | 17 | 1.0053 | -0.032113 | 0.025681 | 0.15372
I3 | A | 18 | 1.0071 | -0.039023 | 0.030644 | 0.15175
I3 | A | 19 | 1.005 | -0.031764 | 0.025594 | 0.1542
I3 | A | 20 | 1.0044 | -0.033007 | 0.027298 | 0.15741
I3 | A | 21 | 1.0052 | -0.033752 | 0.027348 | 0.15127
I3 | A | 22 | 1.0062 | -0.037841 | 0.030294 | 0.15506
I3 | A | 23 | 1.0046 | -0.031486 | 0.025695 | 0.15266
I3 | A | 24 | 1.0064 | -0.037608 | 0.029927 | 0.15214
I3 | A | 25 | 1.0045 | -0.031324 | 0.025666 | 0.16131
I3 | A | 26 | 1.0065 | -0.039666 | 0.03171 | 0.16214
I3 | A | 27 | 1.0049 | -0.033117 | 0.02702 | 0.15621
I3 | A | 28 | 1.0047 | -0.032092 | 0.026175 | 0.15554
I3 | A | 29 | 1.0051 | -0.033301 | 0.026996 | 0.16146
I3 | A | 30 | 1.0052 | -0.034317 | 0.027884 | 0.16025
I3 | A | 31 | 1.0033 | -0.02806 | 0.023582 | 0.16404
I3 | A | 32 | 1.0045 | -0.033546 | 0.027744 | 0.16199
I3 | B | 1 | 1.0013 | -0.021802 | 0.019447 | 0.14639
I3 | B | 2 | 1.0037 | -0.030655 | 0.025703 | 0.14812
I3 | B | 3 | 1.0015 | -0.022976 | 0.020446 | 0.15277
I3 | B | 4 | 1.0975 | -0.51545 | 0.40088 | 2.9563
I3 | B | 5 | 1.0028 | -0.026663 | 0.022722 | 0.16009
I3 | B | 6 | 1.0036 | -0.030731 | 0.025846 | 0.1511
I3 | B | 7 | 1.0011 | -0.02185 | 0.019728 | 0.15006
I3 | B | 8 | 1.0037 | -0.0316 | 0.026651 | 0.15993
I3 | B | 9 | 1.0015 | -0.024321 | 0.021669 | 0.15033
I3 | B | 10 | 1.0017 | -0.026912 | 0.023928 | 0.15345
I3 | B | 11 | 1.0048 | -0.03268 | 0.026688 | 0.14481
I3 | B | 12 | 1.0036 | -0.03189 | 0.026947 | 0.14803
I3 | B | 13 | 1.0021 | -0.02633 | 0.023064 | 0.15279
I3 | B | 14 | 1.004 | -0.032559 | 0.027287 | 0.15056
I3 | B | 15 | 1.0041 | -0.032124 | 0.026751 | 0.15555
I3 | B | 16 | 1.0034 | -0.030989 | 0.026288 | 0.15851
I3 | B | 17 | 1.0037 | -0.03047 | 0.025511 | 0.15563
I3 | B | 18 | 1.0044 | -0.035725 | 0.029876 | 0.15656
I3 | B | 19 | 1.0028 | -0.029484 | 0.025415 | 0.15185
I3 | B | 20 | 1.0021 | -0.027926 | 0.024611 | 0.14968
I3 | B | 21 | 1.0047 | -0.033541 | 0.027582 | 0.15185
I3 | B | 22 | 1.0031 | -0.031803 | 0.027387 | 0.15532
I3 | B | 23 | 1.0018 | -0.026697 | 0.023687 | 0.15709
I3 | B | 24 | 1.0041 | -0.034344 | 0.028901 | 0.15953
I3 | B | 25 | 1.0039 | -0.032335 | 0.027104 | 0.1594
I3 | B | 26 | 1.0031 | -0.031852 | 0.027411 | 0.15891
I3 | B | 27 | 1.0027 | -0.029501 | 0.0255 | 0.14932
I3 | B | 28 | 1.0044 | -0.034892 | 0.029145 | 0.15825
I3 | B | 29 | 1.0023 | -0.027507 | 0.024024 | 0.1506
I3 | B | 30 | 1.0034 | -0.033842 | 0.028995 | 0.16191
I3 | B | 31 | 1.0032 | -0.031107 | 0.026641 | 0.15386
I3 | B | 32 | 1.003 | -0.032179 | 0.027777 | 0.16121
M1 | A | 1 | 1.0085 | 0.02475 | -0.031074 | 0.29045
M1 | A | 2 | 1.0124 | 0.012249 | -0.022771 | 0.29901
M1 | A | 3 | 1.0127 | 0.011574 | -0.022372 | 0.28869
M1 | A | 4 | 1.0089 | 0.021415 | -0.028276 | 0.28492
M1 | A | 5 | 1.0109 | 0.016543 | -0.02544 | 0.2991
M1 | A | 6 | 1.0113 | 0.015978 | -0.025265 | 0.27999
M1 | A | 7 | 1.0143 | 0.0068825 | -0.019331 | 0.2968
M1 | A | 8 | 1.0146 | 0.0073065 | -0.020058 | 0.2861
M1 | A | 9 | 1.0115 | 0.016052 | -0.025542 | 0.28671
M1 | A | 10 | 1.0149 | 0.0058461 | -0.018916 | 0.28225
M1 | A | 11 | 1.0155 | 0.0036277 | -0.017352 | 0.28352
M1 | A | 12 | 1.0158 | 0.0042717 | -0.018229 | 0.28865
M1 | A | 13 | 1.0135 | 0.0095759 | -0.021213 | 0.28994
M1 | A | 14 | 1.0139 | 0.010644 | -0.022542 | 0.28218
M1 | A | 15 | 1.014 | 0.013398 | -0.025247 | 0.32634
M1 | A | 16 | 1.0121 | 0.014938 | -0.025036 | 0.29226
M1 | B | 1 | 0.97672 | 0.082667 | -0.057427 | 0.29497
M1 | B | 2 | 0.98111 | 0.072345 | -0.051591 | 0.30722
M1 | B | 3 | 0.98049 | 0.076886 | -0.05533 | 0.31745
M1 | B | 4 | 0.98508 | 0.061906 | -0.045261 | 0.30695
M1 | B | 5 | 0.98127 | 0.074234 | -0.053517 | 0.30038
M1 | B | 6 | 0.98213 | 0.072223 | -0.052391 | 0.30252
M1 | B | 7 | 0.98005 | 0.081143 | -0.058971 | 0.3085
M1 | B | 8 | 0.98284 | 0.071642 | -0.052473 | 0.29564
M1 | B | 9 | 0.98319 | 0.071293 | -0.052462 | 0.29455
M1 | B | 10 | 0.9851 | 0.064936 | -0.048148 | 0.30153
M1 | B | 11 | 0.98486 | 0.067971 | -0.050808 | 0.30292
M1 | B | 12 | 0.98897 | 0.054966 | -0.042181 | 0.29608
M1 | B | 13 | 0.98792 | 0.058643 | -0.044721 | 0.30859
M1 | B | 14 | 0.98859 | 0.057093 | -0.043853 | 0.29975
M1 | B | 15 | 0.98574 | 0.065064 | -0.048847 | 0.29652
M1 | B | 16 | 0.98659 | 0.060714 | -0.04549 | 0.29617
M2 | A | 1 | 1.0131 | -0.0083811 | -0.0038421 | 0.2167
M2 | A | 2 | 1.0124 | -0.0063711 | -0.0050614 | 0.2155
M2 | A | 3 | 1.0137 | -0.0094353 | -0.0033586 | 0.21589
M2 | A | 4 | 1.0126 | -0.0058358 | -0.0057547 | 0.21097
M2 | A | 5 | 1.0143 | -0.0094531 | -0.0038444 | 0.21344
M2 | A | 6 | 1.012 | -0.0048698 | -0.0061619 | 0.21146
M2 | A | 7 | 1.0136 | -0.0088478 | -0.0038326 | 0.20681
M2 | A | 8 | 1.0154 | -0.013727 | -0.00081058 | 0.20917
M2 | A | 9 | 1.0139 | -0.010206 | -0.0028432 | 0.21283
M2 | A | 10 | 1.013 | -0.0070378 | -0.0050347 | 0.21194
M2 | A | 11 | 1.0133 | -0.0080124 | -0.0043267 | 0.21158
M2 | A | 12 | 1.0163 | -0.017323 | 0.0017607 | 0.21055
M2 | A | 13 | 1.0141 | -0.010199 | -0.0029686 | 0.20567
M2 | A | 14 | 1.0137 | -0.0095209 | -0.0032493 | 0.21905
M2 | A | 15 | 1.0132 | -0.0086872 | -0.0035914 | 0.21273
M2 | A | 16 | 1.0123 | -0.0066911 | -0.0046878 | 0.20777
M2 | B | 1 | 0.99619 | 0.030124 | -0.025121 | 0.21627
M2 | B | 2 | 0.99626 | 0.030008 | -0.025073 | 0.22014
M2 | B | 3 | 0.99679 | 0.029054 | -0.024649 | 0.2205
M2 | B | 4 | 0.99675 | 0.030069 | -0.025571 | 0.21894
M2 | B | 5 | 0.99858 | 0.026118 | -0.02347 | 0.21254
M2 | B | 6 | 1.0007 | 0.020282 | -0.019866 | 0.22378
M2 | B | 7 | 1.0007 | 0.020748 | -0.02032 | 0.22425
M2 | B | 8 | 0.9993 | 0.024703 | -0.022776 | 0.217
M2 | B | 9 | 1 | 0.023829 | -0.02261 | 0.21561
M2 | B | 10 | 1.0003 | 0.02183 | -0.020951 | 0.21733
M2 | B | 11 | 1.0008 | 0.022468 | -0.022025 | 0.21843
M2 | B | 12 | 1.0021 | 0.019305 | -0.020188 | 0.22691
M2 | B | 13 | 0.99904 | 0.025907 | -0.023689 | 0.21382
M2 | B | 14 | 1.0003 | 0.024622 | -0.023588 | 0.22322
M2 | B | 15 | 1.0003 | 0.022758 | -0.021789 | 0.21292
M2 | B | 16 | 1.001 | 0.021664 | -0.021404 | 0.2217
M3 | A | 1 | 1.0168 | -0.030576 | 0.013911 | 0.13445
M3 | A | 2 | 1.0157 | -0.027181 | 0.011625 | 0.13733
M3 | A | 3 | 1.0163 | -0.028678 | 0.01256 | 0.13064
M3 | A | 4 | 1.0165 | -0.028956 | 0.01259 | 0.13358
M3 | A | 5 | 1.0165 | -0.028833 | 0.012536 | 0.13331
M3 | A | 6 | 1.0163 | -0.028346 | 0.012266 | 0.14035
M3 | A | 7 | 1.0167 | -0.030107 | 0.013546 | 0.13238
M3 | A | 8 | 1.0171 | -0.031097 | 0.01414 | 0.13639
M3 | A | 9 | 1.0179 | -0.032569 | 0.01479 | 0.13965
M3 | A | 10 | 1.017 | -0.03108 | 0.014218 | 0.1395
M3 | A | 11 | 1.0172 | -0.031717 | 0.014621 | 0.1418
M3 | A | 12 | 1.018 | -0.033536 | 0.015631 | 0.13862
M3 | A | 13 | 1.0162 | -0.029354 | 0.013279 | 0.13309
M3 | A | 14 | 1.0168 | -0.031113 | 0.01442 | 0.13631
M3 | A | 15 | 1.0167 | -0.031187 | 0.014558 | 0.1355
M3 | A | 16 | 1.0168 | -0.03119 | 0.014502 | 0.14187
M3 | B | 1 | 1.0057 | -0.0010809 | -0.0040874 | 0.14266
M3 | B | 2 | 1.0057 | -0.00072123 | -0.0044491 | 0.13576
M3 | B | 3 | 1.0061 | -0.0016431 | -0.0039264 | 0.14221
M3 | B | 4 | 1.0068 | -0.0032894 | -0.0030091 | 0.13834
M3 | B | 5 | 1.0067 | -0.0029255 | -0.0032796 | 0.1394
M3 | B | 6 | 1.0066 | -0.0020475 | -0.0040146 | 0.13601
M3 | B | 7 | 1.0071 | -0.0030637 | -0.003472 | 0.14277
M3 | B | 8 | 1.0081 | -0.0055833 | -0.001949 | 0.14225
M3 | B | 9 | 1.0075 | -0.0035238 | -0.0033732 | 0.14296
M3 | B | 10 | 1.0081 | -0.0059898 | -0.0015577 | 0.13716
M3 | B | 11 | 1.008 | -0.0059141 | -0.0015736 | 0.14157
M3 | B | 12 | 1.0075 | -0.0038696 | -0.0030712 | 0.1425
M3 | B | 13 | 1.0075 | -0.0042519 | -0.0027011 | 0.14226
M3 | B | 14 | 1.0077 | -0.0048215 | -0.0023072 | 0.1396
M3 | B | 15 | 1.0084 | -0.0069278 | -0.00095718 | 0.14392
M3 | B | 16 | 1.0071 | -0.0032199 | -0.0032985 | 0.1418
M4 | A | 1 | 1.0179 | -0.042874 | 0.024578 | 0.12607
M4 | A | 2 | 1.0174 | -0.040602 | 0.022867 | 0.13526
M4 | A | 3 | 1.0191 | -0.046303 | 0.026706 | 0.12714
M4 | A | 4 | 1.0161 | -0.038096 | 0.02164 | 0.13219
M4 | A | 5 | 1.0193 | -0.047367 | 0.027515 | 0.12596
M4 | A | 6 | 1.0191 | -0.047228 | 0.027569 | 0.13
M4 | A | 7 | 1.0177 | -0.043105 | 0.024913 | 0.12868
M4 | A | 8 | 1.0185 | -0.045231 | 0.026216 | 0.13088
M4 | A | 9 | 1.0199 | -0.049439 | 0.028984 | 0.13068
M4 | A | 10 | 1.0181 | -0.045015 | 0.02636 | 0.12953
M4 | A | 11 | 1.0184 | -0.04566 | 0.026705 | 0.13774
M4 | A | 12 | 1.0176 | -0.044353 | 0.026221 | 0.12911
M4 | A | 13 | 1.0186 | -0.047154 | 0.02798 | 0.13142
M4 | A | 14 | 1.0179 | -0.045782 | 0.027343 | 0.13061
M4 | A | 15 | 1.0191 | -0.049984 | 0.030173 | 0.12758
M4 | A | 16 | 1.017 | -0.044759 | 0.027129 | 0.12911
M4 | B | 1 | 1.011 | -0.022524 | 0.011441 | 0.12863
M4 | B | 2 | 1.0114 | -0.02281 | 0.011401 | 0.12859
M4 | B | 3 | 1.0121 | -0.024293 | 0.012181 | 0.13343
M4 | B | 4 | 1.0126 | -0.026476 | 0.013762 | 0.12604
M4 | B | 5 | 1.0124 | -0.025925 | 0.013393 | 0.12929
M4 | B | 6 | 1.0126 | -0.025947 | 0.013282 | 0.13216
M4 | B | 7 | 1.0134 | -0.028402 | 0.01484 | 0.13282
M4 | B | 8 | 1.013 | -0.027411 | 0.014308 | 0.12687
M4 | B | 9 | 1.0138 | -0.029726 | 0.015777 | 0.13067
M4 | B | 10 | 1.0131 | -0.027828 | 0.014597 | 0.13234
M4 | B | 11 | 1.0118 | -0.024534 | 0.01268 | 0.13004
M4 | B | 12 | 1.0125 | -0.026864 | 0.01426 | 0.13185
M4 | B | 13 | 1.0138 | -0.030547 | 0.016572 | 0.12536
M4 | B | 14 | 1.0132 | -0.029501 | 0.016146 | 0.12455
M4 | B | 15 | 1.013 | -0.029127 | 0.015965 | 0.12758
M4 | B | 16 | 1.0122 | -0.027667 | 0.015231 | 0.12895
M5 | A | 1 | 1.0188 | -0.043291 | 0.024172 | 0.18613
M5 | A | 2 | 1.0181 | -0.042525 | 0.024061 | 0.19026
M5 | A | 3 | 1.0221 | -0.052335 | 0.029775 | 0.18622
M5 | A | 4 | 1.0209 | -0.05103 | 0.029544 | 0.18949
M5 | A | 5 | 1.0189 | -0.045941 | 0.026585 | 0.18324
M5 | A | 6 | 1.0209 | -0.050486 | 0.029082 | 0.18916
M5 | A | 7 | 1.0195 | -0.04627 | 0.026314 | 0.19351
M5 | A | 8 | 1.0198 | -0.047267 | 0.026966 | 0.19098
M5 | A | 9 | 1.0216 | -0.05323 | 0.031038 | 0.19158
M5 | A | 10 | 1.0203 | -0.051328 | 0.030451 | 0.1812
M5 | A | 11 | 1.022 | -0.055391 | 0.03271 | 0.18713
M5 | A | 12 | 1.0186 | -0.045882 | 0.026761 | 0.18785
M5 | A | 13 | 1.021 | -0.052639 | 0.030995 | 0.18229
M5 | A | 14 | 1.0192 | -0.047026 | 0.027298 | 0.19144
M5 | A | 15 | 1.0204 | -0.050903 | 0.029936 | 0.18404
M5 | A | 16 | 1.0202 | -0.050959 | 0.030113 | 0.1975
M5 | B | 1 | 1.017 | -0.048097 | 0.030311 | 0.17392
M5 | B | 2 | 1.0164 | -0.045248 | 0.028128 | 0.18639
M5 | B | 3 | 1.0171 | -0.047718 | 0.029873 | 0.19641
M5 | B | 4 | 1.0188 | -0.052699 | 0.033045 | 0.17857
M5 | B | 5 | 1.0209 | -0.05834 | 0.036556 | 0.19013
M5 | B | 6 | 1.0188 | -0.052329 | 0.032713 | 0.1773
M5 | B | 7 | 1.0209 | -0.060118 | 0.038166 | 0.18153
M5 | B | 8 | 1.0185 | -0.050891 | 0.031637 | 0.19515
M5 | B | 9 | 1.0223 | -0.063519 | 0.040178 | 0.17679
M5 | B | 10 | 1.0191 | -0.054103 | 0.034094 | 0.18142
M5 | B | 11 | 1.0205 | -0.058157 | 0.036743 | 0.19246
M5 | B | 12 | 1.0183 | -0.053508 | 0.034245 | 0.18241
M5 | B | 13 | 1.0194 | -0.055321 | 0.034998 | 0.18815
M5 | B | 14 | 1.0194 | -0.056161 | 0.035785 | 0.18809
M5 | B | 15 | 1.0183 | -0.053642 | 0.034427 | 0.18447
M5 | B | 16 | 1.0194 | -0.058038 | 0.037596 | 0.18597
M6 | A | 1 | 1.016 | -0.033486 | 0.017339 | 0.13947
M6 | A | 2 | 1.0158 | -0.033086 | 0.017126 | 0.13722
M6 | A | 3 | 1.0135 | -0.027457 | 0.013861 | 0.14192
M6 | A | 4 | 1.0164 | -0.033813 | 0.017357 | 0.13753
M6 | A | 5 | 1.0164 | -0.035199 | 0.0186 | 0.1371
M6 | A | 6 | 1.0165 | -0.036567 | 0.019814 | 0.13715
M6 | A | 7 | 1.016 | -0.034108 | 0.017952 | 0.13914
M6 | A | 8 | 1.0172 | -0.037537 | 0.020114 | 0.1377
M6 | A | 9 | 1.0166 | -0.035568 | 0.01877 | 0.14077
M6 | A | 10 | 1.017 | -0.037469 | 0.02023 | 0.13652
M6 | A | 11 | 1.0169 | -0.035924 | 0.01891 | 0.1404
M6 | A | 12 | 1.0165 | -0.035905 | 0.019234 | 0.13952
M6 | A | 13 | 1.0168 | -0.036469 | 0.019496 | 0.142
M6 | A | 14 | 1.0172 | -0.03723 | 0.019821 | 0.14229
M6 | A | 15 | 1.0161 | -0.035374 | 0.019051 | 0.13792
M6 | A | 16 | 1.0156 | -0.033989 | 0.018235 | 0.14473
M6 | B | 1 | 1.0194 | -0.049194 | 0.029218 | 0.13614
M6 | B | 2 | 1.0189 | -0.048472 | 0.028976 | 0.13936
M6 | B | 3 | 1.0191 | -0.049495 | 0.029709 | 0.13808
M6 | B | 4 | 1.0177 | -0.045839 | 0.027542 | 0.1368
M6 | B | 5 | 1.0195 | -0.05106 | 0.030888 | 0.1318
M6 | B | 6 | 1.0181 | -0.046405 | 0.027749 | 0.13795
M6 | B | 7 | 1.019 | -0.04868 | 0.029065 | 0.13155
M6 | B | 8 | 1.0191 | -0.049569 | 0.029804 | 0.13919
M6 | B | 9 | 1.0202 | -0.05347 | 0.032514 | 0.13448
M6 | B | 10 | 1.0202 | -0.052698 | 0.031782 | 0.13549
M6 | B | 11 | 1.0186 | -0.048842 | 0.029608 | 0.13852
M6 | B | 12 | 1.0179 | -0.047349 | 0.028793 | 0.13378
M6 | B | 13 | 1.0196 | -0.051943 | 0.031654 | 0.13859
M6 | B | 14 | 1.0183 | -0.048505 | 0.029506 | 0.14104
M6 | B | 15 | 1.0178 | -0.047751 | 0.029236 | 0.13897
M6 | B | 16 | 1.0176 | -0.04701 | 0.028743 | 0.13577
M7 | A | 1 | 1.0045 | -0.02882 | 0.023312 | 0.10969
M7 | A | 2 | 1.0052 | -0.031414 | 0.025064 | 0.10996
M7 | A | 3 | 1.0055 | -0.033045 | 0.026383 | 0.10837
M7 | A | 4 | 1.0058 | -0.033784 | 0.026777 | 0.10818
M7 | A | 5 | 1.0053 | -0.032771 | 0.026325 | 0.10781
M7 | A | 6 | 1.005 | -0.03204 | 0.025847 | 0.10428
M7 | A | 7 | 1.0057 | -0.034012 | 0.027149 | 0.10475
M7 | A | 8 | 1.0054 | -0.03465 | 0.028013 | 0.1135
M7 | A | 9 | 1.0056 | -0.034856 | 0.027995 | 0.10785
M7 | A | 10 | 1.0051 | -0.033474 | 0.027098 | 0.11066
M7 | A | 11 | 1.0049 | -0.033709 | 0.027575 | 0.11467
M7 | A | 12 | 1.0053 | -0.034955 | 0.028366 | 0.11016
M7 | A | 13 | 1.0058 | -0.036283 | 0.029185 | 0.11132
M7 | A | 14 | 1.0049 | -0.033724 | 0.027526 | 0.11262
M7 | A | 15 | 1.0058 | -0.036544 | 0.02941 | 0.11358
M7 | A | 16 | 1.005 | -0.034564 | 0.028282 | 0.1093
M7 | B | 1 | 1.0125 | -0.027245 | 0.014601 | 0.10945
M7 | B | 2 | 1.0131 | -0.030046 | 0.016698 | 0.1092
M7 | B | 3 | 1.0132 | -0.030398 | 0.01693 | 0.11164
M7 | B | 4 | 1.0124 | -0.028862 | 0.016231 | 0.11451
M7 | B | 5 | 1.0135 | -0.031626 | 0.017835 | 0.10734
M7 | B | 6 | 1.0124 | -0.028933 | 0.016315 | 0.10953
M7 | B | 7 | 1.0122 | -0.028889 | 0.016423 | 0.11027
M7 | B | 8 | 1.0135 | -0.032227 | 0.018427 | 0.10965
M7 | B | 9 | 1.0135 | -0.032689 | 0.018876 | 0.11331
M7 | B | 10 | 1.0141 | -0.03486 | 0.0204 | 0.10845
M7 | B | 11 | 1.0131 | -0.032427 | 0.01896 | 0.11143
M7 | B | 12 | 1.0135 | -0.033297 | 0.019405 | 0.11226
M7 | B | 13 | 1.0138 | -0.033837 | 0.019692 | 0.11483
M7 | B | 14 | 1.0133 | -0.03245 | 0.018772 | 0.11616
M7 | B | 15 | 1.0139 | -0.034438 | 0.020159 | 0.11952
M7 | B | 16 | 1.0132 | -0.033064 | 0.019507 | 0.11119
M8 | A | 1 | 1.0005 | -0.025032 | 0.023312 | 0.082223
M8 | A | 2 | 1.0009 | -0.028024 | 0.025715 | 0.084151
M8 | A | 3 | 1.001 | -0.028779 | 0.026358 | 0.084201
M8 | A | 4 | 1.0017 | -0.031501 | 0.028313 | 0.0844
M8 | A | 5 | 1.0022 | -0.03287 | 0.029187 | 0.087184
M8 | A | 6 | 1.0024 | -0.033911 | 0.03002 | 0.084641
M8 | A | 7 | 1.0027 | -0.035926 | 0.031641 | 0.090047
M8 | A | 8 | 1.0023 | -0.035269 | 0.031382 | 0.089335
M8 | A | 9 | 1.002 | -0.035048 | 0.031457 | 0.09255
M8 | A | 10 | 1.003 | -0.038414 | 0.033726 | 0.091692
M8 | A | 11 | 1.0018 | -0.034674 | 0.031206 | 0.09483
M8 | A | 12 | 1.0015 | -0.034478 | 0.03135 | 0.091598
M8 | A | 13 | 1.0019 | -0.034998 | 0.03149 | 0.094795
M8 | A | 14 | 1.0021 | -0.036172 | 0.032399 | 0.090563
M8 | A | 15 | 1.0021 | -0.036827 | 0.033002 | 0.094062
M8 | A | 16 | 1.0021 | -0.037129 | 0.03328 | 0.095443
M8 | B | 1 | 0.9979 | -0.022776 | 0.023464 | 0.084255
M8 | B | 2 | 0.99917 | -0.027282 | 0.026599 | 0.084835
M8 | B | 3 | 0.99924 | -0.028448 | 0.027639 | 0.083682
M8 | B | 4 | 0.99914 | -0.028226 | 0.027516 | 0.085054
M8 | B | 5 | 0.99981 | -0.030671 | 0.029229 | 0.086724
M8 | B | 6 | 0.9999 | -0.031493 | 0.029927 | 0.088021
M8 | B | 7 | 1.0001 | -0.032912 | 0.031123 | 0.0921
M8 | B | 8 | 1.0001 | -0.033069 | 0.031266 | 0.091969
M8 | B | 9 | 0.99985 | -0.033041 | 0.031445 | 0.094052
M8 | B | 10 | 1.0002 | -0.034842 | 0.032791 | 0.093322
M8 | B | 11 | 1.0001 | -0.03432 | 0.032472 | 0.091052
M8 | B | 12 | 0.99978 | -0.033852 | 0.03227 | 0.094652
M8 | B | 13 | 1.0001 | -0.034373 | 0.03251 | 0.09636
M8 | B | 14 | 0.99986 | -0.033725 | 0.032083 | 0.095104
M8 | B | 15 | 0.99987 | -0.034725 | 0.033022 | 0.09892
M8 | B | 16 | 1.0007 | -0.036994 | 0.034433 | 0.098204
M9 | A | 1 | 1.0247 | -0.093045 | 0.066016 | 0.21577
M9 | A | 2 | 1.0254 | -0.095366 | 0.067591 | 0.21592
M9 | A | 3 | 1.0245 | -0.093339 | 0.066401 | 0.22023
M9 | A | 4 | 1.025 | -0.095433 | 0.067977 | 0.22644
M9 | A | 5 | 1.0252 | -0.096239 | 0.068579 | 0.21926
M9 | A | 6 | 1.0263 | -0.099574 | 0.070741 | 0.22504
M9 | A | 7 | 1.0273 | -0.10278 | 0.072872 | 0.22243
M9 | A | 8 | 1.0259 | -0.099563 | 0.071106 | 0.22466
M9 | A | 9 | 1.0266 | -0.10179 | 0.072587 | 0.22564
M9 | A | 10 | 1.0274 | -0.10398 | 0.073936 | 0.23254
M9 | A | 11 | 1.0268 | -0.10379 | 0.074262 | 0.22671
M9 | A | 12 | 1.0272 | -0.10531 | 0.075339 | 0.23159
M9 | A | 13 | 1.0263 | -0.10215 | 0.073139 | 0.22394
M9 | A | 14 | 1.0258 | -0.10139 | 0.072909 | 0.22976
M9 | A | 15 | 1.0261 | -0.10235 | 0.073497 | 0.23373
M9 | A | 16 | 1.0261 | -0.10252 | 0.073718 | 0.23727
M9 | B | 1 | 1.0226 | -0.090435 | 0.065377 | 0.21748
M9 | B | 2 | 1.0221 | -0.090155 | 0.065542 | 0.21928
M9 | B | 3 | 1.024 | -0.096079 | 0.069533 | 0.21625
M9 | B | 4 | 1.0225 | -0.092109 | 0.067077 | 0.21534
M9 | B | 5 | 1.0235 | -0.095061 | 0.068934 | 0.22012
M9 | B | 6 | 1.0244 | -0.098454 | 0.071419 | 0.22097
M9 | B | 7 | 1.024 | -0.096999 | 0.070366 | 0.22851
M9 | B | 8 | 1.0245 | -0.10019 | 0.072936 | 0.22669
M9 | B | 9 | 1.0243 | -0.09915 | 0.072098 | 0.22644
M9 | B | 10 | 1.0252 | -0.10181 | 0.073857 | 0.22593
M9 | B | 11 | 1.0249 | -0.10211 | 0.074388 | 0.22539
M9 | B | 12 | 1.0247 | -0.10228 | 0.074718 | 0.2219
M9 | B | 13 | 1.0246 | -0.10153 | 0.074085 | 0.21729
M9 | B | 14 | 1.0247 | -0.10305 | 0.075422 | 0.22678
M9 | B | 15 | 1.025 | -0.10385 | 0.075949 | 0.22312
M9 | B | 16 | 1.025 | -0.10334 | 0.075443 | 0.22499
M10 | A | 1 | 1.0045 | -0.029878 | 0.024297 | 0.075199
M10 | A | 2 | 1.0043 | -0.029622 | 0.024178 | 0.073833
M10 | A | 3 | 1.0041 | -0.029516 | 0.024317 | 0.077337
M10 | A | 4 | 1.0051 | -0.034014 | 0.027666 | 0.080464
M10 | A | 5 | 1.0052 | -0.032767 | 0.026376 | 0.080019
M10 | A | 6 | 1.0055 | -0.034807 | 0.028006 | 0.078993
M10 | A | 7 | 1.0055 | -0.033707 | 0.027009 | 0.078774
M10 | A | 8 | 1.005 | -0.033662 | 0.027386 | 0.083128
M10 | A | 9 | 1.005 | -0.033032 | 0.026784 | 0.084163
M10 | A | 10 | 1.0049 | -0.033176 | 0.027069 | 0.082655
M10 | A | 11 | 1.0049 | -0.033115 | 0.026948 | 0.085737
M10 | A | 12 | 1.0057 | -0.03605 | 0.029038 | 0.084549
M10 | A | 13 | 1.0054 | -0.034425 | 0.02778 | 0.084392
M10 | A | 14 | 1.0052 | -0.035017 | 0.028472 | 0.085876
M10 | A | 15 | 1.0052 | -0.034298 | 0.027868 | 0.08582
M10 | A | 16 | 1.0061 | -0.038453 | 0.03094 | 0.087445
M10 | B | 1 | 1.0026 | -0.027519 | 0.023712 | 0.079223
M10 | B | 2 | 1.0028 | -0.028631 | 0.024602 | 0.077767
M10 | B | 3 | 1.0029 | -0.029446 | 0.025312 | 0.076737
M10 | B | 4 | 1.0031 | -0.031012 | 0.026586 | 0.078983
M10 | B | 5 | 1.0036 | -0.031159 | 0.026261 | 0.07897
M10 | B | 6 | 1.0033 | -0.031564 | 0.026972 | 0.081046
M10 | B | 7 | 1.0036 | -0.030869 | 0.026024 | 0.083015
M10 | B | 8 | 1.0047 | -0.035962 | 0.029889 | 0.082864
M10 | B | 9 | 1.0034 | -0.031636 | 0.026931 | 0.082417
M10 | B | 10 | 1.0046 | -0.035422 | 0.029437 | 0.082738
M10 | B | 11 | 1.0027 | -0.029599 | 0.025621 | 0.088178
M10 | B | 12 | 1.0035 | -0.033538 | 0.028604 | 0.08566
M10 | B | 13 | 1.003 | -0.031272 | 0.02696 | 0.081764
M10 | B | 14 | 1.0039 | -0.033946 | 0.028646 | 0.085319
M10 | B | 15 | 1.0028 | -0.031058 | 0.026878 | 0.087304
M10 | B | 16 | 1.0047 | -0.037255 | 0.031094 | 0.086647
M11 | A | 1 | 1.0052 | -0.025823 | 0.019805 | 0.090115
M11 | A | 2 | 1.0048 | -0.025158 | 0.019541 | 0.08938
M11 | A | 3 | 1.0049 | -0.026156 | 0.020406 | 0.093167
M11 | A | 4 | 1.0061 | -0.030778 | 0.023717 | 0.090645
M11 | A | 5 | 1.0056 | -0.029824 | 0.023191 | 0.093828
M11 | A | 6 | 1.0077 | -0.036241 | 0.02743 | 0.094218
M11 | A | 7 | 1.0065 | -0.032939 | 0.025417 | 0.095186
M11 | A | 8 | 1.0058 | -0.03179 | 0.024949 | 0.098078
M11 | A | 9 | 1.0067 | -0.033723 | 0.025895 | 0.094839
M11 | A | 10 | 1.0077 | -0.037563 | 0.028678 | 0.095441
M11 | A | 11 | 1.0075 | -0.037529 | 0.028825 | 0.10192
M11 | A | 12 | 1.0079 | -0.038299 | 0.0292 | 0.10659
M11 | A | 13 | 1.0071 | -0.035652 | 0.027437 | 0.096396
M11 | A | 14 | 1.0081 | -0.04024 | 0.030817 | 0.094223
M11 | A | 15 | 1.0073 | -0.036367 | 0.027896 | 0.096573
M11 | A | 16 | 1.0085 | -0.041086 | 0.031326 | 0.10189
M11 | B | 1 | 1.0047 | -0.025119 | 0.019598 | 0.096023
M11 | B | 2 | 1.0045 | -0.02589 | 0.020461 | 0.086692
M11 | B | 3 | 1.005 | -0.027945 | 0.02201 | 0.094493
M11 | B | 4 | 1.0046 | -0.027677 | 0.02209 | 0.091152
M11 | B | 5 | 1.0045 | -0.027515 | 0.022046 | 0.089238
M11 | B | 6 | 1.0057 | -0.031775 | 0.024948 | 0.093735
M11 | B | 7 | 1.0054 | -0.031316 | 0.024794 | 0.093057
M11 | B | 8 | 1.0052 | -0.032364 | 0.026003 | 0.095877
M11 | B | 9 | 1.0056 | -0.032324 | 0.025575 | 0.093715
M11 | B | 10 | 1.006 | -0.033995 | 0.02684 | 0.09486
M11 | B | 11 | 1.0061 | -0.034424 | 0.027117 | 0.094334
M11 | B | 12 | 1.0065 | -0.036364 | 0.02863 | 0.096805
M11 | B | 13 | 1.0072 | -0.037445 | 0.029023 | 0.092648
M11 | B | 14 | 1.0065 | -0.036139 | 0.028441 | 0.099499
M11 | B | 15 | 1.0072 | -0.037706 | 0.029269 | 0.10225
M11 | B | 16 | 1.0075 | -0.039547 | 0.030763 | 0.09805

## DNB RVS
Band | HAM | Band Median |  |  |  | Band Average
|  | a0 | a1 | a2 | Uncertainty | a0 | a1 | a2 | Uncertainty
DNB | A | 0.99859 | 0.00016523 | -2.433e-06 | 0.036959 | 0.99858 | 0.00017735 | -2.5435e-06 | 0.037009
DNB | B | 0.99941 | 0.00019516 | -3.1578e-06 | 0.037207 | 0.99953 | 0.00019882 | -3.1591e-06 | 0.037458
Band | HAM | Detector | Detector Dependent |  |  |  |  | Notes:
|  |  | a0 | a1 | a2 | Uncertainty |  | 1. This table provides VIIRS F2 DNB RVS.
2. The results were derived from FP-10 using the SIS-100 source.
DNB | A | 1 | 0.99732 | 0.00030246 | -4.2685e-06 | 0.037375 |  | 3. Quadratic fitting was performed on RVS vs HAM AOI.
DNB | A | 2 | 1.0002 | 0.00012719 | -2.1453e-06 | 0.036591 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
DNB | A | 3 | 0.99932 | 0.00015743 | -2.4174e-06 | 0.036409 |  | 5. The test was set to operational day mode.
DNB | A | 4 | 0.99906 | 0.00016421 | -2.4584e-06 | 0.038447
DNB | A | 5 | 0.99881 | 0.00016626 | -2.4225e-06 | 0.035141
DNB | A | 6 | 0.99684 | 0.00025184 | -3.2995e-06 | 0.037136
DNB | A | 7 | 0.99796 | 0.00021461 | -2.9911e-06 | 0.038394
DNB | A | 8 | 0.99928 | 0.00015591 | -2.3814e-06 | 0.034559
DNB | A | 9 | 0.9985 | 0.00017263 | -2.4435e-06 | 0.036781
DNB | A | 10 | 0.99813 | 0.00020518 | -2.8807e-06 | 0.039612
DNB | A | 11 | 0.99799 | 0.00020048 | -2.7672e-06 | 0.039199
DNB | A | 12 | 0.99868 | 0.00016297 | -2.335e-06 | 0.03678
DNB | A | 13 | 0.99799 | 0.00019901 | -2.7422e-06 | 0.03722
DNB | A | 14 | 0.99936 | 0.00013075 | -1.9864e-06 | 0.035146
DNB | A | 15 | 0.99951 | 0.00012373 | -1.9125e-06 | 0.03512
DNB | A | 16 | 0.99833 | 0.00010291 | -1.245e-06 | 0.038232
DNB | B | 1 | 0.9982 | 0.00031996 | -4.7999e-06 | 0.038599
DNB | B | 2 | 1.001 | 0.00015084 | -2.7768e-06 | 0.035389
DNB | B | 3 | 1.0009 | 0.0001492 | -2.7042e-06 | 0.035167
DNB | B | 4 | 0.99992 | 0.00018904 | -3.105e-06 | 0.037191
DNB | B | 5 | 0.99976 | 0.00018508 | -2.9959e-06 | 0.035685
DNB | B | 6 | 0.99868 | 0.00023363 | -3.5033e-06 | 0.041023
DNB | B | 7 | 0.99884 | 0.00023955 | -3.6432e-06 | 0.038733
DNB | B | 8 | 0.99957 | 0.00020127 | -3.2105e-06 | 0.037325
DNB | B | 9 | 0.99913 | 0.00021045 | -3.2414e-06 | 0.037047
DNB | B | 10 | 0.99924 | 0.00020895 | -3.2483e-06 | 0.03983
DNB | B | 11 | 0.99907 | 0.00021696 | -3.3328e-06 | 0.037222
DNB | B | 12 | 0.99903 | 0.00022024 | -3.3769e-06 | 0.037935
DNB | B | 13 | 0.99963 | 0.00018563 | -2.9695e-06 | 0.036636
DNB | B | 14 | 1.0002 | 0.00016274 | -2.7352e-06 | 0.035209
DNB | B | 15 | 1.0006 | 0.00013995 | -2.4899e-06 | 0.036429
DNB | B | 16 | 0.99868 | 0.00016767 | -2.4119e-06 | 0.03991

## TEB NFR
Band | HAM | Baffle Locations |  | Mean | [7 empty] | Median | [7 empty] | Sensor Performance Compared with Requirement
|  |  |  | Harvy Shack Fit Parameters | [7 empty] | Harvy Shack Fit Parameters | [7 empty] | Mean |  | Median |  | Max Allowed Lscat | Angle (mrad) from Target | T Bright Target (K)
|  | Left | Right | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Scattered Radiance
|  |  |  | A | M | A | M | A | M | A | M | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright | Left of Lbright | Right of Lbright
I4 | A | 33 | 120 | 0.99891 | -0.48556 | 10.1951 | -2.0832 | 13.3091 | -1.7455 | 20.708 | -3.0049 | 0.9993 | -0.4419 | 9.4237 | -2.0894 | 12.9558 | -1.763 | 20.0917 | -3.0256 | ~ | ~ | ~ | ` | ~ | ~ | ~
I4 | B | 33 | 120 | 0.93598 | -0.41257 | 11.8851 | -2.0084 | 13.3236 | -1.7459 | 17.4793 | -2.832 | 0.92671 | -0.41051 | 11.3698 | -2.0293 | 12.9482 | -1.7234 | 17.1934 | -2.7777 | ~ | ~ | ~ | ~ | ~ | ~ | ~
I5 | A | -36 | 57 | 5.0277 | 0.077724 | 1.4532 | -3.005 | 41.5618 | -1.1696 | 21.3386 | -1.1113 | 4.7268 | 0.072543 | 1.1526 | -2.9801 | 40.4595 | -1.1713 | 21.3368 | -1.0773 | ~ | ~ | ~ | ~ | ~ | ~ | ~
I5 | B | -36 | 57 | 5.2129 | 0.13797 | 2.4331 | -2.8243 | 46.2017 | -1.1259 | 21.4804 | -1.0552 | 5.0663 | 0.21339 | 1.9585 | -2.8306 | 45.2517 | -1.1213 | 20.7475 | -1.0731 | ~ | ~ | ~ | ~ | ~ | ~ | ~
M12 | A | -12 | 77 | 0.25738 | -0.81389 | 0.39264 | -2.8953 | 3.2175 | -2.0884 | 1.5971 | -2.0274 | 0.26821 | -0.80907 | 0.4242 | -2.8376 | 3.2927 | -2.0834 | 1.4321 | -1.9495 | 1.4314e-05 | 0.00015692 | 1.4684e-05 | 0.00015786 | 0.00047761 | 3 | 295
M12 | B | -12 | 77 | 0.24584 | -0.81646 | 0.31224 | -3.0257 | 3.2477 | -2.057 | 1.587 | -1.9039 | 0.26205 | -0.82833 | 0.31735 | -2.9875 | 3.3257 | -2.0732 | 1.252 | -1.9952 | 1.3391e-05 | 0.00015543 | 1.3724e-05 | 0.00015585 | 0.00047761 | 3 | 295
M13 | A | -8 | 72 | 0.23665 | -0.86867 | 0.00079868 | -5.4595 | 2.8728 | -2.0743 | 6.9508 | -2.3616 | 0.22839 | -0.91883 | 0.00054914 | -5.3881 | 2.7748 | -2.0851 | 0.85316 | -2.5508 | 3.3749e-05 | 0.00073457 | 3.3072e-05 | 0.00072295 | 0.0031468 | 3 | 321
M13 | B | -8 | 72 | 0.31305 | -0.74098 | 0.0012287 | -5.3832 | 3.007 | -2.0478 | 2.6674 | -2.3478 | 0.30746 | -0.77752 | 0.00066691 | -5.3076 | 3.0109 | -2.0617 | 0.80495 | -1.8611 | 4.2585e-05 | 0.00075622 | 4.2309e-05 | 0.00073388 | 0.0031468 | 3 | 321
M14 | A | -39 | 42 | 0.82259 | -0.50566 | 4.1325 | -2.2831 | 3.7391 | -2.0817 | 0.51987 | 0.04242 | 0.74806 | -0.61919 | 4.1148 | -2.2855 | 4.265 | -1.9479 | 0.44768 | 0.129 | ~ | ~ | ~ | ~ | ~ | ~ | ~
M14 | B | -39 | 42 | 1.0594 | -0.5182 | 4.4853 | -2.2389 | 4.8016 | -1.9377 | 0.584 | -0.0053792 | 0.64486 | -0.36001 | 4.3928 | -2.2345 | 4.7384 | -1.9367 | 0.54165 | -0.089347 | ~ | ~ | ~ | ~ | ~ | ~ | ~
M15 | A | -33 | 45 | 1.0061 | -0.60296 | 3.6127 | -2.224 | 3.7963 | -2.2319 | 0.70891 | 0.0096326 | 0.87666 | -0.61651 | 3.5353 | -2.2161 | 3.9657 | -2.1829 | 0.69475 | -0.099283 | 0.00010704 | 0.00011803 | 0.00010743 | 0.00011932 | 0.0096859 | 3 | 160
M15 | B | -33 | 45 | 0.62456 | 0.11801 | 3.5778 | -2.236 | 4.0659 | -2.1001 | 0.7072 | -0.11235 | 0.5411 | 0.06749 | 3.7715 | -2.2266 | 4.2576 | -2.0855 | 0.7278 | -0.23691 | 0.00010585 | 0.00011791 | 0.00010787 | 0.00012032 | 0.0096859 | 3 | 160
M16A | A | -25 | 61 | 1.0354 | 0.032027 | 1.0007 | -2.8885 | 3.8414 | -2.0865 | 1.1627 | 0.19265 | 1.0865 | -0.10405 | 0.63462 | -2.9328 | 3.6456 | -2.1014 | 1.1629 | -0.068591 | 0.0024972 | 0.0065163 | 0.0022194 | 0.0064303 | 0.0089523 | 3 | 318
M16A | B | -25 | 61 | 1.1118 | -0.020227 | 1.1858 | -2.7717 | 3.676 | -2.1288 | 1.2514 | -0.019089 | 1.0851 | -0.099407 | 1.1783 | -2.6291 | 3.7294 | -2.1424 | 0.92458 | 0.12851 | 0.0026697 | 0.0063843 | 0.0028573 | 0.0064969 | 0.0089523 | 3 | 318
M16B | A | -25 | 49 | 1.44 | -0.36167 | 1.7593 | -2.3965 | 3.9223 | -2.1239 | 1.468 | -0.18648 | 1.4504 | -0.38475 | 1.5768 | -2.3416 | 4.0012 | -2.0888 | 1.3165 | -0.21075 | 0.0032435 | 0.0067668 | 0.0033061 | 0.0067662 | 0.0089523 | 3 | 318
M16B | B | -25 | 49 | 1.1865 | -0.27175 | 2.1675 | -2.2459 | 3.7406 | -2.1687 | 1.5868 | -0.38037 | 1.2665 | -0.14799 | 2.1954 | -2.182 | 3.8578 | -2.146 | 1.2839 | -0.32586 | 0.0034515 | 0.0066426 | 0.0034797 | 0.0066612 | 0.0089523 | 3 | 318
Band | HAM | Detector | Harvy Shack Fit Parameters | [7 empty] | Sensor Performance Compared to Requirement | [5 empty] | Notes:
|  |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Max Allowed Lscat | Angle (mrad) from Target | T Bright Target (K) |  | 1. This table provides VIIRS F2 TEB NFR.
|  |  | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright |  |  |  |  | 2. The results were derived from FP-14 using the ScMA.
I4 | A | 1 | 0.82681 | -0.22656 | 3.8672 | -2.4198 | 10.0862 | -1.7668 | 11.9313 | -2.7608 | ~ | ~ | ~ | ~ | ~ |  | 3. Linear fits were peformed of log(dn/(s*t))= log(A) + M*Log(sin(theta)/0.01))
I4 | A | 2 | 0.99657 | -0.27883 | 5.3752 | -2.2766 | 11.5955 | -1.6591 | 9.5441 | -2.4 | ~ | ~ | ~ | ~ | ~ |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I4 | A | 3 | 1.0135 | -0.11547 | 6.0196 | -2.2876 | 12.3021 | -1.7509 | 13.9614 | -2.6217 | ~ | ~ | ~ | ~ | ~ |  | 5. The test was conducted in diagnostic mode.
I4 | A | 4 | 0.99789 | -0.36776 | 6.5969 | -2.2402 | 12.1675 | -1.8928 | 10.4861 | -2.4532 | ~ | ~ | ~ | ~ | ~ |  | 6. Radiance units are W m-2 sr-1 micron-1
I4 | A | 5 | 1.0405 | -0.36464 | 8.0555 | -2.117 | 12.3966 | -1.876 | 16.4131 | -2.9262 | ~ | ~ | ~ | ~ | ~ |  | 7. There is no NFR requirement for bands M14, I4, or I5
I4 | A | 6 | 1.0646 | -0.49217 | 8.6529 | -2.0806 | 12.7624 | -1.8003 | 13.6126 | -2.6446 | ~ | ~ | ~ | ~ | ~
I4 | A | 7 | 0.91091 | -0.3815 | 7.842 | -2.2124 | 13.0561 | -1.6931 | 21.1414 | -3.1681 | ~ | ~ | ~ | ~ | ~
I4 | A | 8 | 1.119 | -0.69333 | 7.8655 | -2.1936 | 13.2862 | -1.7728 | 17.5743 | -2.8767 | ~ | ~ | ~ | ~ | ~
I4 | A | 9 | 1.0099 | -0.63559 | 7.338 | -2.2163 | 13.5697 | -1.7985 | 32.423 | -3.6435 | ~ | ~ | ~ | ~ | ~
I4 | A | 10 | 1.0206 | -0.58894 | 20.2267 | -1.6575 | 13.9757 | -1.6985 | 17.6633 | -2.954 | ~ | ~ | ~ | ~ | ~
I4 | A | 11 | 0.96407 | -0.4797 | 7.4921 | -2.2371 | 13.7621 | -1.7932 | 40.8113 | -3.8689 | ~ | ~ | ~ | ~ | ~
I4 | A | 12 | 0.82919 | -0.3761 | 11.216 | -1.9983 | 12.8555 | -1.8747 | 21.6911 | -3.1168 | ~ | ~ | ~ | ~ | ~
I4 | A | 13 | 0.85179 | -0.26858 | 9.7667 | -2.0476 | 11.5393 | -1.8069 | 17.9927 | -3.1433 | ~ | ~ | ~ | ~ | ~
I4 | A | 14 | 0.79268 | -0.9261 | 9.0064 | -2.0982 | 11.1735 | -1.7593 | 19.155 | -3.1277 | ~ | ~ | ~ | ~ | ~
I4 | A | 15 | 0.73203 | -0.50884 | 7.8015 | -2.1787 | 10.9353 | -1.836 | 14.9804 | -3.0539 | ~ | ~ | ~ | ~ | ~
I4 | A | 16 | 0.76051 | -0.71326 | 7.6038 | -2.1906 | 10.1772 | -1.8517 | 15.839 | -3.0225 | ~ | ~ | ~ | ~ | ~
I4 | A | 17 | 0.69532 | -0.85996 | 10.4941 | -1.9779 | 10.0998 | -1.8482 | 20.3042 | -3.3858 | ~ | ~ | ~ | ~ | ~
I4 | A | 18 | 0.71747 | -0.82871 | 10.1218 | -2.0093 | 10.157 | -1.8159 | 14.9551 | -2.9749 | ~ | ~ | ~ | ~ | ~
I4 | A | 19 | 0.65967 | -0.39716 | 12.415 | -1.9121 | 10.5865 | -1.8499 | 20.9041 | -3.284 | ~ | ~ | ~ | ~ | ~
I4 | A | 20 | 0.7826 | -0.43936 | 9.0807 | -2.1162 | 11.5325 | -1.8243 | 23.118 | -3.3017 | ~ | ~ | ~ | ~ | ~
I4 | A | 21 | 0.9025 | -0.60882 | 9.0464 | -2.1478 | 12.1819 | -1.8108 | 19.8791 | -3.0288 | ~ | ~ | ~ | ~ | ~
I4 | A | 22 | 1.0007 | -0.61171 | 19.0019 | -1.7241 | 13.4316 | -1.6412 | 20.99 | -3.124 | ~ | ~ | ~ | ~ | ~
I4 | A | 23 | 0.96588 | -0.44443 | 15.4674 | -1.8882 | 15.2026 | -1.4864 | 17.466 | -2.7596 | ~ | ~ | ~ | ~ | ~
I4 | A | 24 | 1.3079 | -0.3427 | 13.5062 | -2.0007 | 16.1577 | -1.7268 | 28.239 | -3.2232 | ~ | ~ | ~ | ~ | ~
I4 | A | 25 | 1.1386 | -0.34254 | 12.6858 | -2.0369 | 17.0259 | -1.6309 | 28.1471 | -3.0571 | ~ | ~ | ~ | ~ | ~
I4 | A | 26 | 1.2829 | -0.79195 | 12.1086 | -2.0646 | 17.1289 | -1.6505 | 30.9715 | -3.2083 | ~ | ~ | ~ | ~ | ~
I4 | A | 27 | 1.287 | -0.36844 | 12.7768 | -2.0278 | 16.5577 | -1.7141 | 24.4888 | -2.8389 | ~ | ~ | ~ | ~ | ~
I4 | A | 28 | 1.4187 | -0.27182 | 10.8941 | -2.1312 | 16.7121 | -1.7068 | 27.3394 | -3.0132 | ~ | ~ | ~ | ~ | ~
I4 | A | 29 | 1.3712 | -0.52019 | 13.0299 | -1.9753 | 16.7689 | -1.6963 | 21.4374 | -2.7072 | ~ | ~ | ~ | ~ | ~
I4 | A | 30 | 1.2065 | -0.31073 | 13.3412 | -1.9565 | 16.547 | -1.5893 | 26.9548 | -3.0312 | ~ | ~ | ~ | ~ | ~
I4 | A | 31 | 1.295 | -0.69019 | 10.0315 | -2.0583 | 15.5883 | -1.6179 | 25.7346 | -2.8672 | ~ | ~ | ~ | ~ | ~
I4 | A | 32 | 1.0025 | -0.2919 | 7.5144 | -2.1822 | 14.5719 | -1.6157 | 16.5066 | -2.5693 | ~ | ~ | ~ | ~ | ~
I4 | B | 1 | 0.75609 | -0.32156 | 4.4478 | -2.3575 | 10.2814 | -1.782 | 7.9493 | -2.3733 | ~ | ~ | ~ | ~ | ~
I4 | B | 2 | 0.85737 | -0.18686 | 6.2844 | -2.2102 | 11.682 | -1.4645 | 7.7668 | -2.1798 | ~ | ~ | ~ | ~ | ~
I4 | B | 3 | 0.939 | -0.20516 | 6.7354 | -2.2452 | 12.5811 | -1.7333 | 13.4242 | -2.707 | ~ | ~ | ~ | ~ | ~
I4 | B | 4 | 0.90457 | -0.17137 | 7.7646 | -2.1657 | 12.677 | -1.8504 | 9.6481 | -2.3577 | ~ | ~ | ~ | ~ | ~
I4 | B | 5 | 1.0314 | -0.45612 | 9.3047 | -2.0526 | 12.9646 | -1.8416 | 18.1059 | -2.9873 | ~ | ~ | ~ | ~ | ~
I4 | B | 6 | 0.91443 | -0.37617 | 10.5562 | -1.982 | 12.9318 | -1.7216 | 13.5646 | -2.6598 | ~ | ~ | ~ | ~ | ~
I4 | B | 7 | 0.95186 | -0.19344 | 8.3345 | -2.1935 | 12.8914 | -1.6927 | 15.4866 | -2.7245 | ~ | ~ | ~ | ~ | ~
I4 | B | 8 | 1.079 | -0.54584 | 12.9734 | -1.9107 | 13.2209 | -1.7083 | 24.1738 | -3.2686 | ~ | ~ | ~ | ~ | ~
I4 | B | 9 | 0.9564 | -0.32812 | 12.2935 | -1.9422 | 13.8506 | -1.7976 | 17.8945 | -2.8309 | ~ | ~ | ~ | ~ | ~
I4 | B | 10 | 0.99934 | -0.65695 | 20.9742 | -1.6458 | 14.0473 | -1.7009 | 19.7985 | -3.0344 | ~ | ~ | ~ | ~ | ~
I4 | B | 11 | 0.8806 | -0.45452 | 14.3282 | -1.8732 | 13.5806 | -1.8174 | 25.1751 | -3.1791 | ~ | ~ | ~ | ~ | ~
I4 | B | 12 | 0.82186 | -0.38334 | 11.7505 | -1.9865 | 12.7204 | -1.9057 | 25.0544 | -3.3542 | ~ | ~ | ~ | ~ | ~
I4 | B | 13 | 0.86204 | -0.46624 | 9.8901 | -2.0523 | 11.8141 | -1.7937 | 17.4179 | -2.9878 | ~ | ~ | ~ | ~ | ~
I4 | B | 14 | 0.69082 | -0.36904 | 9.8412 | -2.0546 | 11.3107 | -1.7178 | 16.6306 | -3.0545 | ~ | ~ | ~ | ~ | ~
I4 | B | 15 | 0.74292 | -0.39717 | 8.4176 | -2.1465 | 10.7916 | -1.85 | 14.7001 | -2.9742 | ~ | ~ | ~ | ~ | ~
I4 | B | 16 | 0.63731 | -0.30365 | 7.9794 | -2.1698 | 10.1911 | -1.8666 | 10.1441 | -2.7083 | ~ | ~ | ~ | ~ | ~
I4 | B | 17 | 0.60659 | -0.77699 | 12.838 | -1.8655 | 10.0878 | -1.8489 | 17.7241 | -3.217 | ~ | ~ | ~ | ~ | ~
I4 | B | 18 | 0.63981 | -0.53609 | 11.2164 | -1.9567 | 10.0293 | -1.8505 | 25.5462 | -3.6963 | ~ | ~ | ~ | ~ | ~
I4 | B | 19 | 0.71437 | -0.51716 | 17.4853 | -1.7088 | 10.6573 | -1.8549 | 18.4738 | -3.1966 | ~ | ~ | ~ | ~ | ~
I4 | B | 20 | 0.68164 | -0.36252 | 9.1696 | -2.115 | 11.1061 | -1.8528 | 12.0724 | -2.6854 | ~ | ~ | ~ | ~ | ~
I4 | B | 21 | 0.68905 | -0.48169 | 8.5328 | -2.1877 | 12.1895 | -1.8231 | 24.3093 | -3.3161 | ~ | ~ | ~ | ~ | ~
I4 | B | 22 | 0.907 | -0.25771 | 23.6301 | -1.5916 | 13.5961 | -1.723 | 17.9446 | -2.8999 | ~ | ~ | ~ | ~ | ~
I4 | B | 23 | 0.97069 | -0.76317 | 25.4661 | -1.5972 | 14.896 | -1.6708 | 16.969 | -2.6828 | ~ | ~ | ~ | ~ | ~
I4 | B | 24 | 1.0012 | -0.54136 | 14.1119 | -1.9778 | 16.2549 | -1.7237 | 14.1625 | -2.4461 | ~ | ~ | ~ | ~ | ~
I4 | B | 25 | 1.3178 | -0.44885 | 13.1788 | -2.0201 | 16.8073 | -1.6628 | 24.3708 | -2.8639 | ~ | ~ | ~ | ~ | ~
I4 | B | 26 | 1.1423 | -0.49606 | 12.767 | -2.0386 | 16.9904 | -1.6683 | 16.3054 | -2.559 | ~ | ~ | ~ | ~ | ~
I4 | B | 27 | 1.2831 | -0.2208 | 14.288 | -1.9676 | 16.9268 | -1.6964 | 30.3807 | -3.1765 | ~ | ~ | ~ | ~ | ~
I4 | B | 28 | 1.2332 | -0.42385 | 11.9536 | -2.0787 | 16.9057 | -1.6998 | 17.4882 | -2.5304 | ~ | ~ | ~ | ~ | ~
I4 | B | 29 | 1.2196 | -0.26741 | 10.9135 | -2.084 | 16.8381 | -1.6879 | 19.1258 | -2.5619 | ~ | ~ | ~ | ~ | ~
I4 | B | 30 | 1.2047 | -0.58958 | 13.4972 | -1.9546 | 16.2927 | -1.6129 | 15.8215 | -2.4319 | ~ | ~ | ~ | ~ | ~
I4 | B | 31 | 1.0814 | -0.27848 | 11.5231 | -1.9789 | 15.4471 | -1.6161 | 15.6445 | -2.4181 | ~ | ~ | ~ | ~ | ~
I4 | B | 32 | 1.2338 | -0.42495 | 7.876 | -2.1583 | 13.792 | -1.6325 | 16.0655 | -2.5595 | ~ | ~ | ~ | ~ | ~
I5 | A | 1 | 10.2292 | 0.049001 | 0.51797 | -3.5836 | 69.523 | -0.72349 | 27.972 | -1.0961 | ~ | ~ | ~ | ~ | ~
I5 | A | 2 | 3.7041 | 0.4533 | 3.159 | -2.7742 | 48.9591 | -0.97055 | 21.7549 | -0.90841 | ~ | ~ | ~ | ~ | ~
I5 | A | 3 | 9.3653 | -0.27521 | 0.2452 | -3.4583 | 41.2218 | -1.1585 | 22.3967 | -0.95092 | ~ | ~ | ~ | ~ | ~
I5 | A | 4 | 4.8185 | 0.37883 | 2.6893 | -2.6709 | 29.9915 | -1.306 | 21.6882 | -1.1234 | ~ | ~ | ~ | ~ | ~
I5 | A | 5 | 5.5045 | -0.29966 | 1.0369 | -3.0634 | 40.2902 | -1.1485 | 14.6017 | -0.959 | ~ | ~ | ~ | ~ | ~
I5 | A | 6 | 4.635 | -0.073635 | 3.0852 | -2.6073 | 43.9758 | -1.2415 | 25.3417 | -1.3738 | ~ | ~ | ~ | ~ | ~
I5 | A | 7 | 4.138 | 0.063095 | 0.92279 | -3.0381 | 35.9064 | -1.1739 | 29.3361 | -1.4992 | ~ | ~ | ~ | ~ | ~
I5 | A | 8 | 3.8346 | 0.12568 | 3.1828 | -2.6064 | 50.5579 | -1.1418 | 22.3 | -1.1727 | ~ | ~ | ~ | ~ | ~
I5 | A | 9 | 6.9737 | 0.43241 | 0.66539 | -3.1431 | 40.9198 | -1.0484 | 20.8985 | -0.7049 | ~ | ~ | ~ | ~ | ~
I5 | A | 10 | 4.491 | 0.11493 | 2.1292 | -2.7849 | 33.7161 | -1.3115 | 24.4564 | -1.2687 | ~ | ~ | ~ | ~ | ~
I5 | A | 11 | 4.5146 | 0.36818 | 0.75563 | -3.1127 | 34.5845 | -1.3 | 19.7976 | -1.0401 | ~ | ~ | ~ | ~ | ~
I5 | A | 12 | 3.13 | 0.86888 | 1.3507 | -2.9592 | 53.5909 | -1.0994 | 20.3268 | -1.3629 | ~ | ~ | ~ | ~ | ~
I5 | A | 13 | 3.119 | 0.55218 | 0.91662 | -3.0443 | 32.8871 | -1.3231 | 14.1493 | -0.9233 | ~ | ~ | ~ | ~ | ~
I5 | A | 14 | 7.7818 | 0.18053 | 3.8922 | -2.49 | 26.4957 | -1.3603 | 28.8536 | -0.91161 | ~ | ~ | ~ | ~ | ~
I5 | A | 15 | 3.6368 | -0.34444 | 0.21102 | -3.5419 | 40.6288 | -1.2615 | 14.9372 | -0.95348 | ~ | ~ | ~ | ~ | ~
I5 | A | 16 | 5.6305 | -0.16927 | 1.0667 | -3.0711 | 42.6428 | -1.1771 | 26.5927 | -1.6685 | ~ | ~ | ~ | ~ | ~
I5 | A | 17 | 4.9491 | 0.14563 | 1.2218 | -2.959 | 40.9124 | -1.2343 | 21.7564 | -1.5447 | ~ | ~ | ~ | ~ | ~
I5 | A | 18 | 3.2926 | 0.081991 | 2.1288 | -2.7648 | 27.4898 | -1.3211 | 17.7388 | -1.0179 | ~ | ~ | ~ | ~ | ~
I5 | A | 19 | 3.049 | 0.30092 | 0.45773 | -3.2965 | 40.8374 | -1.1682 | 20.9576 | -1.0584 | ~ | ~ | ~ | ~ | ~
I5 | A | 20 | 1.9988 | 0.46369 | 0.62635 | -3.3031 | 37.5423 | -1.1673 | 16.9803 | -1.2923 | ~ | ~ | ~ | ~ | ~
I5 | A | 21 | 5.1877 | -0.088004 | 0.76376 | -3.1934 | 42.7376 | -1.217 | 13.9421 | -1.0035 | ~ | ~ | ~ | ~ | ~
I5 | A | 22 | 3.1697 | 0.43859 | 2.3903 | -2.7016 | 33.8528 | -1.2151 | 28.0782 | -1.5158 | ~ | ~ | ~ | ~ | ~
I5 | A | 23 | 5.6726 | -0.34259 | 0.20175 | -3.5839 | 39.0751 | -1.0882 | 15.0374 | -0.92512 | ~ | ~ | ~ | ~ | ~
I5 | A | 24 | 5.7388 | -0.73912 | 0.72406 | -3.2122 | 47.694 | -1.0812 | 25.7094 | -1.1166 | ~ | ~ | ~ | ~ | ~
I5 | A | 25 | 6.6641 | -0.15427 | 1.9414 | -2.7704 | 45.0378 | -1.1534 | 28.6114 | -1.1063 | ~ | ~ | ~ | ~ | ~
I5 | A | 26 | 4.2555 | 0.034777 | 1.7746 | -2.8363 | 46.2569 | -1.1395 | 20.9855 | -1.1397 | ~ | ~ | ~ | ~ | ~
I5 | A | 27 | 5.0901 | -0.39587 | 1.8448 | -2.8029 | 35.7727 | -1.1687 | 19.1602 | -0.86627 | ~ | ~ | ~ | ~ | ~
I5 | A | 28 | 7.9704 | 0.12281 | 1.3947 | -2.9324 | 36.2573 | -1.1589 | 24.6386 | -1.1279 | ~ | ~ | ~ | ~ | ~
I5 | A | 29 | 3.6014 | -0.11833 | 1.7103 | -2.8061 | 36.7714 | -1.1977 | 14.504 | -0.86558 | ~ | ~ | ~ | ~ | ~
I5 | A | 30 | 4.0289 | 0.64434 | 0.6795 | -3.251 | 38.758 | -1.1964 | 24.765 | -1.3572 | ~ | ~ | ~ | ~ | ~
I5 | A | 31 | 5.7963 | 0.046373 | 1.0834 | -3.001 | 35.6914 | -1.189 | 15.1035 | -0.80896 | ~ | ~ | ~ | ~ | ~
I5 | A | 32 | 4.9138 | -0.37855 | 1.7325 | -2.7968 | 79.3991 | -0.98514 | 19.4624 | -0.89845 | ~ | ~ | ~ | ~ | ~
I5 | B | 1 | 4.2726 | 0.36504 | 2.4215 | -2.709 | 67.8595 | -0.82947 | 23.4808 | -1.2666 | ~ | ~ | ~ | ~ | ~
I5 | B | 2 | 5.3907 | 0.29316 | 3.9387 | -2.5406 | 47.0396 | -1.1979 | 37.0387 | -1.5053 | ~ | ~ | ~ | ~ | ~
I5 | B | 3 | 6.4034 | 0.24825 | 2.3344 | -2.733 | 31.0528 | -1.2802 | 21.9802 | -1.3186 | ~ | ~ | ~ | ~ | ~
I5 | B | 4 | 5.1797 | 0.02765 | 2.7685 | -2.6758 | 43.6687 | -1.1754 | 17.3695 | -0.95204 | ~ | ~ | ~ | ~ | ~
I5 | B | 5 | 8.6265 | 0.27923 | 1.2849 | -2.9611 | 55.9498 | -1.0448 | 33.2969 | -1.3222 | ~ | ~ | ~ | ~ | ~
I5 | B | 6 | 4.5165 | -0.074293 | 6.0949 | -2.3082 | 29.8278 | -1.3355 | 32.5863 | -1.5112 | ~ | ~ | ~ | ~ | ~
I5 | B | 7 | 4.357 | 0.36206 | 2.6406 | -2.7024 | 80.4365 | -0.87821 | 15.0641 | -0.81024 | ~ | ~ | ~ | ~ | ~
I5 | B | 8 | 7.407 | -0.51003 | 4.1664 | -2.5207 | 48.5968 | -1.1385 | 25.5341 | -1.4413 | ~ | ~ | ~ | ~ | ~
I5 | B | 9 | 3.6068 | -0.24345 | 1.7723 | -2.8445 | 40.4388 | -1.1314 | 17.5406 | -0.71336 | ~ | ~ | ~ | ~ | ~
I5 | B | 10 | 5.7974 | 0.26706 | 3.1166 | -2.6449 | 55.6208 | -1.1094 | 30.2195 | -0.90701 | ~ | ~ | ~ | ~ | ~
I5 | B | 11 | 3.5267 | 0.14711 | 1.3265 | -2.9773 | 43.1843 | -1.1845 | 17.9575 | -0.97858 | ~ | ~ | ~ | ~ | ~
I5 | B | 12 | 2.1118 | 1.2845 | 1.9162 | -2.8189 | 42.0649 | -1.1719 | 12.63 | -1.0065 | ~ | ~ | ~ | ~ | ~
I5 | B | 13 | 4.8966 | 0.38037 | 1.0539 | -3.0378 | 46.3294 | -1.212 | 23.6262 | -1.0702 | ~ | ~ | ~ | ~ | ~
I5 | B | 14 | 10.1804 | 0.0072436 | 6.0674 | -2.3088 | 35.9367 | -1.2303 | 30.0992 | -1.076 | ~ | ~ | ~ | ~ | ~
I5 | B | 15 | 3.8785 | -0.063148 | 0.92588 | -3.0976 | 20.6062 | -1.5079 | 19.3349 | -1.0767 | ~ | ~ | ~ | ~ | ~
I5 | B | 16 | 2.6141 | 0.58962 | 0.41774 | -3.395 | 49.2911 | -1.1053 | 17.4543 | -1.4074 | ~ | ~ | ~ | ~ | ~
I5 | B | 17 | 2.7506 | 1.0118 | 1.7498 | -2.8651 | 44.0048 | -1.2042 | 13.7264 | -1.1171 | ~ | ~ | ~ | ~ | ~
I5 | B | 18 | 7.1804 | -0.44812 | 4.4213 | -2.4654 | 48.1282 | -1.1072 | 25.6519 | -1.0455 | ~ | ~ | ~ | ~ | ~
I5 | B | 19 | 5.0668 | -0.40426 | 0.39546 | -3.3954 | 42.3511 | -1.1145 | 10.9634 | -0.44958 | ~ | ~ | ~ | ~ | ~
I5 | B | 20 | 3.392 | -0.1228 | 4.6431 | -2.4313 | 29.6035 | -1.1813 | 22.1142 | -1.1821 | ~ | ~ | ~ | ~ | ~
I5 | B | 21 | 3.0817 | 0.5765 | 0.79159 | -3.1554 | 24.8945 | -1.3947 | 16.6294 | -1.1531 | ~ | ~ | ~ | ~ | ~
I5 | B | 22 | 7.3845 | 0.19593 | 5.7198 | -2.333 | 51.029 | -1.0792 | 30.7129 | -0.89687 | ~ | ~ | ~ | ~ | ~
I5 | B | 23 | 5.0659 | -0.12975 | 1.2448 | -2.9906 | 52.7582 | -1.0271 | 19.5311 | -1.1339 | ~ | ~ | ~ | ~ | ~
I5 | B | 24 | 6.2158 | -0.5661 | 3.6108 | -2.5295 | 43.2179 | -1.0606 | 19.801 | -1.1873 | ~ | ~ | ~ | ~ | ~
I5 | B | 25 | 4.0275 | 0.088304 | 0.21917 | -3.5904 | 52.8879 | -1.0681 | 13.9486 | -0.82261 | ~ | ~ | ~ | ~ | ~
I5 | B | 26 | 5.6407 | 0.23085 | 3.9786 | -2.4989 | 47.9602 | -1.0799 | 24.1518 | -1.3084 | ~ | ~ | ~ | ~ | ~
I5 | B | 27 | 5.4414 | 0.3424 | 0.6578 | -3.2263 | 44.174 | -1.1355 | 24.6602 | -1.0424 | ~ | ~ | ~ | ~ | ~
I5 | B | 28 | 6.2598 | 0.46017 | 1.53 | -2.9078 | 51.3706 | -1.0596 | 21.6939 | -1.128 | ~ | ~ | ~ | ~ | ~
I5 | B | 29 | 4.211 | 0.33378 | 2.8944 | -2.6708 | 31.3097 | -0.84667 | 16.2466 | -0.97042 | ~ | ~ | ~ | ~ | ~
I5 | B | 30 | 6.8451 | 0.25527 | 2.0008 | -2.8423 | 42.2839 | -1.1281 | 25.7689 | -0.93238 | ~ | ~ | ~ | ~ | ~
I5 | B | 31 | 4.2718 | -0.41243 | 0.81732 | -3.1266 | 47.8076 | -1.0799 | 12.9193 | -0.44279 | ~ | ~ | ~ | ~ | ~
I5 | B | 32 | 7.2122 | -0.35677 | 0.93722 | -3.0717 | 86.7684 | -0.92911 | 13.6389 | -0.59234 | ~ | ~ | ~ | ~ | ~
M12 | A | 1 | 0.25212 | -0.48312 | 0.26167 | -2.9827 | 3.2205 | -2.1066 | 2.6491 | -2.7274 | 1.1746e-05 | 0.0001591 | 0.00047761 | 3 | 295
M12 | A | 2 | 0.29306 | -0.81023 | 0.44359 | -2.7899 | 3.3083 | -2.0159 | 3.8663 | -2.8166 | 1.5477e-05 | 0.00015483 | 0.00047761 | 3 | 295
M12 | A | 3 | 0.28762 | -0.82565 | 0.44952 | -2.802 | 3.7338 | -2.1939 | 1.5602 | -2.1132 | 1.5599e-05 | 0.0001946 | 0.00047761 | 3 | 295
M12 | A | 4 | 0.2958 | -0.81948 | 0.49776 | -2.7934 | 3.5204 | -1.9983 | 0.70776 | -1.2734 | 1.6613e-05 | 0.00016307 | 0.00047761 | 3 | 295
M12 | A | 5 | 0.25585 | -0.73174 | 0.43365 | -2.8258 | 3.7138 | -2.1324 | 1.4991 | -2.0201 | 1.4577e-05 | 0.00018637 | 0.00047761 | 3 | 295
M12 | A | 6 | 0.29531 | -0.89522 | 0.41906 | -2.8495 | 3.3285 | -2.0603 | 0.94074 | -1.709 | 1.5791e-05 | 0.00015992 | 0.00047761 | 3 | 295
M12 | A | 7 | 0.23887 | -0.72187 | 0.63099 | -2.6334 | 3.0597 | -2.1649 | 1.0561 | -1.8506 | 1.5631e-05 | 0.00015663 | 0.00047761 | 3 | 295
M12 | A | 8 | 0.17329 | -1.0733 | 0.42934 | -2.8091 | 2.5106 | -2.1595 | 2.1398 | -2.6547 | 1.2808e-05 | 0.00012809 | 0.00047761 | 3 | 295
M12 | A | 9 | 0.20317 | -0.74816 | 0.50777 | -2.7302 | 2.4652 | -2.2373 | 1.3494 | -2.2634 | 1.3783e-05 | 0.00013201 | 0.00047761 | 3 | 295
M12 | A | 10 | 0.1819 | -0.96074 | 0.31625 | -2.9602 | 2.4206 | -2.1963 | 1.5352 | -2.3352 | 1.1811e-05 | 0.00012634 | 0.00047761 | 3 | 295
M12 | A | 11 | 0.2106 | -0.78929 | 0.4429 | -2.8201 | 2.8626 | -2.1756 | 1.7794 | -2.2621 | 1.3699e-05 | 0.0001475 | 0.00047761 | 3 | 295
M12 | A | 12 | 0.26809 | -0.91227 | 0.36459 | -2.9361 | 3.2771 | -2.0122 | 1.4338 | -1.8789 | 1.4791e-05 | 0.00015304 | 0.00047761 | 3 | 295
M12 | A | 13 | 0.29333 | -0.80791 | 0.32292 | -3.006 | 3.6169 | -2.0508 | 1.4304 | -1.6634 | 1.4863e-05 | 0.0001728 | 0.00047761 | 3 | 295
M12 | A | 14 | 0.30651 | -0.98893 | 0.32024 | -3.0371 | 3.5896 | -1.9655 | 1.3743 | -1.8092 | 1.5806e-05 | 0.00016312 | 0.00047761 | 3 | 295
M12 | A | 15 | 0.29422 | -0.75857 | 0.26296 | -3.1183 | 3.7533 | -2.01 | 1.0034 | -1.3565 | 1.4192e-05 | 0.00017505 | 0.00047761 | 3 | 295
M12 | A | 16 | 0.26832 | -0.69568 | 0.179 | -3.2304 | 3.0987 | -1.9342 | 1.2291 | -1.7044 | 1.1836e-05 | 0.0001383 | 0.00047761 | 3 | 295
M12 | B | 1 | 0.26185 | -0.49315 | 0.18924 | -3.158 | 3.2044 | -2.0884 | 1.0703 | -1.8773 | 1.1244e-05 | 0.00015658 | 0.00047761 | 3 | 295
M12 | B | 2 | 0.29292 | -0.62935 | 0.31738 | -2.97 | 3.3811 | -1.982 | 0.37917 | -0.84511 | 1.4089e-05 | 0.00015513 | 0.00047761 | 3 | 295
M12 | B | 3 | 0.2638 | -0.86529 | 0.31732 | -2.987 | 3.8163 | -2.0981 | 1.7382 | -2.1131 | 1.3957e-05 | 0.00018757 | 0.00047761 | 3 | 295
M12 | B | 4 | 0.29284 | -0.73327 | 0.39362 | -2.9266 | 3.5169 | -1.9722 | 1.742 | -2.2353 | 1.5595e-05 | 0.00016044 | 0.00047761 | 3 | 295
M12 | B | 5 | 0.26421 | -0.76425 | 0.32163 | -2.9881 | 3.7837 | -2.0579 | 2.382 | -2.5265 | 1.3891e-05 | 0.00018154 | 0.00047761 | 3 | 295
M12 | B | 6 | 0.26225 | -0.99881 | 0.34739 | -2.9596 | 3.3208 | -2.0464 | 2.8862 | -2.7662 | 1.4646e-05 | 0.00015824 | 0.00047761 | 3 | 295
M12 | B | 7 | 0.21471 | -0.84957 | 0.52564 | -2.7399 | 3.0267 | -2.1512 | 2.2947 | -2.5684 | 1.4622e-05 | 0.00015364 | 0.00047761 | 3 | 295
M12 | B | 8 | 0.18759 | -0.98063 | 0.35063 | -2.9276 | 2.4986 | -2.1295 | 3.7355 | -3.1951 | 1.2482e-05 | 0.00012516 | 0.00047761 | 3 | 295
M12 | B | 9 | 0.18115 | -0.7999 | 0.40299 | -2.8619 | 2.4954 | -2.2073 | 1.2319 | -2.1912 | 1.2595e-05 | 0.00013114 | 0.00047761 | 3 | 295
M12 | B | 10 | 0.18006 | -1.0492 | 0.26683 | -3.0625 | 2.4923 | -2.1399 | 0.57577 | -1.398 | 1.1504e-05 | 0.00012564 | 0.00047761 | 3 | 295
M12 | B | 11 | 0.21596 | -0.80709 | 0.35193 | -2.9524 | 2.9332 | -2.1413 | 1.0879 | -1.8249 | 1.3175e-05 | 0.000148 | 0.00047761 | 3 | 295
M12 | B | 12 | 0.24067 | -0.91011 | 0.3129 | -3.0311 | 3.3307 | -1.9717 | 1.272 | -1.714 | 1.3726e-05 | 0.0001519 | 0.00047761 | 3 | 295
M12 | B | 13 | 0.26361 | -0.86879 | 0.26326 | -3.1257 | 3.6722 | -2.0947 | 0.743 | -1.2427 | 1.3723e-05 | 0.00018012 | 0.00047761 | 3 | 295
M12 | B | 14 | 0.29237 | -0.91397 | 0.27064 | -3.1408 | 3.5633 | -1.9358 | 0.28241 | -0.42406 | 1.4863e-05 | 0.00015918 | 0.00047761 | 3 | 295
M12 | B | 15 | 0.26358 | -0.75706 | 0.21528 | -3.2372 | 3.8283 | -1.9913 | 0.52266 | -0.83787 | 1.2987e-05 | 0.00017659 | 0.00047761 | 3 | 295
M12 | B | 16 | 0.25587 | -0.64284 | 0.1492 | -3.3434 | 3.0995 | -1.9049 | 3.4476 | -2.7023 | 1.1159e-05 | 0.00013604 | 0.00047761 | 3 | 295
M13 | A | 1 | 0.22992 | -0.39476 | 2.225e-05 | -6.0802 | 2.6091 | -2.1159 | 0.042878 | 0.61524 | 2.8345e-05 | 0.00068356 | 0.0031468 | 3 | 321
M13 | A | 2 | 0.23438 | -0.87514 | 0.0011109 | -4.9842 | 2.7459 | -2.0048 | 0.54021 | -2.3868 | 3.357e-05 | 0.00067326 | 0.0031468 | 3 | 321
M13 | A | 3 | 0.21909 | -0.96864 | 3.7146e-05 | -5.9247 | 2.9984 | -2.152 | 2.4779 | -3.1108 | 3.2488e-05 | 0.00080304 | 0.0031468 | 3 | 321
M13 | A | 4 | 0.24522 | -0.65255 | 0.0012035 | -4.9942 | 2.7913 | -2.0181 | 0.15784 | -0.91135 | 3.2574e-05 | 0.00068976 | 0.0031468 | 3 | 321
M13 | A | 5 | 0.2049 | -1.0786 | 3.9439e-05 | -5.925 | 3.1185 | -2.1249 | 2.2506 | -3.5301 | 3.1719e-05 | 0.00082148 | 0.0031468 | 3 | 321
M13 | A | 6 | 0.2788 | -0.82375 | 0.0017002 | -4.9382 | 2.7583 | -2.033 | 3.1404 | -3.9784 | 3.9209e-05 | 0.00068763 | 0.0031468 | 3 | 321
M13 | A | 7 | 0.21311 | -1.1248 | 3.3423e-05 | -5.9936 | 2.6807 | -2.147 | 1.2813 | -2.7148 | 3.3618e-05 | 0.00071576 | 0.0031468 | 3 | 321
M13 | A | 8 | 0.15918 | -0.90824 | 0.0010078 | -5.1144 | 2.1993 | -2.0454 | 1.1661 | -3.607 | 2.3076e-05 | 0.00055233 | 0.0031468 | 3 | 321
M13 | A | 9 | 0.18815 | -1.1222 | 2.5393e-05 | -6.0965 | 2.3691 | -2.1423 | 12.4045 | -5.3702 | 2.9649e-05 | 0.00063073 | 0.0031468 | 3 | 321
M13 | A | 10 | 0.16816 | -1.2018 | 0.0012343 | -5.0695 | 2.2461 | -2.0543 | 0.0099294 | 1.062 | 2.7401e-05 | 0.00056707 | 0.0031468 | 3 | 321
M13 | A | 11 | 0.19397 | -0.92508 | 3.2862e-05 | -6.065 | 2.7109 | -2.1612 | 34.8285 | -6.5268 | 2.8295e-05 | 0.00073014 | 0.0031468 | 3 | 321
M13 | A | 12 | 0.22686 | -1.0416 | 0.0019107 | -4.9137 | 3.0193 | -1.9814 | 0.28491 | -1.0572 | 3.4604e-05 | 0.00073026 | 0.0031468 | 3 | 321
M13 | A | 13 | 0.23953 | -0.99713 | 6.0277e-05 | -5.8106 | 3.226 | -2.1383 | 52.2115 | -6.5778 | 3.5908e-05 | 0.00085677 | 0.0031468 | 3 | 321
M13 | A | 14 | 0.28702 | -0.50912 | 0.0022912 | -4.8699 | 3.0677 | -2.0045 | 0.17724 | 0 | 3.6515e-05 | 0.00075204 | 0.0031468 | 3 | 321
M13 | A | 15 | 0.33408 | -0.91258 | 9.0523e-05 | -5.6618 | 3.9062 | -2.117 | 0.075195 | 0.48487 | 4.8508e-05 | 0.0010241 | 0.0031468 | 3 | 321
M13 | A | 16 | 0.36401 | -0.36274 | 0.001979 | -4.9107 | 3.5179 | -1.9492 | 0.16322 | -0.17681 | 4.4502e-05 | 0.00083515 | 0.0031468 | 3 | 321
M13 | B | 1 | 0.24352 | -0.78824 | 2.6729e-05 | -6.0045 | 2.7018 | -2.0983 | 4.5466 | -3.6774 | 3.3828e-05 | 0.00070034 | 0.0031468 | 3 | 321
M13 | B | 2 | 0.30274 | -0.75356 | 0.0042559 | -4.5765 | 2.8384 | -1.9806 | 0.40225 | -1.5505 | 4.1562e-05 | 0.00068618 | 0.0031468 | 3 | 321
M13 | B | 3 | 0.31858 | -0.73909 | 4.3845e-05 | -5.8404 | 3.1664 | -2.136 | 6.4179 | -3.9353 | 4.3524e-05 | 0.00083977 | 0.0031468 | 3 | 321
M13 | B | 4 | 0.31224 | -0.81669 | 0.0034121 | -4.6803 | 3.0334 | -1.965 | 1.1309 | -2.8265 | 4.3802e-05 | 0.00072673 | 0.0031468 | 3 | 321
M13 | B | 5 | 0.30857 | -0.77868 | 0.00011865 | -5.5434 | 3.3997 | -2.0783 | 0.78938 | -2.1717 | 4.2723e-05 | 0.00087072 | 0.0031468 | 3 | 321
M13 | B | 6 | 0.25722 | -0.81961 | 0.0019416 | -4.9012 | 2.9884 | -1.9822 | 0.82052 | -2.231 | 3.6121e-05 | 0.00072313 | 0.0031468 | 3 | 321
M13 | B | 7 | 0.30636 | -0.76272 | 3.88e-05 | -5.9194 | 2.7853 | -2.124 | 9.8773 | -5.1177 | 4.2189e-05 | 0.0007333 | 0.0031468 | 3 | 321
M13 | B | 8 | 0.24445 | -0.80572 | 0.0012661 | -5.0376 | 2.1923 | -2.0451 | 0.28447 | -1.0649 | 3.4162e-05 | 0.00055047 | 0.0031468 | 3 | 321
M13 | B | 9 | 0.21968 | -0.94566 | 2.7221e-05 | -6.0446 | 2.443 | -2.1298 | 0.27829 | -1.2959 | 3.2293e-05 | 0.00064548 | 0.0031468 | 3 | 321
M13 | B | 10 | 0.25003 | -0.77635 | 0.0012152 | -5.0717 | 2.3346 | -2.0319 | 0.10616 | -0.78542 | 3.4591e-05 | 0.00058164 | 0.0031468 | 3 | 321
M13 | B | 11 | 0.29954 | -0.84416 | 4.0206e-05 | -5.9811 | 2.7458 | -2.15 | 12.3746 | -4.8557 | 4.2429e-05 | 0.00073446 | 0.0031468 | 3 | 321
M13 | B | 12 | 0.31226 | -0.67544 | 0.0025543 | -4.8556 | 3.082 | -1.966 | 0.42764 | -1.4455 | 4.1783e-05 | 0.00073877 | 0.0031468 | 3 | 321
M13 | B | 13 | 0.31414 | -0.78706 | 4.0256e-05 | -6.0135 | 3.5055 | -2.0851 | 3.5277 | -3.5462 | 4.362e-05 | 0.00090149 | 0.0031468 | 3 | 321
M13 | B | 14 | 0.49067 | -0.59082 | 0.0027305 | -4.8548 | 3.2882 | -1.979 | 0.88213 | -1.4769 | 6.3946e-05 | 0.00079419 | 0.0031468 | 3 | 321
M13 | B | 15 | 0.43119 | -0.37873 | 5.6431e-05 | -5.8733 | 3.8582 | -2.1017 | 0.22903 | -0.45813 | 5.2933e-05 | 0.0010022 | 0.0031468 | 3 | 321
M13 | B | 16 | 0.39756 | -0.59322 | 0.0018913 | -4.9331 | 3.7489 | -1.911 | 0.58335 | -1.126 | 5.185e-05 | 0.00087072 | 0.0031468 | 3 | 321
M14 | A | 1 | 0.56481 | 0.48569 | 4.0069 | -2.3136 | 4.3906 | -1.9784 | 1.2628 | -0.68092 | ~ | ~ | ~ | ~ | ~
M14 | A | 2 | 0.55942 | 0.73429 | 4.684 | -2.2388 | 1.4116 | -2.7383 | 0.45038 | 0.15694 | ~ | ~ | ~ | ~ | ~
M14 | A | 3 | 1.6858 | -2.1337 | 4.1347 | -2.2895 | 4.8409 | -1.9322 | 0.87324 | -0.68348 | ~ | ~ | ~ | ~ | ~
M14 | A | 4 | 0.89462 | -0.56789 | 3.9696 | -2.2616 | 2.2442 | -2.3581 | 0.25326 | 0.45414 | ~ | ~ | ~ | ~ | ~
M14 | A | 5 | 0.68183 | -0.5693 | 3.8654 | -2.3012 | 4.6783 | -1.9017 | 0.29267 | 0.42516 | ~ | ~ | ~ | ~ | ~
M14 | A | 6 | 0.85976 | -1.2273 | 4.2507 | -2.2088 | 4.2928 | -1.9398 | 0.1883 | 0.53873 | ~ | ~ | ~ | ~ | ~
M14 | A | 7 | 0.50142 | -0.67142 | 4.0697 | -2.3147 | 5.1726 | -1.8863 | 0.54498 | 0.21561 | ~ | ~ | ~ | ~ | ~
M14 | A | 8 | 0.48507 | -0.36821 | 4.2594 | -2.2197 | 4.2373 | -1.956 | 0.66535 | -0.59398 | ~ | ~ | ~ | ~ | ~
M14 | A | 9 | 1.6074 | -2.6966 | 3.4693 | -2.3384 | 4.2188 | -1.917 | 0.34537 | 0.36295 | ~ | ~ | ~ | ~ | ~
M14 | A | 10 | 0.94749 | -1.0928 | 4.2836 | -2.2005 | 2.6089 | -2.2363 | 0.33838 | 0.10106 | ~ | ~ | ~ | ~ | ~
M14 | A | 11 | 0.27607 | 1.7292 | 4.0948 | -2.3339 | 5.1011 | -1.8604 | 0.28261 | 0.5464 | ~ | ~ | ~ | ~ | ~
M14 | A | 12 | 1.3956 | -1.3011 | 4.6931 | -2.2362 | 2.8954 | -2.1965 | 0.47087 | 0.074704 | ~ | ~ | ~ | ~ | ~
M14 | A | 13 | 0.81428 | -0.66908 | 3.743 | -2.3491 | 4.6115 | -1.8974 | 0.44498 | -0.053612 | ~ | ~ | ~ | ~ | ~
M14 | A | 14 | 0.86808 | -0.77287 | 4.2577 | -2.2814 | 3.4824 | -2.0202 | 0.3453 | 0.20459 | ~ | ~ | ~ | ~ | ~
M14 | A | 15 | 0.4211 | 0.53495 | 3.8454 | -2.4121 | 4.4361 | -1.8808 | 0.8384 | -0.4128 | ~ | ~ | ~ | ~ | ~
M14 | A | 16 | 0.59866 | 0.49564 | 4.4931 | -2.2304 | 1.2039 | -2.6072 | 0.72096 | 0.02322 | ~ | ~ | ~ | ~ | ~
M14 | B | 1 | 0.78826 | 0.22426 | 4.8012 | -2.2287 | 4.6189 | -1.9984 | 0.53036 | 0.38834 | ~ | ~ | ~ | ~ | ~
M14 | B | 2 | 1.2717 | -1.178 | 5.4142 | -2.1776 | 3.0172 | -2.2612 | 0.67116 | 0.021774 | ~ | ~ | ~ | ~ | ~
M14 | B | 3 | 0.5328 | -0.25539 | 4.0461 | -2.3072 | 4.8776 | -1.973 | 0.31333 | 0.58845 | ~ | ~ | ~ | ~ | ~
M14 | B | 4 | 0.24724 | 0.36601 | 3.3925 | -2.3352 | 3.2838 | -2.1585 | 0.61145 | -0.32104 | ~ | ~ | ~ | ~ | ~
M14 | B | 5 | 0.89811 | -0.76211 | 3.9003 | -2.2988 | 4.8579 | -1.9253 | 0.6902 | -0.41537 | ~ | ~ | ~ | ~ | ~
M14 | B | 6 | 0.61638 | -0.68558 | 4.4439 | -2.1898 | 2.4976 | -2.3689 | 0.65123 | -0.18848 | ~ | ~ | ~ | ~ | ~
M14 | B | 7 | 0.40959 | 0.5363 | 4.1515 | -2.2866 | 5.2709 | -1.9095 | 0.49414 | 0.016846 | ~ | ~ | ~ | ~ | ~
M14 | B | 8 | 1.7999 | -2.0261 | 4.3992 | -2.2022 | 7.7634 | -1.5761 | 0.48894 | 0.024831 | ~ | ~ | ~ | ~ | ~
M14 | B | 9 | 0.53573 | -0.85551 | 3.943 | -2.2558 | 4.2944 | -1.9481 | 0.52521 | -0.14726 | ~ | ~ | ~ | ~ | ~
M14 | B | 10 | 0.58436 | 0.2152 | 4.3865 | -2.1734 | 3.4399 | -2.0731 | 0.54289 | -0.10167 | ~ | ~ | ~ | ~ | ~
M14 | B | 11 | 0.67334 | 0.13293 | 4.068 | -2.3376 | 5.3947 | -1.8679 | 0.78331 | -0.31178 | ~ | ~ | ~ | ~ | ~
M14 | B | 12 | 0.3995 | 0.29467 | 4.6173 | -2.2402 | 7.6847 | -1.558 | 0.5404 | -0.29044 | ~ | ~ | ~ | ~ | ~
M14 | B | 13 | 1.4962 | -1.3736 | 4.0578 | -2.3024 | 4.4809 | -1.9136 | 0.76725 | -0.077025 | ~ | ~ | ~ | ~ | ~
M14 | B | 14 | 0.91395 | -0.46462 | 6.0272 | -2.0688 | 4.2011 | -1.9587 | 0.50807 | 0.34153 | ~ | ~ | ~ | ~ | ~
M14 | B | 15 | 0.53276 | 0.62683 | 5.4461 | -2.2156 | 5.3088 | -1.8 | 0.47064 | 0.56818 | ~ | ~ | ~ | ~ | ~
M14 | B | 16 | 5.2505 | -3.0865 | 4.6702 | -2.2019 | 5.8337 | -1.7128 | 0.75549 | -0.18295 | ~ | ~ | ~ | ~ | ~
M15 | A | 1 | 1.6445 | -1.2861 | 3.4239 | -2.2794 | 5.2459 | -2.0406 | 0.32731 | 0.98669 | 0.00010813 | 0.00014765 | 0.0096859 | 3 | 160
M15 | A | 2 | 1.3671 | -0.85035 | 3.7077 | -2.1859 | 3.32 | -2.3564 | 0.63798 | 0.15696 | 0.00010878 | 0.00011437 | 0.0096859 | 3 | 160
M15 | A | 3 | 1.8165 | -2.0718 | 3.4898 | -2.2324 | 4.8052 | -2.0216 | 0.95306 | -0.40001 | 0.00010628 | 0.00013403 | 0.0096859 | 3 | 160
M15 | A | 4 | 0.78674 | -0.61816 | 3.9176 | -2.1479 | 3.9527 | -2.2496 | 0.66631 | -0.041957 | 0.00010932 | 0.00012696 | 0.0096859 | 3 | 160
M15 | A | 5 | 2.2177 | -2.5056 | 3.523 | -2.2724 | 4.8822 | -2.0512 | 1.2886 | -0.63753 | 0.00011098 | 0.00013878 | 0.0096859 | 3 | 160
M15 | A | 6 | 1.2655 | -1.578 | 4.0889 | -2.1707 | 2.9735 | -2.4842 | 0.76584 | -0.21406 | 0.00011692 | 0.00011162 | 0.0096859 | 3 | 160
M15 | A | 7 | 1.0314 | -0.61487 | 3.8041 | -2.2417 | 4.9408 | -2.1065 | 1.0326 | -0.45592 | 0.00011453 | 0.00014513 | 0.0096859 | 3 | 160
M15 | A | 8 | 0.88105 | -0.93996 | 3.8984 | -2.1577 | 3.6053 | -2.2677 | 0.72319 | -0.15798 | 0.00010971 | 0.00011722 | 0.0096859 | 3 | 160
M15 | A | 9 | 0.47164 | 0.18052 | 3.3372 | -2.2687 | 4.121 | -2.1161 | 0.2558 | 0.54038 | 0.00010079 | 0.00012142 | 0.0096859 | 3 | 160
M15 | A | 10 | 0.20157 | 0.74431 | 3.8601 | -2.1655 | 2.8132 | -2.4196 | 0.44432 | 0.28213 | 0.00010672 | 0.00010098 | 0.0096859 | 3 | 160
M15 | A | 11 | 1.4166 | -1.3898 | 3.3342 | -2.3025 | 4.581 | -2.052 | 0.91524 | -0.41162 | 0.00010613 | 0.00013015 | 0.0096859 | 3 | 160
M15 | A | 12 | 0.5045 | -0.18267 | 3.9752 | -2.1641 | 1.9578 | -2.6671 | 0.4029 | 0.40637 | 0.00011105 | 8.3361e-05 | 0.0096859 | 3 | 160
M15 | A | 13 | 0.36364 | 1.0228 | 3.2552 | -2.2704 | 3.9786 | -2.0656 | 0.5551 | 0.11556 | 9.8316e-05 | 0.00011388 | 0.0096859 | 3 | 160
M15 | A | 14 | 0.62446 | 0.11013 | 3.5383 | -2.1948 | 2.1162 | -2.5323 | 0.42166 | 0.36646 | 0.00010208 | 8.2051e-05 | 0.0096859 | 3 | 160
M15 | A | 15 | 0.87227 | -0.20662 | 3.1171 | -2.3297 | 4.5271 | -2.0056 | 0.93253 | -0.22472 | 0.00010028 | 0.00012514 | 0.0096859 | 3 | 160
M15 | A | 16 | 0.63245 | 0.53876 | 3.5323 | -2.1998 | 2.9199 | -2.2752 | 1.0201 | -0.15661 | 0.00010264 | 9.5743e-05 | 0.0096859 | 3 | 160
M15 | B | 1 | 1.0198 | 0.47105 | 3.8062 | -2.22 | 5.4111 | -1.9222 | 1.045 | -0.25331 | 0.00011407 | 0.00014234 | 0.0096859 | 3 | 160
M15 | B | 2 | 0.60013 | 0.59835 | 3.7849 | -2.168 | 3.8097 | -2.1999 | 0.79805 | -0.21254 | 0.00010719 | 0.00011866 | 0.0096859 | 3 | 160
M15 | B | 3 | 0.35928 | 0.79253 | 3.8503 | -2.1801 | 4.3864 | -2.0605 | 1.1421 | -0.6141 | 0.00010855 | 0.00012539 | 0.0096859 | 3 | 160
M15 | B | 4 | 0.79437 | -0.69115 | 3.9127 | -2.1515 | 4.369 | -2.1031 | 0.48287 | 0.31304 | 0.00010944 | 0.00012787 | 0.0096859 | 3 | 160
M15 | B | 5 | 0.53199 | -0.093184 | 3.3448 | -2.3041 | 4.5508 | -2.068 | 0.7662 | -0.35397 | 0.00010374 | 0.00013045 | 0.0096859 | 3 | 160
M15 | B | 6 | 0.61645 | -0.2056 | 3.8486 | -2.2065 | 4.1462 | -2.1752 | 0.4672 | 0.049541 | 0.00011138 | 0.00012687 | 0.0096859 | 3 | 160
M15 | B | 7 | 0.47481 | -0.03693 | 3.758 | -2.2542 | 4.9419 | -1.9983 | 0.83416 | -0.38104 | 0.00011191 | 0.00013583 | 0.0096859 | 3 | 160
M15 | B | 8 | 0.30851 | 0.96729 | 3.9261 | -2.1572 | 3.6166 | -2.1931 | 0.23462 | 0.76924 | 0.00010865 | 0.00011183 | 0.0096859 | 3 | 160
M15 | B | 9 | 0.54178 | 0.11158 | 2.9698 | -2.3389 | 3.8038 | -2.0669 | 0.68941 | -0.23087 | 9.4902e-05 | 0.00010903 | 0.0096859 | 3 | 160
M15 | B | 10 | 0.33619 | 0.63713 | 3.6421 | -2.2073 | 3.0731 | -2.2892 | 0.42947 | 0.15908 | 0.00010457 | 0.00010121 | 0.0096859 | 3 | 160
M15 | B | 11 | 0.534 | -0.17528 | 3.8637 | -2.2332 | 4.5109 | -1.993 | 0.83106 | -0.34517 | 0.00011352 | 0.00012365 | 0.0096859 | 3 | 160
M15 | B | 12 | 0.66621 | 0.023395 | 3.9168 | -2.1824 | 3.6485 | -2.1436 | 0.56587 | -0.10399 | 0.00011183 | 0.00010957 | 0.0096859 | 3 | 160
M15 | B | 13 | 0.54041 | 0.15296 | 2.9895 | -2.3232 | 4.4032 | -1.9455 | 0.4376 | 0.36287 | 9.4471e-05 | 0.00011718 | 0.0096859 | 3 | 160
M15 | B | 14 | 0.45378 | 0.29126 | 3.2687 | -2.2497 | 2.3511 | -2.3943 | 0.62407 | -0.24295 | 9.7421e-05 | 8.3122e-05 | 0.0096859 | 3 | 160
M15 | B | 15 | 0.99803 | -0.51284 | 3.0951 | -2.3424 | 4.5789 | -1.9434 | 0.99821 | -0.41288 | 0.0001008 | 0.00012199 | 0.0096859 | 3 | 160
M15 | B | 16 | 1.2172 | -0.44239 | 3.2675 | -2.2578 | 3.4534 | -2.1049 | 0.96938 | -0.30061 | 0.0001012 | 0.00010155 | 0.0096859 | 3 | 160
M16A | A | 1 | 1.6231 | -0.055635 | 0.62613 | -2.8966 | 5.0371 | -2.018 | 0.74518 | 0.36471 | 0.0023212 | 0.0082214 | 0.0089523 | 3 | 318
M16A | A | 2 | 1.02 | -0.21545 | 2.5651 | -2.2693 | 4.9984 | -1.9941 | 0.49312 | 0.81287 | 0.0043053 | 0.0080445 | 0.0089523 | 3 | 318
M16A | A | 3 | 1.4405 | -0.16133 | 0.94149 | -2.7373 | 4.155 | -2.1259 | 1.6679 | -0.17749 | 0.0027276 | 0.0072338 | 0.0089523 | 3 | 318
M16A | A | 4 | 1.2428 | -0.050874 | 2.6205 | -2.2887 | 4.4564 | -2.0914 | 1.393 | -0.27957 | 0.0045599 | 0.0075984 | 0.0089523 | 3 | 318
M16A | A | 5 | 0.80843 | 0.58191 | 0.10474 | -3.5963 | 3.5692 | -2.2333 | 1.35 | -0.50809 | 0.00088131 | 0.0066406 | 0.0089523 | 3 | 318
M16A | A | 6 | 0.81648 | 0.43776 | 0.6431 | -2.9691 | 3.6552 | -2.1547 | 0.97655 | 0.040305 | 0.0021177 | 0.0064769 | 0.0089523 | 3 | 318
M16A | A | 7 | 1.0082 | -0.15247 | 0.57101 | -2.9786 | 3.6284 | -2.1256 | 0.24333 | 1.1332 | 0.0020002 | 0.0063162 | 0.0089523 | 3 | 318
M16A | A | 8 | 1.1863 | -0.59654 | 1.2655 | -2.6618 | 4.0216 | -2.0341 | 0.22165 | 1.0587 | 0.0031362 | 0.0066266 | 0.0089523 | 3 | 318
M16A | A | 9 | 1.153 | -0.46779 | 0.25124 | -3.3919 | 3.6361 | -2.1396 | 2.8419 | -1.1097 | 0.0014789 | 0.0063837 | 0.0089523 | 3 | 318
M16A | A | 10 | 0.63475 | 0.30062 | 2.2859 | -2.3454 | 3.295 | -2.1114 | 1.3493 | -0.28926 | 0.0039463 | 0.0056866 | 0.0089523 | 3 | 318
M16A | A | 11 | 1.399 | -1.0184 | 0.25567 | -3.3133 | 3.7957 | -2.0671 | 0.60012 | 0.41522 | 0.001515 | 0.0063783 | 0.0089523 | 3 | 318
M16A | A | 12 | 0.41156 | 1.3773 | 1.1524 | -2.6786 | 3.116 | -2.1339 | 1.6948 | -0.18912 | 0.0026427 | 0.0054515 | 0.0089523 | 3 | 318
M16A | A | 13 | 1.2271 | -0.54384 | 0.53122 | -3.0156 | 2.9392 | -2.1582 | 2.5799 | -0.9415 | 0.0020194 | 0.0052193 | 0.0089523 | 3 | 318
M16A | A | 14 | 0.9178 | -0.49465 | 1.4603 | -2.5536 | 3.4432 | -2.024 | 0.23826 | 1.3366 | 0.0031659 | 0.0056399 | 0.0089523 | 3 | 318
M16A | A | 15 | 1.1613 | 0.68597 | 0.16569 | -3.5101 | 4.2531 | -1.9528 | 2.0201 | -0.17895 | 0.0012876 | 0.0066829 | 0.0089523 | 3 | 318
M16A | A | 16 | 0.51656 | 0.88592 | 0.57183 | -3.0095 | 3.4635 | -2.0201 | 0.18819 | 1.5946 | 0.0018498 | 0.0056599 | 0.0089523 | 3 | 318
M16A | B | 1 | 1.0222 | 0.72849 | 1.6237 | -2.4222 | 4.3662 | -2.0166 | 0.57389 | 0.73321 | 0.0032674 | 0.0071204 | 0.0089523 | 3 | 318
M16A | B | 2 | 1.0719 | 0.93598 | 2.1704 | -2.2792 | 4.2196 | -2.164 | 1.7409 | -0.26684 | 0.0038302 | 0.0075199 | 0.0089523 | 3 | 318
M16A | B | 3 | 0.48873 | 0.90807 | 0.40447 | -3.0214 | 3.3731 | -2.2596 | 0.93479 | 0.11587 | 0.0013836 | 0.0063802 | 0.0089523 | 3 | 318
M16A | B | 4 | 1.0352 | 0.049831 | 3.1748 | -2.1189 | 3.454 | -2.2521 | 1.7175 | -0.60961 | 0.0046964 | 0.0065026 | 0.0089523 | 3 | 318
M16A | B | 5 | 1.5021 | -0.40362 | 0.071894 | -3.7814 | 3.3734 | -2.2867 | 3.0295 | -1.1605 | 0.0010585 | 0.0064913 | 0.0089523 | 3 | 318
M16A | B | 6 | 1.1592 | -0.23329 | 0.96401 | -2.7425 | 3.6671 | -2.1494 | 0.91438 | -0.14293 | 0.0026554 | 0.0064769 | 0.0089523 | 3 | 318
M16A | B | 7 | 0.72893 | -0.48164 | 0.5859 | -2.9031 | 3.7918 | -2.1354 | 1.6055 | -0.4365 | 0.0018058 | 0.0066401 | 0.0089523 | 3 | 318
M16A | B | 8 | 1.0518 | 0.026931 | 1.3109 | -2.5812 | 4.3329 | -2.0295 | 0.65351 | 0.24837 | 0.003022 | 0.0071201 | 0.0089523 | 3 | 318
M16A | B | 9 | 1.4541 | -0.13579 | 0.60573 | -2.9207 | 4.1349 | -2.0621 | 0.4672 | 0.71176 | 0.0022188 | 0.0069277 | 0.0089523 | 3 | 318
M16A | B | 10 | 1.5264 | -1.4019 | 2.1385 | -2.3229 | 2.9296 | -2.2161 | 1.2276 | -0.45062 | 0.0039818 | 0.005392 | 0.0089523 | 3 | 318
M16A | B | 11 | 0.59824 | -0.063028 | 0.25942 | -3.2657 | 4.265 | -1.9809 | 0.47451 | 0.41809 | 0.0011711 | 0.0068116 | 0.0089523 | 3 | 318
M16A | B | 12 | 1.6501 | -0.97418 | 1.7207 | -2.4397 | 4.5437 | -1.9685 | 0.85942 | 0.41469 | 0.0036647 | 0.0072045 | 0.0089523 | 3 | 318
M16A | B | 13 | 0.81383 | 1.1145 | 1.2495 | -2.5964 | 3.7942 | -2.0524 | 0.8107 | 0.14116 | 0.0028767 | 0.0063202 | 0.0089523 | 3 | 318
M16A | B | 14 | 1.221 | -0.27589 | 1.1071 | -2.6617 | 3.1276 | -2.1787 | 0.46133 | 0.5678 | 0.002838 | 0.0056243 | 0.0089523 | 3 | 318
M16A | B | 15 | 1.3665 | -0.53434 | 0.076014 | -3.8088 | 2.4663 | -2.2499 | 1.1767 | 0.2245 | 0.0010247 | 0.0046365 | 0.0089523 | 3 | 318
M16A | B | 16 | 1.0984 | 0.4162 | 1.51 | -2.4819 | 2.9772 | -2.0594 | 3.3755 | -0.81389 | 0.0032208 | 0.0049801 | 0.0089523 | 3 | 318
M16B | A | 1 | 1.9923 | -0.15447 | 3.0776 | -2.0295 | 4.5904 | -2.0718 | 1.3699 | 0.14669 | 0.0047116 | 0.0077354 | 0.0089523 | 3 | 318
M16B | A | 2 | 1.7221 | -1.6237 | 1.4692 | -2.4821 | 2.8988 | -2.4176 | 1.1949 | -0.28785 | 0.0033279 | 0.0060717 | 0.0089523 | 3 | 318
M16B | A | 3 | 1.2191 | 0.11875 | 1.9707 | -2.2244 | 4.1387 | -2.0899 | 1.0733 | -0.46548 | 0.0034172 | 0.0070503 | 0.0089523 | 3 | 318
M16B | A | 4 | 1.533 | -0.59821 | 0.58914 | -2.8593 | 4.1695 | -2.1888 | 1.3939 | -0.027396 | 0.0021078 | 0.0075454 | 0.0089523 | 3 | 318
M16B | A | 5 | 0.85407 | 0.065161 | 3.0077 | -2.0993 | 4.096 | -2.0922 | 1.7449 | -0.55032 | 0.0043293 | 0.0069874 | 0.0089523 | 3 | 318
M16B | A | 6 | 1.3695 | -0.51738 | 1.6843 | -1.908 | 2.9226 | -2.3061 | 0.95305 | 0.077309 | 0.0024953 | 0.005694 | 0.0089523 | 3 | 318
M16B | A | 7 | 1.0605 | -0.11342 | 2.4642 | -2.1877 | 3.9064 | -2.0834 | 1.8745 | -0.54497 | 0.0039419 | 0.0066286 | 0.0089523 | 3 | 318
M16B | A | 8 | 1.2615 | -0.32539 | 0.98695 | -2.5407 | 3.7155 | -2.1698 | 0.87937 | 0.076945 | 0.0024101 | 0.0066451 | 0.0089523 | 3 | 318
M16B | A | 9 | 1.6559 | -0.74142 | 3.2133 | -2.0313 | 4.5436 | -2.0295 | 2.212 | -0.42967 | 0.0046964 | 0.0074663 | 0.0089523 | 3 | 318
M16B | A | 10 | 1.119 | -0.44411 | 0.12633 | -3.4092 | 3.0802 | -2.2912 | 0.69729 | 0.078957 | 0.00099001 | 0.0059442 | 0.0089523 | 3 | 318
M16B | A | 11 | 1.6283 | -1.2518 | 2.5287 | -2.1469 | 3.5494 | -2.0878 | 1.2056 | -0.22127 | 0.0041165 | 0.0060387 | 0.0089523 | 3 | 318
M16B | A | 12 | 1.5441 | -0.012465 | 0.8124 | -2.7111 | 4.1059 | -2.0641 | 1.3446 | -0.20022 | 0.0024614 | 0.0068873 | 0.0089523 | 3 | 318
M16B | A | 13 | 1.2042 | -0.85547 | 2.6969 | -2.1163 | 3.845 | -2.0263 | 2.5166 | -0.77667 | 0.0040914 | 0.0063064 | 0.0089523 | 3 | 318
M16B | A | 14 | 1.5314 | 0.26387 | 0.98489 | -2.637 | 5.3353 | -1.9199 | 3.1238 | -0.61843 | 0.0027209 | 0.0082267 | 0.0089523 | 3 | 318
M16B | A | 15 | 1.2948 | 1.0688 | 1.1683 | -2.5018 | 4.7897 | -1.8643 | 1.2884 | -0.015534 | 0.0027944 | 0.0071568 | 0.0089523 | 3 | 318
M16B | A | 16 | 2.0506 | -0.66534 | 1.3684 | -2.4587 | 3.0699 | -2.2805 | 0.615 | 0.77426 | 0.0032842 | 0.0058843 | 0.0089523 | 3 | 318
M16B | B | 1 | 1.4047 | -0.41996 | 0.34292 | -2.9812 | 3.3393 | -2.2389 | 1.223 | -0.21751 | 0.0015509 | 0.0062346 | 0.0089523 | 3 | 318
M16B | B | 2 | 1.8402 | -0.24035 | 1.594 | -2.3436 | 2.8004 | -2.4304 | 1.0293 | -0.18105 | 0.003366 | 0.005915 | 0.0089523 | 3 | 318
M16B | B | 3 | 0.86141 | 0.49311 | 2.5885 | -2.0524 | 4.3135 | -2.0737 | 1.007 | -0.11916 | 0.0036798 | 0.0072773 | 0.0089523 | 3 | 318
M16B | B | 4 | 1.4197 | 0.0077809 | 1.4391 | -2.4445 | 3.7837 | -2.2743 | 0.6968 | 0.37034 | 0.0031421 | 0.0072237 | 0.0089523 | 3 | 318
M16B | B | 5 | 1.2966 | -1.3521 | 3.5942 | -1.9342 | 4.1991 | -2.1579 | 0.81417 | 0.28471 | 0.0046643 | 0.0074553 | 0.0089523 | 3 | 318
M16B | B | 6 | 0.76927 | 0.27938 | 2.1854 | -2.1994 | 3.8953 | -2.2106 | 1.3945 | -0.89919 | 0.0034613 | 0.007145 | 0.0089523 | 3 | 318
M16B | B | 7 | 0.47943 | 0.38008 | 3.5062 | -1.9436 | 3.7069 | -2.1341 | 1.3448 | -0.43421 | 0.0042834 | 0.0064862 | 0.0089523 | 3 | 318
M16B | B | 8 | 1.4348 | -1.5625 | 1.7359 | -2.3419 | 2.4681 | -2.5013 | 0.89462 | -0.16914 | 0.0033516 | 0.0054655 | 0.0089523 | 3 | 318
M16B | B | 9 | 0.79321 | 0.56316 | 3.4179 | -1.9631 | 4.3426 | -2.0173 | 3.4308 | -1.269 | 0.004401 | 0.0070849 | 0.0089523 | 3 | 318
M16B | B | 10 | 1.2365 | -1.1299 | 1.4218 | -2.3967 | 3.2433 | -2.2366 | 0.34819 | 0.85857 | 0.0028847 | 0.0060466 | 0.0089523 | 3 | 318
M16B | B | 11 | 1.0285 | -0.05563 | 3.3792 | -1.9391 | 3.8203 | -2.0491 | 2.7057 | -1.0295 | 0.0043639 | 0.0063509 | 0.0089523 | 3 | 318
M16B | B | 12 | 1.5006 | -0.49242 | 0.68867 | -2.7685 | 3.3119 | -2.2045 | 1.9141 | -0.631 | 0.0022185 | 0.0060518 | 0.0089523 | 3 | 318
M16B | B | 13 | 0.89439 | -0.0235 | 2.2448 | -2.1646 | 4.3125 | -1.9786 | 1.1878 | -0.067324 | 0.0035039 | 0.0068781 | 0.0089523 | 3 | 318
M16B | B | 14 | 1.6075 | -0.83572 | 1.2972 | -2.4815 | 4.3795 | -2.0722 | 2.691 | -1.0207 | 0.0030014 | 0.0073819 | 0.0089523 | 3 | 318
M16B | B | 15 | 0.58903 | 0.79969 | 3.0389 | -1.9591 | 3.9816 | -2.0047 | 2.5324 | -0.99175 | 0.0038535 | 0.006448 | 0.0089523 | 3 | 318
M16B | B | 16 | 1.8277 | -0.7591 | 2.2055 | -2.0206 | 3.9521 | -2.1152 | 2.1745 | -0.57011 | 0.0034981 | 0.0068363 | 0.0089523 | 3 | 318

## RSB NFR
Band | HAM | Baffle Locations |  | Mean | [7 empty] | Median | [7 empty] | Sensor Performance Compared with Requirement
|  |  |  | Harvy Shack Fit Parameters | [7 empty] | Harvy Shack Fit Parameters | [7 empty] | Mean |  | Median |  | Max Allowed Lscat | Angle (mrad) from Target | L Bright Target
|  | Left | Right | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Scattered Radiance
|  |  |  | A | M | A | M | A | M | A | M | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright | Left of Lbright | Right of Lbright
I1 | A | -13 | 61 | 0.53531 | -1.6031 | 5.115 | -1.4647 | 5.6944 | -1.4441 | 1.5684 | -2.0141 | 0.54 | -1.56 | 5.255 | -1.46 | 5.675 | -1.45 | 1.4 | -2.07 | 0.0065494 | 0.051384 | 0.00672 | 0.052585 | 0.22 | 6 | 115
I1 | B | -13 | 61 | 0.6025 | -1.5859 | 5.5766 | -1.4581 | 6.1719 | -1.4606 | 1.4734 | -1.8431 | 0.62 | -1.57 | 5.585 | -1.45 | 6.215 | -1.465 | 1.3 | -1.905 | 0.0074228 | 0.049287 | 0.007735 | 0.050495 | 0.22 | 6 | 115
I2 | A | -15 | 59 | 0.5775 | -1.6231 | 4.2328 | -1.5209 | 5.2403 | -1.3816 | 4.57 | -3.0113 | 0.58 | -1.63 | 4.32 | -1.51 | 5.28 | -1.395 | 4.48 | -3.075 | 0.0074981 | 0.050084 | 0.00746 | 0.05163 | 0.25 | 6 | 124
I2 | B | -15 | 59 | 0.6475 | -1.6056 | 4.5934 | -1.5156 | 5.685 | -1.3987 | 5.1488 | -2.9922 | 0.65 | -1.61 | 4.68 | -1.505 | 5.72 | -1.41 | 5.025 | -3.02 | 0.0084763 | 0.04804 | 0.008515 | 0.04959 | 0.25 | 6 | 124
I3 | A | -20 | 54 | 0.1575 | -1.2544 | 2.4944 | -1.5544 | 2.7038 | -1.3878 | 6.8419 | -4.545 | 0.165 | -1.3 | 2.335 | -1.555 | 2.62 | -1.38 | 6.885 | -4.66 | 0.0022325 | 0.020925 | 0.00194 | 0.021615 | 0.073 | 6 | 86.1
I3 | B | -20 | 54 | 0.16594 | -1.2622 | 2.5628 | -1.5672 | 2.8147 | -1.3888 | 7.4731 | -4.5756 | 0.15 | -1.27 | 2.225 | -1.55 | 2.43 | -1.37 | 7.065 | -4.63 | 0.0023813 | 0.019956 | 0.002265 | 0.020865 | 0.073 | 6 | 86.1
M1 | A | -4 | 72 | 1.1187 | -1.4987 | 0 | -5.3131 | 15.4094 | -1.2619 | 9.0675 | -2.4269 | 1.125 | -1.495 | 0 | -5.4 | 15.495 | -1.265 | 4.82 | -2.39 | 0.013504 | 0.1708 | 0.013635 | 0.17319 | 0.449 | 6 | 162
M1 | B | -4 | 72 | 1.1694 | -1.4663 | 0 | -4.9031 | 15.3506 | -1.26 | 9.7087 | -2.3244 | 1.21 | -1.47 | 0 | -4.96 | 15.21 | -1.26 | 5.555 | -2.625 | 0.014637 | 0.16154 | 0.015195 | 0.1632 | 0.449 | 6 | 162
M2 | A | -6 | 69 | 1.2194 | -1.0631 | 0.045625 | -3.7881 | 13.1306 | -1.3019 | 19.2019 | -1.2663 | 1.2 | -1.075 | 0.05 | -3.76 | 12.96 | -1.31 | 2.795 | -1.55 | 0.016414 | 0.17365 | 0.016065 | 0.17078 | 0.4 | 6 | 180
M2 | B | -6 | 69 | 1.2875 | -1.0456 | 0.065 | -3.5962 | 12.9913 | -1.3125 | 12.3362 | -1.3819 | 1.305 | -1.06 | 0.065 | -3.57 | 12.91 | -1.33 | 1.915 | -1.24 | 0.017989 | 0.16338 | 0.018385 | 0.1626 | 0.4 | 6 | 180
M3 | A | -10 | 63 | 0.78813 | -1.6219 | 3.8994 | -1.9237 | 9.65 | -1.3456 | 1.16 | -1.6194 | 0.805 | -1.64 | 4.005 | -1.89 | 9.75 | -1.365 | 1.085 | -1.535 | 0.010664 | 0.11899 | 0.010845 | 0.12189 | 0.32 | 6 | 160
M3 | B | -10 | 63 | 0.79563 | -1.5981 | 4.3844 | -1.8325 | 9.515 | -1.3513 | 1.2075 | -1.5956 | 0.815 | -1.605 | 4.47 | -1.8 | 9.48 | -1.365 | 1.235 | -1.705 | 0.011179 | 0.11156 | 0.01134 | 0.11368 | 0.32 | 6 | 160
M4 | A | -7 | 66 | 0.645 | -1.6537 | 4.1688 | -1.7394 | 8.2706 | -1.3031 | 0.8475 | -1.4025 | 0.655 | -1.66 | 4.255 | -1.73 | 8.42 | -1.315 | 0.81 | -1.35 | 0.009425 | 0.10983 | 0.009615 | 0.11195 | 0.22 | 6 | 160
M4 | B | -7 | 66 | 0.65438 | -1.6306 | 4.3 | -1.705 | 8.0944 | -1.3106 | 1.0581 | -1.6731 | 0.665 | -1.63 | 4.35 | -1.7 | 8.185 | -1.32 | 0.95 | -1.675 | 0.009915 | 0.10233 | 0.0102 | 0.10498 | 0.22 | 6 | 160
M5 | A | -20 | 54 | 0.605 | -1.6706 | 5.3525 | -1.5144 | 6.1638 | -1.3763 | 0.80187 | -1.3838 | 0.62 | -1.695 | 5.465 | -1.51 | 6.27 | -1.39 | 0.81 | -1.395 | 0.011316 | 0.062433 | 0.01138 | 0.065065 | 0.1 | 6 | 115
M5 | B | -20 | 54 | 0.60563 | -1.6644 | 5.1631 | -1.5194 | 5.9744 | -1.3931 | 0.81937 | -1.3862 | 0.62 | -1.685 | 5.21 | -1.515 | 6.1 | -1.41 | 0.79 | -1.385 | 0.011525 | 0.058254 | 0.01169 | 0.060425 | 0.1 | 6 | 115
M6 | A | -22 | 51 | 0.62375 | -1.6569 | 5.6487 | -1.4263 | 5.6856 | -1.4031 | 0.84313 | -1.3769 | 0.645 | -1.67 | 5.855 | -1.42 | 5.85 | -1.42 | 0.84 | -1.375 | 0.0019056 | 0.013476 | 0.001945 | 0.013955 | 0.096 | 6 | 147
M6 | B | -22 | 51 | 0.63 | -1.6569 | 5.4487 | -1.4344 | 5.5188 | -1.4206 | 0.83562 | -1.3506 | 0.655 | -1.66 | 5.635 | -1.43 | 5.6 | -1.445 | 0.835 | -1.345 | 0.002 | 0.012369 | 0.002035 | 0.01273 | 0.096 | 6 | 147
M7 | A | -16 | 57 | 0.5325 | -1.5875 | 4.7206 | -1.4244 | 4.7906 | -1.3419 | 0.89312 | -1.55 | 0.56 | -1.59 | 4.87 | -1.42 | 4.925 | -1.36 | 0.91 | -1.55 | 0.0065125 | 0.055036 | 0.006745 | 0.0577 | 0.064 | 6 | 124
M7 | B | -16 | 57 | 0.5325 | -1.5837 | 4.5619 | -1.425 | 4.6356 | -1.355 | 0.90062 | -1.5488 | 0.56 | -1.59 | 4.715 | -1.415 | 4.72 | -1.375 | 0.905 | -1.54 | 0.0067981 | 0.050812 | 0.00701 | 0.053025 | 0.064 | 6 | 124
M8 | A | -37 | 36 | 0.64187 | -1.58 | 5.9613 | -1.5463 | 6.8137 | -1.3925 | 0.7325 | -1.4269 | 0.63 | -1.565 | 6.315 | -1.55 | 7.12 | -1.38 | 0.73 | -1.42 | 0.030064 | 0.03247 | 0.030845 | 0.03374 | 0.054 | 6 | 57
M8 | B | -37 | 36 | 0.64438 | -1.57 | 5.8881 | -1.5444 | 6.7612 | -1.4075 | 0.74312 | -1.4338 | 0.63 | -1.58 | 6.19 | -1.55 | 7.12 | -1.4 | 0.745 | -1.44 | 0.030874 | 0.030666 | 0.031485 | 0.03182 | 0.054 | 6 | 57
M9 | A | -43 | 31 | 0.49875 | -1.8738 | 4.9406 | -1.6113 | 4.8894 | -1.6106 | 0.395 | -1.3744 | 0.505 | -1.895 | 5.15 | -1.62 | 5.08 | -1.6 | 0.38 | -1.345 | ~ | ~ | ~ | ~ | ~ | ~ | ~
M9 | B | -43 | 31 | 0.505 | -1.87 | 4.8794 | -1.6188 | 4.8394 | -1.6319 | 0.39563 | -1.3606 | 0.5 | -1.84 | 5.085 | -1.63 | 5.02 | -1.62 | 0.39 | -1.355 | ~ | ~ | ~ | ~ | ~ | ~ | ~
M10 | A | -25 | 49 | 0.27125 | -1.2556 | 4.5906 | -1.5925 | 4.8388 | -1.52 | 4.0281 | -3.5475 | 0.28 | -1.26 | 4.82 | -1.59 | 5.055 | -1.52 | 3.78 | -3.61 | 0.014759 | 0.038132 | 0.01374 | 0.039975 | 0.073 | 6 | 86.1
M10 | B | -25 | 49 | 0.27188 | -1.2544 | 4.5656 | -1.5844 | 4.8256 | -1.4956 | 4.0281 | -3.5312 | 0.28 | -1.27 | 4.735 | -1.585 | 4.985 | -1.495 | 4 | -3.595 | 0.015248 | 0.036022 | 0.014295 | 0.037655 | 0.073 | 6 | 86.1
M11 | A | -31 | 43 | 0.24375 | -0.9025 | 3.4675 | -1.7519 | 3.4281 | -1.7631 | 1.5244 | -2.6287 | 0.235 | -0.885 | 3.63 | -1.735 | 3.515 | -1.76 | 1.415 | -2.65 | 0.00030125 | 0.0003725 | 0.0003 | 0.000385 | 0.0012 | 6 | 1.2
M11 | B | -31 | 43 | 0.24875 | -0.8875 | 3.4656 | -1.74 | 3.405 | -1.8031 | 1.4363 | -2.6888 | 0.235 | -0.975 | 3.62 | -1.715 | 3.505 | -1.8 | 1.28 | -2.555 | 0.00031062 | 0.00035313 | 0.000305 | 0.00037 | 0.0012 | 6 | 1.2
Band | HAM | Detector | Harvy Shack Fit Parameters | [7 empty] | Sensor Performance Compared to Requirement | [5 empty] | Notes:
|  |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Max Allowed Lscat | Angle (mrad) from Target | T Bright Target (K) |  | 1. This table provides VIIRS F2 RSB NFR.
|  |  | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright |  |  |  |  | 2. The results were derived from FP-14 using the ScMA.
I1 | A | 1 | 0.59 | -1.66 | 4.74 | -1.48 | 6.26 | -1.37 | 2.16 | -2.43 | 0.0074 | 0.04671 | 0.22 | 6 | 115 |  | 3. Linear fits were peformed of log(dn/(s*t))= log(A) + M*Log(sin(theta)/0.01))
I1 | A | 2 | 0.53 | -1.32 | 4.37 | -1.51 | 5.69 | -1.4 | 1.42 | -1.91 | 0.00629 | 0.04362 | 0.22 | 6 | 115 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | A | 3 | 0.54 | -1.5 | 4.36 | -1.52 | 5.59 | -1.43 | 1.28 | -1.78 | 0.00655 | 0.04393 | 0.22 | 6 | 115 |  | 5. The test was conducted in diagnostic mode.
I1 | A | 4 | 0.49 | -1.78 | 4.63 | -1.51 | 5.66 | -1.44 | 0.34 | -0.86 | 0.00619 | 0.04521 | 0.22 | 6 | 115 |  | 6. Radiance units are W m-2 sr-1 micron-1
I1 | A | 5 | 0.46 | -1.57 | 4.93 | -1.5 | 5.77 | -1.45 | 0.52 | -1.24 | 0.0056 | 0.04723 | 0.22 | 6 | 115 |  | 7. There is no NFR requirement for band M9
I1 | A | 6 | 0.65 | -1.53 | 5.33 | -1.47 | 6.16 | -1.42 | 2.32 | -2.33 | 0.00791 | 0.05134 | 0.22 | 6 | 115
I1 | A | 7 | 0.63 | -1.22 | 5.57 | -1.46 | 5.84 | -1.49 | 2.19 | -2.15 | 0.00725 | 0.0497 | 0.22 | 6 | 115
I1 | A | 8 | 0.48 | -2.12 | 5.76 | -1.44 | 6.25 | -1.44 | 2.16 | -3.19 | 0.00654 | 0.05336 | 0.22 | 6 | 115
I1 | A | 9 | 0.52 | -1.86 | 5.85 | -1.44 | 6.11 | -1.47 | 0.91 | -1.66 | 0.00667 | 0.05281 | 0.22 | 6 | 115
I1 | A | 10 | 0.58 | -1.43 | 5.88 | -1.43 | 6.12 | -1.45 | 3.09 | -2.86 | 0.00687 | 0.05317 | 0.22 | 6 | 115
I1 | A | 11 | 0.55 | -1.45 | 5.95 | -1.42 | 6.04 | -1.46 | 1.49 | -2.08 | 0.00658 | 0.05274 | 0.22 | 6 | 115
I1 | A | 12 | 0.59 | -1.56 | 5.92 | -1.42 | 6.02 | -1.47 | 1.01 | -1.54 | 0.00714 | 0.05311 | 0.22 | 6 | 115
I1 | A | 13 | 0.54 | -1.82 | 5.86 | -1.43 | 6.09 | -1.45 | 2.03 | -2.63 | 0.00689 | 0.05408 | 0.22 | 6 | 115
I1 | A | 14 | 0.3 | -1.17 | 5.7 | -1.43 | 5.39 | -1.56 | 0.22 | -0.56 | 0.00334 | 0.04893 | 0.22 | 6 | 115
I1 | A | 15 | 0.54 | -1.75 | 5.78 | -1.43 | 5.84 | -1.48 | 0.69 | -1.43 | 0.00677 | 0.0528 | 0.22 | 6 | 115
I1 | A | 16 | 0.69 | -1.51 | 5.82 | -1.42 | 5.98 | -1.47 | 0.92 | -1.32 | 0.00824 | 0.05431 | 0.22 | 6 | 115
I1 | A | 17 | 0.62 | -1.51 | 5.74 | -1.43 | 5.7 | -1.5 | 1.38 | -1.69 | 0.00746 | 0.05245 | 0.22 | 6 | 115
I1 | A | 18 | 0.6 | -1.54 | 5.66 | -1.43 | 5.9 | -1.43 | 2.75 | -2.8 | 0.00725 | 0.05431 | 0.22 | 6 | 115
I1 | A | 19 | 0.7 | -1.48 | 5.5 | -1.44 | 5.54 | -1.5 | 3.07 | -2.5 | 0.0083 | 0.05189 | 0.22 | 6 | 115
I1 | A | 20 | 0.59 | -1.79 | 5.47 | -1.44 | 5.7 | -1.47 | 1.16 | -2.06 | 0.00745 | 0.05358 | 0.22 | 6 | 115
I1 | A | 21 | 0.52 | -1.9 | 5.28 | -1.45 | 5.55 | -1.47 | 3.72 | -3.64 | 0.00664 | 0.05249 | 0.22 | 6 | 115
I1 | A | 22 | 0.54 | -1.79 | 5.23 | -1.45 | 5.56 | -1.47 | 2.41 | -3.23 | 0.00679 | 0.05291 | 0.22 | 6 | 115
I1 | A | 23 | 0.44 | -2.09 | 5.04 | -1.46 | 5.61 | -1.43 | 0.65 | -2.22 | 0.00583 | 0.05346 | 0.22 | 6 | 115
I1 | A | 24 | 0.48 | -1.83 | 4.96 | -1.46 | 5.46 | -1.45 | 1.42 | -2.55 | 0.00605 | 0.05237 | 0.22 | 6 | 115
I1 | A | 25 | 0.5 | -1.82 | 4.83 | -1.47 | 5.46 | -1.43 | 0.52 | -1.38 | 0.00626 | 0.05257 | 0.22 | 6 | 115
I1 | A | 26 | 0.4 | -1.73 | 4.64 | -1.48 | 5.3 | -1.44 | 0.43 | -1.21 | 0.00501 | 0.05157 | 0.22 | 6 | 115
I1 | A | 27 | 0.58 | -1.56 | 4.49 | -1.49 | 5.39 | -1.41 | 2.47 | -2.61 | 0.00702 | 0.0526 | 0.22 | 6 | 115
I1 | A | 28 | 0.59 | -1.42 | 4.36 | -1.5 | 5.36 | -1.4 | 1.54 | -1.94 | 0.00687 | 0.05279 | 0.22 | 6 | 115
I1 | A | 29 | 0.46 | -1.96 | 4.22 | -1.51 | 5.29 | -1.4 | 1.01 | -2.21 | 0.00594 | 0.05244 | 0.22 | 6 | 115
I1 | A | 30 | 0.22 | -0.73 | 3.96 | -1.53 | 4.76 | -1.48 | 0.1 | 0.15 | 0.00227 | 0.04802 | 0.22 | 6 | 115
I1 | A | 31 | 0.56 | -1.35 | 3.98 | -1.51 | 5.38 | -1.35 | 0.94 | -1.51 | 0.00645 | 0.05352 | 0.22 | 6 | 115
I1 | A | 32 | 0.65 | -1.55 | 3.87 | -1.51 | 5.45 | -1.33 | 3.87 | -3.08 | 0.00776 | 0.05426 | 0.22 | 6 | 115
I1 | B | 1 | 0.74 | -1.52 | 5.51 | -1.46 | 7.19 | -1.37 | 1.74 | -1.94 | 0.00904 | 0.04534 | 0.22 | 6 | 115
I1 | B | 2 | 0.62 | -1.42 | 5.05 | -1.49 | 6.39 | -1.42 | 1.92 | -1.98 | 0.0074 | 0.04164 | 0.22 | 6 | 115
I1 | B | 3 | 0.65 | -1.49 | 4.97 | -1.51 | 6.3 | -1.43 | 2 | -2.21 | 0.00786 | 0.04221 | 0.22 | 6 | 115
I1 | B | 4 | 0.54 | -1.99 | 5.19 | -1.5 | 6.35 | -1.44 | 0.81 | -1.62 | 0.00724 | 0.04345 | 0.22 | 6 | 115
I1 | B | 5 | 0.52 | -1.67 | 5.52 | -1.49 | 6.45 | -1.46 | 1.01 | -1.81 | 0.00654 | 0.04532 | 0.22 | 6 | 115
I1 | B | 6 | 0.76 | -1.58 | 5.95 | -1.46 | 6.86 | -1.44 | 2.3 | -2.13 | 0.00944 | 0.04934 | 0.22 | 6 | 115
I1 | B | 7 | 0.71 | -1.17 | 6.19 | -1.45 | 6.46 | -1.51 | 2.15 | -1.91 | 0.0082 | 0.04773 | 0.22 | 6 | 115
I1 | B | 8 | 0.54 | -2.21 | 6.39 | -1.44 | 6.86 | -1.47 | 1.88 | -2.96 | 0.00748 | 0.05103 | 0.22 | 6 | 115
I1 | B | 9 | 0.62 | -1.73 | 6.46 | -1.43 | 6.76 | -1.48 | 2.28 | -2.58 | 0.00788 | 0.05094 | 0.22 | 6 | 115
I1 | B | 10 | 0.66 | -1.39 | 6.45 | -1.43 | 6.69 | -1.48 | 3.27 | -2.96 | 0.00787 | 0.05095 | 0.22 | 6 | 115
I1 | B | 11 | 0.64 | -1.39 | 6.48 | -1.42 | 6.62 | -1.47 | 0.67 | -0.94 | 0.00765 | 0.0506 | 0.22 | 6 | 115
I1 | B | 12 | 0.68 | -1.54 | 6.45 | -1.42 | 6.54 | -1.49 | 1.94 | -2.09 | 0.00825 | 0.05072 | 0.22 | 6 | 115
I1 | B | 13 | 0.63 | -1.75 | 6.35 | -1.43 | 6.63 | -1.46 | 2.89 | -3.01 | 0.00792 | 0.0519 | 0.22 | 6 | 115
I1 | B | 14 | 0.34 | -1 | 6.11 | -1.44 | 5.79 | -1.58 | 0.22 | -0.38 | 0.0038 | 0.04657 | 0.22 | 6 | 115
I1 | B | 15 | 0.61 | -1.63 | 6.25 | -1.43 | 6.34 | -1.49 | 1.05 | -1.77 | 0.00753 | 0.05068 | 0.22 | 6 | 115
I1 | B | 16 | 0.77 | -1.49 | 6.28 | -1.42 | 6.45 | -1.47 | 1.82 | -2.03 | 0.00931 | 0.05194 | 0.22 | 6 | 115
I1 | B | 17 | 0.69 | -1.33 | 6.17 | -1.43 | 6.13 | -1.51 | 1.64 | -1.9 | 0.00811 | 0.0502 | 0.22 | 6 | 115
I1 | B | 18 | 0.65 | -1.66 | 6.04 | -1.43 | 6.32 | -1.45 | 0.9 | -1.43 | 0.00809 | 0.05192 | 0.22 | 6 | 115
I1 | B | 19 | 0.77 | -1.42 | 5.94 | -1.43 | 5.94 | -1.51 | 5.14 | -3.01 | 0.00917 | 0.04977 | 0.22 | 6 | 115
I1 | B | 20 | 0.65 | -1.75 | 5.85 | -1.43 | 6.08 | -1.48 | 1.58 | -2.26 | 0.00822 | 0.05117 | 0.22 | 6 | 115
I1 | B | 21 | 0.56 | -1.88 | 5.62 | -1.45 | 5.91 | -1.48 | 1.17 | -2.24 | 0.00718 | 0.05022 | 0.22 | 6 | 115
I1 | B | 22 | 0.61 | -1.84 | 5.55 | -1.45 | 5.93 | -1.49 | 0.72 | -1.56 | 0.00786 | 0.05076 | 0.22 | 6 | 115
I1 | B | 23 | 0.46 | -2.18 | 5.42 | -1.45 | 5.94 | -1.46 | 0.36 | -1.22 | 0.00636 | 0.05103 | 0.22 | 6 | 115
I1 | B | 24 | 0.54 | -1.77 | 5.29 | -1.46 | 5.77 | -1.47 | 0.72 | -1.69 | 0.00687 | 0.04991 | 0.22 | 6 | 115
I1 | B | 25 | 0.58 | -1.76 | 5.18 | -1.46 | 5.79 | -1.45 | 0.98 | -1.93 | 0.00725 | 0.0504 | 0.22 | 6 | 115
I1 | B | 26 | 0.47 | -1.72 | 4.95 | -1.48 | 5.65 | -1.45 | 0.45 | -1.23 | 0.00593 | 0.04964 | 0.22 | 6 | 115
I1 | B | 27 | 0.65 | -1.52 | 4.81 | -1.48 | 5.73 | -1.43 | 1.43 | -1.83 | 0.00782 | 0.05071 | 0.22 | 6 | 115
I1 | B | 28 | 0.53 | -1.56 | 4.67 | -1.49 | 5.6 | -1.44 | 0.73 | -1.37 | 0.00638 | 0.05028 | 0.22 | 6 | 115
I1 | B | 29 | 0.51 | -1.87 | 4.56 | -1.49 | 5.59 | -1.42 | 0.34 | -0.98 | 0.00658 | 0.05059 | 0.22 | 6 | 115
I1 | B | 30 | 0.27 | -0.79 | 4.26 | -1.52 | 4.99 | -1.51 | 0.27 | -0.69 | 0.00291 | 0.04602 | 0.22 | 6 | 115
I1 | B | 31 | 0.6 | -1.23 | 4.31 | -1.5 | 5.67 | -1.38 | 0.78 | -1.17 | 0.00684 | 0.05151 | 0.22 | 6 | 115
I1 | B | 32 | 0.71 | -1.5 | 4.23 | -1.49 | 5.78 | -1.35 | 1.99 | -2.15 | 0.00855 | 0.05269 | 0.22 | 6 | 115
I2 | A | 1 | 0.55 | -1.54 | 3.62 | -1.56 | 5.31 | -1.31 | 6.62 | -3.3 | 0.00716 | 0.04181 | 0.25 | 6 | 124
I2 | A | 2 | 0.54 | -1.52 | 3.38 | -1.58 | 4.96 | -1.34 | 3.76 | -2.72 | 0.007 | 0.04005 | 0.25 | 6 | 124
I2 | A | 3 | 0.55 | -1.23 | 3.37 | -1.6 | 4.84 | -1.37 | 2.29 | -2.03 | 0.00678 | 0.04011 | 0.25 | 6 | 124
I2 | A | 4 | 0.46 | -1.75 | 3.6 | -1.59 | 4.94 | -1.39 | 4.61 | -3.32 | 0.00614 | 0.04185 | 0.25 | 6 | 124
I2 | A | 5 | 0.53 | -1.5 | 3.9 | -1.57 | 5.14 | -1.4 | 4.79 | -3.14 | 0.00677 | 0.04457 | 0.25 | 6 | 124
I2 | A | 6 | 0.6 | -1.3 | 4.25 | -1.53 | 5.31 | -1.4 | 2.5 | -2.18 | 0.00738 | 0.04716 | 0.25 | 6 | 124
I2 | A | 7 | 0.56 | -1.63 | 4.57 | -1.51 | 5.5 | -1.41 | 5.06 | -3.2 | 0.00735 | 0.04957 | 0.25 | 6 | 124
I2 | A | 8 | 0.6 | -1.64 | 4.82 | -1.49 | 5.68 | -1.39 | 4.72 | -3.07 | 0.00794 | 0.05164 | 0.25 | 6 | 124
I2 | A | 9 | 0.62 | -1.59 | 4.94 | -1.48 | 5.74 | -1.4 | 5.13 | -3.12 | 0.00809 | 0.05226 | 0.25 | 6 | 124
I2 | A | 10 | 0.64 | -1.61 | 4.95 | -1.47 | 5.64 | -1.4 | 5.78 | -3.25 | 0.00832 | 0.05175 | 0.25 | 6 | 124
I2 | A | 11 | 0.58 | -1.67 | 4.93 | -1.47 | 5.56 | -1.42 | 4.41 | -3.05 | 0.00768 | 0.05142 | 0.25 | 6 | 124
I2 | A | 12 | 0.62 | -1.85 | 4.9 | -1.47 | 5.57 | -1.4 | 6.41 | -3.45 | 0.00847 | 0.05176 | 0.25 | 6 | 124
I2 | A | 13 | 0.56 | -1.69 | 4.83 | -1.48 | 5.43 | -1.42 | 4.02 | -2.95 | 0.00735 | 0.05099 | 0.25 | 6 | 124
I2 | A | 14 | 0.58 | -1.63 | 4.8 | -1.48 | 5.39 | -1.42 | 3.33 | -2.67 | 0.00757 | 0.05095 | 0.25 | 6 | 124
I2 | A | 15 | 0.58 | -1.59 | 4.85 | -1.48 | 5.42 | -1.42 | 5.31 | -3.3 | 0.00748 | 0.05162 | 0.25 | 6 | 124
I2 | A | 16 | 0.62 | -1.72 | 4.88 | -1.48 | 5.48 | -1.42 | 6.09 | -3.38 | 0.00813 | 0.05247 | 0.25 | 6 | 124
I2 | A | 17 | 0.57 | -1.61 | 4.83 | -1.48 | 5.41 | -1.42 | 3.68 | -2.84 | 0.00736 | 0.05233 | 0.25 | 6 | 124
I2 | A | 18 | 0.6 | -1.71 | 4.75 | -1.48 | 5.38 | -1.41 | 4.79 | -3.11 | 0.00789 | 0.05253 | 0.25 | 6 | 124
I2 | A | 19 | 0.58 | -1.58 | 4.65 | -1.49 | 5.38 | -1.4 | 4.55 | -3.15 | 0.00744 | 0.05287 | 0.25 | 6 | 124
I2 | A | 20 | 0.62 | -1.78 | 4.54 | -1.5 | 5.32 | -1.4 | 5.33 | -3.22 | 0.00819 | 0.05276 | 0.25 | 6 | 124
I2 | A | 21 | 0.59 | -1.78 | 4.44 | -1.5 | 5.25 | -1.4 | 8.25 | -3.7 | 0.00789 | 0.05241 | 0.25 | 6 | 124
I2 | A | 22 | 0.61 | -1.79 | 4.36 | -1.51 | 5.2 | -1.39 | 4.31 | -3 | 0.00809 | 0.05231 | 0.25 | 6 | 124
I2 | A | 23 | 0.62 | -1.55 | 4.28 | -1.51 | 5.16 | -1.39 | 3.51 | -2.58 | 0.00789 | 0.05212 | 0.25 | 6 | 124
I2 | A | 24 | 0.61 | -1.67 | 4.16 | -1.51 | 5.06 | -1.39 | 4.29 | -2.83 | 0.0079 | 0.05158 | 0.25 | 6 | 124
I2 | A | 25 | 0.64 | -1.45 | 4.06 | -1.52 | 5.07 | -1.38 | 3.13 | -2.45 | 0.00796 | 0.0519 | 0.25 | 6 | 124
I2 | A | 26 | 0.55 | -1.69 | 3.91 | -1.54 | 4.98 | -1.38 | 3.9 | -2.99 | 0.00714 | 0.05137 | 0.25 | 6 | 124
I2 | A | 27 | 0.59 | -1.44 | 3.76 | -1.55 | 5.04 | -1.34 | 3.45 | -2.62 | 0.00733 | 0.05231 | 0.25 | 6 | 124
I2 | A | 28 | 0.53 | -1.63 | 3.63 | -1.56 | 4.86 | -1.36 | 4.06 | -3.06 | 0.00683 | 0.05092 | 0.25 | 6 | 124
I2 | A | 29 | 0.52 | -1.9 | 3.51 | -1.57 | 4.85 | -1.35 | 5.48 | -3.54 | 0.00704 | 0.05109 | 0.25 | 6 | 124
I2 | A | 30 | 0.56 | -1.63 | 3.43 | -1.57 | 4.95 | -1.31 | 3.61 | -2.87 | 0.00717 | 0.05222 | 0.25 | 6 | 124
I2 | A | 31 | 0.52 | -1.63 | 3.32 | -1.57 | 4.89 | -1.31 | 3.92 | -3.08 | 0.00672 | 0.0516 | 0.25 | 6 | 124
I2 | A | 32 | 0.58 | -1.64 | 3.23 | -1.57 | 4.98 | -1.27 | 5.15 | -3.19 | 0.00749 | 0.05238 | 0.25 | 6 | 124
I2 | B | 1 | 0.67 | -1.61 | 4.16 | -1.54 | 6.07 | -1.32 | 7.14 | -3.26 | 0.00884 | 0.04048 | 0.25 | 6 | 124
I2 | B | 2 | 0.64 | -1.45 | 3.84 | -1.58 | 5.62 | -1.35 | 3.77 | -2.58 | 0.00827 | 0.03846 | 0.25 | 6 | 124
I2 | B | 3 | 0.64 | -1.14 | 3.79 | -1.59 | 5.43 | -1.38 | 2.34 | -1.86 | 0.00774 | 0.03835 | 0.25 | 6 | 124
I2 | B | 4 | 0.52 | -1.61 | 4.02 | -1.58 | 5.5 | -1.41 | 5.29 | -3.31 | 0.00689 | 0.03994 | 0.25 | 6 | 124
I2 | B | 5 | 0.59 | -1.53 | 4.35 | -1.56 | 5.75 | -1.41 | 4.42 | -2.91 | 0.0077 | 0.04279 | 0.25 | 6 | 124
I2 | B | 6 | 0.67 | -1.26 | 4.72 | -1.53 | 5.9 | -1.42 | 2.74 | -2.1 | 0.0083 | 0.04514 | 0.25 | 6 | 124
I2 | B | 7 | 0.65 | -1.58 | 5.05 | -1.51 | 6.11 | -1.42 | 4.5 | -2.84 | 0.00848 | 0.04764 | 0.25 | 6 | 124
I2 | B | 8 | 0.7 | -1.55 | 5.29 | -1.49 | 6.29 | -1.41 | 5.03 | -2.95 | 0.00911 | 0.04969 | 0.25 | 6 | 124
I2 | B | 9 | 0.71 | -1.56 | 5.42 | -1.48 | 6.34 | -1.41 | 5.37 | -3.02 | 0.00924 | 0.05028 | 0.25 | 6 | 124
I2 | B | 10 | 0.73 | -1.65 | 5.41 | -1.47 | 6.2 | -1.42 | 5.03 | -2.93 | 0.0097 | 0.04966 | 0.25 | 6 | 124
I2 | B | 11 | 0.65 | -1.61 | 5.39 | -1.47 | 6.09 | -1.43 | 5.85 | -3.21 | 0.00855 | 0.04924 | 0.25 | 6 | 124
I2 | B | 12 | 0.69 | -1.83 | 5.35 | -1.47 | 6.07 | -1.42 | 11.55 | -3.95 | 0.00951 | 0.04952 | 0.25 | 6 | 124
I2 | B | 13 | 0.64 | -1.69 | 5.24 | -1.48 | 5.9 | -1.44 | 3.17 | -2.49 | 0.00857 | 0.04879 | 0.25 | 6 | 124
I2 | B | 14 | 0.66 | -1.61 | 5.22 | -1.48 | 5.83 | -1.44 | 4.83 | -3.01 | 0.00865 | 0.04868 | 0.25 | 6 | 124
I2 | B | 15 | 0.61 | -1.86 | 5.22 | -1.48 | 5.84 | -1.44 | 6.79 | -3.53 | 0.0084 | 0.04917 | 0.25 | 6 | 124
I2 | B | 16 | 0.69 | -1.7 | 5.27 | -1.47 | 5.92 | -1.43 | 6.25 | -3.24 | 0.00917 | 0.05015 | 0.25 | 6 | 124
I2 | B | 17 | 0.63 | -1.61 | 5.17 | -1.48 | 5.82 | -1.44 | 4.99 | -3.09 | 0.00829 | 0.05 | 0.25 | 6 | 124
I2 | B | 18 | 0.67 | -1.8 | 5.09 | -1.48 | 5.77 | -1.43 | 5.32 | -3.12 | 0.00902 | 0.05014 | 0.25 | 6 | 124
I2 | B | 19 | 0.63 | -1.64 | 4.97 | -1.49 | 5.77 | -1.41 | 4.51 | -3.02 | 0.00824 | 0.05056 | 0.25 | 6 | 124
I2 | B | 20 | 0.68 | -1.75 | 4.85 | -1.49 | 5.69 | -1.42 | 5.34 | -3.11 | 0.00905 | 0.05049 | 0.25 | 6 | 124
I2 | B | 21 | 0.67 | -1.74 | 4.71 | -1.5 | 5.59 | -1.42 | 6.75 | -3.32 | 0.00895 | 0.05008 | 0.25 | 6 | 124
I2 | B | 22 | 0.67 | -1.73 | 4.65 | -1.5 | 5.53 | -1.41 | 6.11 | -3.32 | 0.00892 | 0.05003 | 0.25 | 6 | 124
I2 | B | 23 | 0.69 | -1.49 | 4.55 | -1.5 | 5.49 | -1.41 | 4.93 | -2.87 | 0.00882 | 0.04994 | 0.25 | 6 | 124
I2 | B | 24 | 0.69 | -1.55 | 4.42 | -1.51 | 5.39 | -1.41 | 8 | -3.42 | 0.00885 | 0.04939 | 0.25 | 6 | 124
I2 | B | 25 | 0.72 | -1.41 | 4.31 | -1.52 | 5.39 | -1.4 | 3.45 | -2.41 | 0.00903 | 0.04971 | 0.25 | 6 | 124
I2 | B | 26 | 0.6 | -1.61 | 4.15 | -1.53 | 5.28 | -1.4 | 5.02 | -3.13 | 0.00781 | 0.04927 | 0.25 | 6 | 124
I2 | B | 27 | 0.65 | -1.42 | 4.01 | -1.54 | 5.34 | -1.36 | 3.29 | -2.48 | 0.00817 | 0.05036 | 0.25 | 6 | 124
I2 | B | 28 | 0.58 | -1.63 | 3.87 | -1.55 | 5.15 | -1.38 | 3.52 | -2.75 | 0.00761 | 0.04905 | 0.25 | 6 | 124
I2 | B | 29 | 0.55 | -1.89 | 3.74 | -1.56 | 5.14 | -1.37 | 5.85 | -3.46 | 0.0075 | 0.04928 | 0.25 | 6 | 124
I2 | B | 30 | 0.61 | -1.65 | 3.69 | -1.56 | 5.25 | -1.33 | 4.15 | -2.93 | 0.00797 | 0.05045 | 0.25 | 6 | 124
I2 | B | 31 | 0.57 | -1.64 | 3.56 | -1.56 | 5.18 | -1.33 | 3.8 | -2.92 | 0.00743 | 0.04989 | 0.25 | 6 | 124
I2 | B | 32 | 0.65 | -1.58 | 3.51 | -1.55 | 5.28 | -1.29 | 5.66 | -3.21 | 0.00846 | 0.05066 | 0.25 | 6 | 124
I3 | A | 1 | 0.17 | -1.18 | 2.57 | -1.53 | 2.92 | -1.22 | 9.14 | -4.96 | 0.00166 | 0.01435 | 0.073 | 6 | 86.1
I3 | A | 2 | 0.09 | -1.22 | 1.31 | -1.69 | 1.51 | -1.29 | 2.93 | -4.44 | 0.0019 | 0.01429 | 0.073 | 6 | 86.1
I3 | A | 3 | 0.16 | -1.27 | 1.97 | -1.75 | 2.73 | -1.51 | 7.1 | -4.69 | 0.00161 | 0.01476 | 0.073 | 6 | 86.1
I3 | A | 4 | 0.17 | -0.28 | 1.59 | -1.58 | 1.58 | -1.58 | 0.44 | -1.48 | 0.00246 | 0.01583 | 0.073 | 6 | 86.1
I3 | A | 5 | 0.17 | -1.19 | 2.75 | -1.54 | 2.98 | -1.31 | 7.23 | -4.61 | 0.00169 | 0.01629 | 0.073 | 6 | 86.1
I3 | A | 6 | 0.11 | -1.24 | 1.81 | -1.61 | 1.91 | -1.34 | 3.25 | -4.25 | 0.00243 | 0.01832 | 0.073 | 6 | 86.1
I3 | A | 7 | 0.2 | -1.3 | 3.34 | -1.47 | 3.36 | -1.37 | 9.46 | -4.75 | 0.00195 | 0.01937 | 0.073 | 6 | 86.1
I3 | A | 8 | 0.13 | -1.36 | 2.15 | -1.56 | 2.18 | -1.46 | 4.65 | -4.54 | 0.00277 | 0.02131 | 0.073 | 6 | 86.1
I3 | A | 9 | 0.2 | -1.26 | 3.31 | -1.53 | 3.55 | -1.52 | 10.24 | -4.85 | 0.00193 | 0.02135 | 0.073 | 6 | 86.1
I3 | A | 10 | 0.13 | -1.29 | 2.25 | -1.57 | 2.31 | -1.52 | 5.49 | -4.72 | 0.00286 | 0.02275 | 0.073 | 6 | 86.1
I3 | A | 11 | 0.19 | -1.23 | 3.48 | -1.43 | 3.41 | -1.4 | 10.65 | -4.9 | 0.00189 | 0.02068 | 0.073 | 6 | 86.1
I3 | A | 12 | 0.13 | -1.18 | 2.29 | -1.53 | 2.25 | -1.41 | 4.82 | -4.51 | 0.00283 | 0.02182 | 0.073 | 6 | 86.1
I3 | A | 13 | 0.18 | -1.34 | 3.3 | -1.45 | 3.32 | -1.51 | 8.67 | -4.74 | 0.00175 | 0.02105 | 0.073 | 6 | 86.1
I3 | A | 14 | 0.13 | -1.33 | 2.23 | -1.58 | 2.29 | -1.39 | 4.7 | -4.4 | 0.00282 | 0.02219 | 0.073 | 6 | 86.1
I3 | A | 15 | 0.17 | -1.31 | 3.27 | -1.43 | 3.29 | -1.52 | 10.55 | -4.97 | 0.00174 | 0.02139 | 0.073 | 6 | 86.1
I3 | A | 16 | 0.14 | -1.37 | 2.26 | -1.59 | 2.37 | -1.41 | 7.01 | -4.92 | 0.00286 | 0.02298 | 0.073 | 6 | 86.1
I3 | A | 17 | 0.18 | -1.31 | 3.2 | -1.46 | 3.25 | -1.42 | 14.85 | -5.35 | 0.0018 | 0.02129 | 0.073 | 6 | 86.1
I3 | A | 18 | 0.14 | -1.33 | 2.37 | -1.55 | 2.43 | -1.38 | 5.12 | -4.47 | 0.00295 | 0.02347 | 0.073 | 6 | 86.1
I3 | A | 19 | 0.18 | -1.4 | 3.12 | -1.46 | 3.26 | -1.51 | 10.04 | -4.87 | 0.0018 | 0.02204 | 0.073 | 6 | 86.1
I3 | A | 20 | 0.14 | -1.34 | 2.3 | -1.59 | 2.42 | -1.37 | 6.38 | -4.81 | 0.00289 | 0.0234 | 0.073 | 6 | 86.1
I3 | A | 21 | 0.18 | -1.38 | 2.98 | -1.48 | 3.13 | -1.38 | 8.16 | -4.67 | 0.00179 | 0.02127 | 0.073 | 6 | 86.1
I3 | A | 22 | 0.14 | -1.32 | 2.23 | -1.61 | 2.4 | -1.37 | 4.8 | -4.31 | 0.00284 | 0.02321 | 0.073 | 6 | 86.1
I3 | A | 23 | 0.18 | -1.33 | 2.87 | -1.48 | 3.06 | -1.34 | 8.16 | -4.71 | 0.00175 | 0.0211 | 0.073 | 6 | 86.1
I3 | A | 24 | 0.14 | -1.27 | 2.13 | -1.66 | 2.41 | -1.37 | 4.91 | -4.41 | 0.00277 | 0.02323 | 0.073 | 6 | 86.1
I3 | A | 25 | 0.17 | -1.3 | 2.78 | -1.5 | 3.11 | -1.55 | 7.88 | -4.65 | 0.00166 | 0.02228 | 0.073 | 6 | 86.1
I3 | A | 26 | 0.15 | -1.35 | 2.16 | -1.62 | 2.42 | -1.31 | 5.6 | -4.6 | 0.00279 | 0.02314 | 0.073 | 6 | 86.1
I3 | A | 27 | 0.18 | -1.33 | 2.7 | -1.5 | 3.04 | -1.3 | 9.17 | -4.95 | 0.00174 | 0.02141 | 0.073 | 6 | 86.1
I3 | A | 28 | 0.14 | -1.24 | 1.96 | -1.71 | 2.43 | -1.26 | 4.88 | -4.42 | 0.00268 | 0.02317 | 0.073 | 6 | 86.1
I3 | A | 29 | 0.17 | -1.2 | 2.62 | -1.51 | 3.16 | -1.47 | 7.73 | -4.7 | 0.00163 | 0.0231 | 0.073 | 6 | 86.1
I3 | A | 30 | 0.15 | -1.31 | 2.07 | -1.62 | 2.45 | -1.25 | 4.55 | -4.29 | 0.00277 | 0.02326 | 0.073 | 6 | 86.1
I3 | A | 31 | 0.18 | -1.18 | 2.42 | -1.56 | 3.08 | -1.2 | 6.76 | -4.47 | 0.00171 | 0.02197 | 0.073 | 6 | 86.1
I3 | A | 32 | 0.15 | -1.2 | 2.03 | -1.59 | 2.51 | -1.17 | 3.62 | -4.03 | 0.00272 | 0.02354 | 0.073 | 6 | 86.1
I3 | B | 1 | 0.15 | -1.23 | 2.19 | -1.55 | 2.55 | -1.25 | 8.51 | -5.01 | 0.00202 | 0.01407 | 0.073 | 6 | 86.1
I3 | B | 2 | 0.09 | -1.23 | 1.29 | -1.69 | 1.5 | -1.27 | 3.19 | -4.54 | 0.00195 | 0.01379 | 0.073 | 6 | 86.1
I3 | B | 3 | 0.15 | -1.27 | 1.73 | -1.81 | 2.54 | -1.28 | 9.11 | -5.1 | 0.00185 | 0.01356 | 0.073 | 6 | 86.1
I3 | B | 4 | 0.22 | -0.72 | 1.31 | -1.74 | 1.58 | -1.69 | 0.47 | -1.74 | 0.00293 | 0.01572 | 0.073 | 6 | 86.1
I3 | B | 5 | 0.18 | -1.27 | 2.79 | -1.54 | 3.04 | -1.36 | 9.74 | -4.95 | 0.002 | 0.01578 | 0.073 | 6 | 86.1
I3 | B | 6 | 0.11 | -1.24 | 1.6 | -1.71 | 1.86 | -1.35 | 3.63 | -4.35 | 0.00239 | 0.01758 | 0.073 | 6 | 86.1
I3 | B | 7 | 0.22 | -1.24 | 3.72 | -1.46 | 3.75 | -1.38 | 10.58 | -4.8 | 0.00231 | 0.0188 | 0.073 | 6 | 86.1
I3 | B | 8 | 0.12 | -1.24 | 2 | -1.6 | 2.09 | -1.37 | 4.22 | -4.43 | 0.00274 | 0.01991 | 0.073 | 6 | 86.1
I3 | B | 9 | 0.22 | -1.3 | 4.09 | -1.46 | 4.13 | -1.49 | 13.12 | -5.02 | 0.00221 | 0.02073 | 0.073 | 6 | 86.1
I3 | B | 10 | 0.13 | -1.31 | 2.03 | -1.63 | 2.18 | -1.39 | 4.95 | -4.63 | 0.00282 | 0.02076 | 0.073 | 6 | 86.1
I3 | B | 11 | 0.22 | -1.24 | 3.73 | -1.51 | 3.99 | -1.52 | 13.59 | -4.96 | 0.00222 | 0.02019 | 0.073 | 6 | 86.1
I3 | B | 12 | 0.13 | -1.27 | 2.05 | -1.61 | 2.16 | -1.41 | 4.28 | -4.35 | 0.00283 | 0.02073 | 0.073 | 6 | 86.1
I3 | B | 13 | 0.21 | -1.31 | 3.84 | -1.42 | 3.75 | -1.42 | 14.14 | -5.19 | 0.00207 | 0.01989 | 0.073 | 6 | 86.1
I3 | B | 14 | 0.13 | -1.29 | 2.2 | -1.53 | 2.18 | -1.42 | 5.26 | -4.63 | 0.00284 | 0.0209 | 0.073 | 6 | 86.1
I3 | B | 15 | 0.21 | -1.32 | 3.45 | -1.51 | 3.76 | -1.57 | 12.15 | -5 | 0.00206 | 0.02107 | 0.073 | 6 | 86.1
I3 | B | 16 | 0.13 | -1.27 | 2.26 | -1.53 | 2.24 | -1.44 | 6.16 | -4.8 | 0.00287 | 0.02159 | 0.073 | 6 | 86.1
I3 | B | 17 | 0.21 | -1.25 | 3.6 | -1.46 | 3.78 | -1.57 | 11.42 | -4.83 | 0.00205 | 0.02136 | 0.073 | 6 | 86.1
I3 | B | 18 | 0.13 | -1.35 | 2.25 | -1.55 | 2.29 | -1.41 | 5 | -4.52 | 0.0029 | 0.02202 | 0.073 | 6 | 86.1
I3 | B | 19 | 0.21 | -1.32 | 3.51 | -1.46 | 3.68 | -1.51 | 13.58 | -5.07 | 0.00208 | 0.02135 | 0.073 | 6 | 86.1
I3 | B | 20 | 0.14 | -1.31 | 2.16 | -1.58 | 2.27 | -1.41 | 4.97 | -4.52 | 0.00286 | 0.0219 | 0.073 | 6 | 86.1
I3 | B | 21 | 0.2 | -1.42 | 3.38 | -1.47 | 3.56 | -1.36 | 11.47 | -4.89 | 0.00206 | 0.02083 | 0.073 | 6 | 86.1
I3 | B | 22 | 0.14 | -1.38 | 2.05 | -1.63 | 2.28 | -1.36 | 4.84 | -4.49 | 0.00282 | 0.02183 | 0.073 | 6 | 86.1
I3 | B | 23 | 0.2 | -1.3 | 2.99 | -1.57 | 3.47 | -1.37 | 9.21 | -4.64 | 0.00199 | 0.02065 | 0.073 | 6 | 86.1
I3 | B | 24 | 0.13 | -1.28 | 2.08 | -1.6 | 2.27 | -1.36 | 4.4 | -4.37 | 0.00275 | 0.02179 | 0.073 | 6 | 86.1
I3 | B | 25 | 0.19 | -1.35 | 3.14 | -1.5 | 3.51 | -1.46 | 8.72 | -4.63 | 0.0019 | 0.02073 | 0.073 | 6 | 86.1
I3 | B | 26 | 0.13 | -1.27 | 2 | -1.63 | 2.26 | -1.34 | 4.09 | -4.29 | 0.00273 | 0.02168 | 0.073 | 6 | 86.1
I3 | B | 27 | 0.2 | -1.33 | 3.04 | -1.51 | 3.45 | -1.31 | 7.97 | -4.5 | 0.00199 | 0.02101 | 0.073 | 6 | 86.1
I3 | B | 28 | 0.14 | -1.27 | 1.97 | -1.62 | 2.27 | -1.3 | 4.44 | -4.39 | 0.00274 | 0.02168 | 0.073 | 6 | 86.1
I3 | B | 29 | 0.19 | -1.17 | 2.95 | -1.52 | 3.47 | -1.27 | 8.59 | -4.62 | 0.00188 | 0.02132 | 0.073 | 6 | 86.1
I3 | B | 30 | 0.14 | -1.2 | 1.94 | -1.61 | 2.32 | -1.22 | 3.96 | -4.15 | 0.00269 | 0.02185 | 0.073 | 6 | 86.1
I3 | B | 31 | 0.2 | -1.24 | 2.8 | -1.53 | 3.57 | -1.37 | 8.9 | -4.65 | 0.00202 | 0.02166 | 0.073 | 6 | 86.1
I3 | B | 32 | 0.14 | -1.2 | 1.87 | -1.61 | 2.32 | -1.21 | 4.48 | -4.36 | 0.00263 | 0.02187 | 0.073 | 6 | 86.1
M1 | A | 1 | 1.11 | -1.36 | 0 | -4.88 | 17.51 | -1.17 | 32.53 | -4.45 | 0.0137 | 0.18325 | 0.449 | 6 | 162
M1 | A | 2 | 0.98 | -1.43 | 0 | -5.12 | 13.89 | -1.21 | 7.29 | -3.18 | 0.01211 | 0.14739 | 0.449 | 6 | 162
M1 | A | 3 | 0.96 | -1.43 | 0 | -5.14 | 13.21 | -1.26 | 26.58 | -4.82 | 0.01175 | 0.14206 | 0.449 | 6 | 162
M1 | A | 4 | 1.17 | -1.49 | 0 | -5.1 | 15.01 | -1.23 | 8.56 | -3 | 0.01442 | 0.16215 | 0.449 | 6 | 162
M1 | A | 5 | 1.11 | -1.56 | 0 | -5.1 | 15.87 | -1.25 | 2.76 | -2.28 | 0.01374 | 0.17231 | 0.449 | 6 | 162
M1 | A | 6 | 1.13 | -1.44 | 0 | -5.22 | 14.89 | -1.33 | 0.81 | -0.74 | 0.01356 | 0.16404 | 0.449 | 6 | 162
M1 | A | 7 | 1.04 | -1.58 | 0 | -5.29 | 15.31 | -1.32 | 0.25 | 0.45 | 0.01269 | 0.16939 | 0.449 | 6 | 162
M1 | A | 8 | 1.11 | -1.57 | 0 | -5.37 | 15.45 | -1.31 | 1.01 | -0.76 | 0.01349 | 0.17195 | 0.449 | 6 | 162
M1 | A | 9 | 1.12 | -1.54 | 0 | -5.43 | 15.56 | -1.32 | 31.88 | -4.6 | 0.01348 | 0.17408 | 0.449 | 6 | 162
M1 | A | 10 | 1.2 | -1.47 | 0 | -5.45 | 15.73 | -1.31 | 6.89 | -3.21 | 0.01426 | 0.17636 | 0.449 | 6 | 162
M1 | A | 11 | 1.13 | -1.57 | 0 | -5.46 | 16.16 | -1.27 | 10.72 | -3.62 | 0.01358 | 0.18138 | 0.449 | 6 | 162
M1 | A | 12 | 1.14 | -1.56 | 0 | -5.48 | 15.52 | -1.27 | 1.83 | -1.54 | 0.01369 | 0.17536 | 0.449 | 6 | 162
M1 | A | 13 | 1.08 | -1.62 | 0 | -5.53 | 14.97 | -1.31 | 5.3 | -2.5 | 0.01307 | 0.17108 | 0.449 | 6 | 162
M1 | A | 14 | 1.2 | -1.5 | 0 | -5.5 | 15.47 | -1.25 | 3.79 | -2.18 | 0.01421 | 0.17624 | 0.449 | 6 | 162
M1 | A | 15 | 1.25 | -1.4 | 0 | -5.48 | 15.86 | -1.2 | 0.54 | -0.13 | 0.01454 | 0.18115 | 0.449 | 6 | 162
M1 | A | 16 | 1.17 | -1.46 | 0 | -5.46 | 16.14 | -1.18 | 4.34 | -2.27 | 0.01377 | 0.18462 | 0.449 | 6 | 162
M1 | B | 1 | 1.14 | -1.4 | 0 | -4.48 | 17.88 | -1.18 | 1.1 | -1.04 | 0.01466 | 0.17636 | 0.449 | 6 | 162
M1 | B | 2 | 1.04 | -1.41 | 0 | -4.73 | 14.04 | -1.21 | 15.95 | -4.09 | 0.01327 | 0.14072 | 0.449 | 6 | 162
M1 | B | 3 | 0.94 | -1.42 | 0 | -4.73 | 13.28 | -1.26 | 0.28 | 0.21 | 0.01194 | 0.13492 | 0.449 | 6 | 162
M1 | B | 4 | 1.21 | -1.39 | 0 | -4.74 | 14.99 | -1.23 | 6.17 | -2.88 | 0.0152 | 0.15352 | 0.449 | 6 | 162
M1 | B | 5 | 1.11 | -1.49 | 0 | -4.74 | 15.73 | -1.26 | 19.53 | -4.25 | 0.01411 | 0.16208 | 0.449 | 6 | 162
M1 | B | 6 | 1.13 | -1.44 | 0 | -4.84 | 14.86 | -1.31 | 13.36 | -3.63 | 0.01423 | 0.15499 | 0.449 | 6 | 162
M1 | B | 7 | 1.23 | -1.51 | 0 | -4.88 | 15.13 | -1.33 | 0.65 | -0.46 | 0.01554 | 0.15915 | 0.449 | 6 | 162
M1 | B | 8 | 1.22 | -1.48 | 0 | -4.93 | 15.17 | -1.32 | 16.65 | -3.95 | 0.01519 | 0.16058 | 0.449 | 6 | 162
M1 | B | 9 | 1.25 | -1.46 | 0 | -5.02 | 15.54 | -1.3 | 1.8 | -1.39 | 0.01551 | 0.16487 | 0.449 | 6 | 162
M1 | B | 10 | 1.09 | -1.56 | 0 | -5.03 | 15.64 | -1.3 | 0.36 | 0.01 | 0.01372 | 0.16684 | 0.449 | 6 | 162
M1 | B | 11 | 1.23 | -1.54 | 0 | -5.03 | 16 | -1.26 | 23.53 | -4.67 | 0.01546 | 0.17085 | 0.449 | 6 | 162
M1 | B | 12 | 1.24 | -1.48 | 0 | -5.08 | 15.19 | -1.29 | 8.75 | -3.01 | 0.01531 | 0.16377 | 0.449 | 6 | 162
M1 | B | 13 | 1.1 | -1.58 | 0 | -5.08 | 15 | -1.28 | 40.89 | -4.7 | 0.01387 | 0.16264 | 0.449 | 6 | 162
M1 | B | 14 | 1.28 | -1.35 | 0 | -5.09 | 15.23 | -1.25 | 0.51 | -0.16 | 0.01542 | 0.16547 | 0.449 | 6 | 162
M1 | B | 15 | 1.21 | -1.59 | 0 | -5.06 | 15.96 | -1.19 | 0.87 | -0.81 | 0.01521 | 0.17351 | 0.449 | 6 | 162
M1 | B | 16 | 1.29 | -1.36 | 0 | -4.99 | 15.97 | -1.19 | 4.94 | -2.37 | 0.01556 | 0.17443 | 0.449 | 6 | 162
M2 | A | 1 | 1.4 | -1.01 | 0.06 | -3.55 | 15.39 | -1.2 | 0.21 | 1.58 | 0.01957 | 0.1904 | 0.4 | 6 | 180
M2 | A | 2 | 1.11 | -0.85 | 0.05 | -3.67 | 12.83 | -1.2 | 154.72 | -5.81 | 0.01494 | 0.16069 | 0.4 | 6 | 180
M2 | A | 3 | 0.93 | -0.9 | 0.05 | -3.66 | 11.85 | -1.3 | 1 | -0.61 | 0.01257 | 0.15138 | 0.4 | 6 | 180
M2 | A | 4 | 1.1 | -1.11 | 0.05 | -3.68 | 12.91 | -1.27 | 0.92 | -0.42 | 0.01524 | 0.16611 | 0.4 | 6 | 180
M2 | A | 5 | 1.08 | -1.06 | 0.06 | -3.66 | 13.01 | -1.34 | 3.29 | -1.88 | 0.01478 | 0.16991 | 0.4 | 6 | 180
M2 | A | 6 | 1.37 | -1.03 | 0.07 | -3.61 | 14.12 | -1.29 | 4.82 | -2.32 | 0.01847 | 0.18414 | 0.4 | 6 | 180
M2 | A | 7 | 1.21 | -1.19 | 0.05 | -3.78 | 12.75 | -1.36 | 0.26 | 1.02 | 0.0166 | 0.16869 | 0.4 | 6 | 180
M2 | A | 8 | 1.19 | -1.03 | 0.05 | -3.74 | 13.18 | -1.34 | 0.49 | 0.09 | 0.01587 | 0.17508 | 0.4 | 6 | 180
M2 | A | 9 | 1.48 | -1.09 | 0.06 | -3.7 | 14.17 | -1.3 | 2.3 | -1.22 | 0.01988 | 0.18824 | 0.4 | 6 | 180
M2 | A | 10 | 1.43 | -1.05 | 0.04 | -3.83 | 13.82 | -1.31 | 11.53 | -2.95 | 0.01893 | 0.18481 | 0.4 | 6 | 180
M2 | A | 11 | 1.13 | -1.27 | 0.04 | -3.87 | 13.26 | -1.36 | 39.44 | -4.33 | 0.01552 | 0.17947 | 0.4 | 6 | 180
M2 | A | 12 | 1.12 | -1.14 | 0.04 | -3.89 | 13.46 | -1.31 | 0.45 | 0.24 | 0.01501 | 0.18174 | 0.4 | 6 | 180
M2 | A | 13 | 1.03 | -1.15 | 0.03 | -3.96 | 12.46 | -1.34 | 27.99 | -3.79 | 0.01381 | 0.17012 | 0.4 | 6 | 180
M2 | A | 14 | 1.23 | -1.11 | 0.03 | -3.98 | 12.07 | -1.35 | 10.23 | -3.04 | 0.01626 | 0.16615 | 0.4 | 6 | 180
M2 | A | 15 | 1.27 | -1.14 | 0.03 | -3.93 | 12.37 | -1.33 | 0 | 7.63 | 0.01691 | 0.17086 | 0.4 | 6 | 180
M2 | A | 16 | 1.43 | -0.88 | 0.02 | -4.1 | 12.44 | -1.23 | 49.58 | -4.45 | 0.01827 | 0.17069 | 0.4 | 6 | 180
M2 | B | 1 | 1.37 | -0.68 | 0.09 | -3.38 | 15.87 | -1.19 | 35.16 | -4.06 | 0.01887 | 0.18474 | 0.4 | 6 | 180
M2 | B | 2 | 1.28 | -0.87 | 0.06 | -3.52 | 13.23 | -1.2 | 0.94 | -0.46 | 0.01798 | 0.15623 | 0.4 | 6 | 180
M2 | B | 3 | 0.98 | -0.96 | 0.07 | -3.47 | 11.92 | -1.3 | 10.58 | -3.63 | 0.01388 | 0.14391 | 0.4 | 6 | 180
M2 | B | 4 | 1.16 | -1.11 | 0.06 | -3.55 | 12.96 | -1.27 | 0.29 | 0.69 | 0.01669 | 0.15788 | 0.4 | 6 | 180
M2 | B | 5 | 1.33 | -1.05 | 0.07 | -3.53 | 12.69 | -1.37 | 0.14 | 1.55 | 0.01883 | 0.15759 | 0.4 | 6 | 180
M2 | B | 6 | 1.26 | -1.06 | 0.1 | -3.4 | 14.08 | -1.3 | 42.84 | -4.53 | 0.01785 | 0.17435 | 0.4 | 6 | 180
M2 | B | 7 | 1.08 | -1 | 0.08 | -3.51 | 12.9 | -1.34 | 2.53 | -1.55 | 0.01504 | 0.16187 | 0.4 | 6 | 180
M2 | B | 8 | 1.28 | -1.07 | 0.07 | -3.59 | 12.92 | -1.35 | 53.03 | -4.73 | 0.01788 | 0.16342 | 0.4 | 6 | 180
M2 | B | 9 | 1.47 | -1.1 | 0.08 | -3.54 | 13.68 | -1.33 | 2.38 | -1.45 | 0.02049 | 0.17354 | 0.4 | 6 | 180
M2 | B | 10 | 1.43 | -1.04 | 0.07 | -3.6 | 13.38 | -1.35 | 0.25 | 1.13 | 0.01966 | 0.17131 | 0.4 | 6 | 180
M2 | B | 11 | 1.37 | -1.04 | 0.06 | -3.69 | 12.9 | -1.37 | 1.45 | -1.03 | 0.01879 | 0.16641 | 0.4 | 6 | 180
M2 | B | 12 | 1.35 | -1.2 | 0.05 | -3.72 | 13.14 | -1.33 | 5.51 | -2.26 | 0.01906 | 0.16983 | 0.4 | 6 | 180
M2 | B | 13 | 1.43 | -1.06 | 0.05 | -3.7 | 11.78 | -1.38 | 0.08 | 2.1 | 0.01973 | 0.15461 | 0.4 | 6 | 180
M2 | B | 14 | 1.11 | -1.2 | 0.05 | -3.73 | 11.91 | -1.35 | 40.83 | -4.09 | 0.01552 | 0.15639 | 0.4 | 6 | 180
M2 | B | 15 | 1.42 | -1.14 | 0.05 | -3.74 | 12.41 | -1.32 | 0.99 | -0.37 | 0.01978 | 0.16334 | 0.4 | 6 | 180
M2 | B | 16 | 1.28 | -1.15 | 0.03 | -3.87 | 12.09 | -1.25 | 0.38 | 0.58 | 0.01777 | 0.15863 | 0.4 | 6 | 180
M3 | A | 1 | 0.81 | -1.52 | 4.75 | -1.74 | 10.87 | -1.23 | 1.39 | -1.74 | 0.01142 | 0.12412 | 0.32 | 6 | 160
M3 | A | 2 | 0.63 | -1.5 | 3.56 | -1.86 | 8.68 | -1.3 | 1.57 | -2.13 | 0.00864 | 0.10133 | 0.32 | 6 | 160
M3 | A | 3 | 0.62 | -1.67 | 3.58 | -1.88 | 8.47 | -1.33 | 0.93 | -1.55 | 0.00878 | 0.10016 | 0.32 | 6 | 160
M3 | A | 4 | 0.71 | -1.65 | 4.23 | -1.85 | 9.22 | -1.36 | 1.35 | -1.96 | 0.00985 | 0.11065 | 0.32 | 6 | 160
M3 | A | 5 | 0.82 | -1.6 | 4.72 | -1.82 | 9.84 | -1.37 | 1.08 | -1.52 | 0.01121 | 0.11887 | 0.32 | 6 | 160
M3 | A | 6 | 0.78 | -1.64 | 4.69 | -1.83 | 9.86 | -1.37 | 0.75 | -1.13 | 0.01075 | 0.11991 | 0.32 | 6 | 160
M3 | A | 7 | 0.82 | -1.65 | 4.56 | -1.86 | 9.72 | -1.4 | 1.71 | -2.06 | 0.01115 | 0.11959 | 0.32 | 6 | 160
M3 | A | 8 | 0.79 | -1.63 | 4.37 | -1.88 | 9.62 | -1.4 | 1.09 | -1.48 | 0.01067 | 0.11942 | 0.32 | 6 | 160
M3 | A | 9 | 0.86 | -1.65 | 4.35 | -1.9 | 9.92 | -1.39 | 1.06 | -1.5 | 0.01162 | 0.12343 | 0.32 | 6 | 160
M3 | A | 10 | 0.8 | -1.67 | 4.14 | -1.93 | 9.87 | -1.4 | 1.27 | -1.83 | 0.01078 | 0.12375 | 0.32 | 6 | 160
M3 | A | 11 | 0.8 | -1.64 | 3.87 | -1.96 | 9.82 | -1.37 | 0.87 | -1.42 | 0.01062 | 0.12349 | 0.32 | 6 | 160
M3 | A | 12 | 0.78 | -1.72 | 3.61 | -1.99 | 9.63 | -1.38 | 1.1 | -1.6 | 0.01056 | 0.12208 | 0.32 | 6 | 160
M3 | A | 13 | 0.81 | -1.69 | 3.38 | -2.02 | 9.57 | -1.35 | 1.76 | -2.13 | 0.01091 | 0.1217 | 0.32 | 6 | 160
M3 | A | 14 | 0.83 | -1.61 | 3.15 | -2.05 | 9.73 | -1.32 | 0.89 | -1.35 | 0.01094 | 0.12399 | 0.32 | 6 | 160
M3 | A | 15 | 0.87 | -1.51 | 2.89 | -2.08 | 9.81 | -1.29 | 0.71 | -1.17 | 0.01116 | 0.12569 | 0.32 | 6 | 160
M3 | A | 16 | 0.88 | -1.6 | 2.54 | -2.13 | 9.77 | -1.27 | 1.03 | -1.34 | 0.01157 | 0.12561 | 0.32 | 6 | 160
M3 | B | 1 | 0.86 | -1.46 | 5.42 | -1.66 | 11.03 | -1.23 | 1.3 | -1.68 | 0.01238 | 0.11866 | 0.32 | 6 | 160
M3 | B | 2 | 0.63 | -1.54 | 4.03 | -1.78 | 8.74 | -1.3 | 0.7 | -1.27 | 0.00914 | 0.09625 | 0.32 | 6 | 160
M3 | B | 3 | 0.64 | -1.6 | 4.05 | -1.79 | 8.5 | -1.33 | 1.46 | -2.06 | 0.00928 | 0.09504 | 0.32 | 6 | 160
M3 | B | 4 | 0.72 | -1.58 | 4.75 | -1.76 | 9.18 | -1.35 | 1.36 | -1.86 | 0.01032 | 0.10429 | 0.32 | 6 | 160
M3 | B | 5 | 0.84 | -1.57 | 5.3 | -1.74 | 9.77 | -1.38 | 1.82 | -2.05 | 0.01194 | 0.11201 | 0.32 | 6 | 160
M3 | B | 6 | 0.79 | -1.67 | 5.3 | -1.74 | 9.78 | -1.37 | 1.51 | -1.92 | 0.01133 | 0.11294 | 0.32 | 6 | 160
M3 | B | 7 | 0.83 | -1.57 | 5.13 | -1.76 | 9.65 | -1.39 | 1.14 | -1.73 | 0.01163 | 0.11265 | 0.32 | 6 | 160
M3 | B | 8 | 0.77 | -1.69 | 4.91 | -1.79 | 9.46 | -1.41 | 2.34 | -2.45 | 0.01092 | 0.11175 | 0.32 | 6 | 160
M3 | B | 9 | 0.86 | -1.61 | 4.88 | -1.81 | 9.72 | -1.39 | 1.14 | -1.55 | 0.01208 | 0.11515 | 0.32 | 6 | 160
M3 | B | 10 | 0.81 | -1.64 | 4.64 | -1.84 | 9.67 | -1.4 | 1.62 | -2.03 | 0.01128 | 0.1155 | 0.32 | 6 | 160
M3 | B | 11 | 0.78 | -1.7 | 4.3 | -1.87 | 9.56 | -1.38 | 0.57 | -0.96 | 0.01097 | 0.11485 | 0.32 | 6 | 160
M3 | B | 12 | 0.8 | -1.61 | 4 | -1.9 | 9.41 | -1.38 | 0.88 | -1.31 | 0.01107 | 0.11375 | 0.32 | 6 | 160
M3 | B | 13 | 0.82 | -1.66 | 3.75 | -1.93 | 9.36 | -1.36 | 0.54 | -0.75 | 0.01138 | 0.11361 | 0.32 | 6 | 160
M3 | B | 14 | 0.82 | -1.61 | 3.52 | -1.95 | 9.45 | -1.35 | 1.24 | -1.74 | 0.01135 | 0.11528 | 0.32 | 6 | 160
M3 | B | 15 | 0.84 | -1.6 | 3.26 | -1.98 | 9.48 | -1.32 | 0.47 | -0.56 | 0.01156 | 0.11638 | 0.32 | 6 | 160
M3 | B | 16 | 0.92 | -1.46 | 2.91 | -2.02 | 9.48 | -1.28 | 1.23 | -1.61 | 0.01223 | 0.11677 | 0.32 | 6 | 160
M4 | A | 1 | 0.63 | -1.58 | 4.48 | -1.65 | 8.87 | -1.23 | 1.4 | -1.94 | 0.00956 | 0.11042 | 0.22 | 6 | 160
M4 | A | 2 | 0.51 | -1.64 | 3.35 | -1.76 | 7.1 | -1.28 | 0.53 | -1.21 | 0.00773 | 0.08997 | 0.22 | 6 | 160
M4 | A | 3 | 0.52 | -1.59 | 3.43 | -1.77 | 7 | -1.31 | 0.81 | -1.54 | 0.00781 | 0.0898 | 0.22 | 6 | 160
M4 | A | 4 | 0.58 | -1.7 | 4.15 | -1.72 | 7.77 | -1.32 | 0.99 | -1.68 | 0.00872 | 0.10073 | 0.22 | 6 | 160
M4 | A | 5 | 0.62 | -1.71 | 4.86 | -1.67 | 8.37 | -1.33 | 0.7 | -1.37 | 0.00927 | 0.10915 | 0.22 | 6 | 160
M4 | A | 6 | 0.65 | -1.64 | 4.96 | -1.66 | 8.43 | -1.33 | 0.64 | -1.11 | 0.00962 | 0.11053 | 0.22 | 6 | 160
M4 | A | 7 | 0.63 | -1.71 | 4.86 | -1.67 | 8.33 | -1.34 | 1.82 | -2.38 | 0.00931 | 0.11023 | 0.22 | 6 | 160
M4 | A | 8 | 0.64 | -1.69 | 4.79 | -1.68 | 8.32 | -1.34 | 0.94 | -1.66 | 0.00931 | 0.11067 | 0.22 | 6 | 160
M4 | A | 9 | 0.66 | -1.68 | 4.85 | -1.69 | 8.56 | -1.33 | 0.81 | -1.4 | 0.00961 | 0.11429 | 0.22 | 6 | 160
M4 | A | 10 | 0.7 | -1.62 | 4.68 | -1.7 | 8.54 | -1.33 | 0.84 | -1.45 | 0.01006 | 0.11463 | 0.22 | 6 | 160
M4 | A | 11 | 0.67 | -1.68 | 4.36 | -1.74 | 8.44 | -1.32 | 0.55 | -1.08 | 0.00968 | 0.11404 | 0.22 | 6 | 160
M4 | A | 12 | 0.68 | -1.66 | 4.12 | -1.76 | 8.41 | -1.31 | 0.81 | -1.33 | 0.00976 | 0.11421 | 0.22 | 6 | 160
M4 | A | 13 | 0.68 | -1.66 | 3.91 | -1.78 | 8.3 | -1.3 | 0.64 | -1.03 | 0.00982 | 0.11324 | 0.22 | 6 | 160
M4 | A | 14 | 0.69 | -1.67 | 3.64 | -1.81 | 8.44 | -1.28 | 0.51 | -0.82 | 0.0099 | 0.11547 | 0.22 | 6 | 160
M4 | A | 15 | 0.72 | -1.59 | 3.36 | -1.86 | 8.64 | -1.27 | 0.71 | -1.23 | 0.01019 | 0.11875 | 0.22 | 6 | 160
M4 | A | 16 | 0.74 | -1.64 | 2.9 | -1.91 | 8.81 | -1.23 | 0.86 | -1.21 | 0.01045 | 0.12111 | 0.22 | 6 | 160
M4 | B | 1 | 0.66 | -1.54 | 4.74 | -1.61 | 8.96 | -1.23 | 1.78 | -2.21 | 0.01028 | 0.10521 | 0.22 | 6 | 160
M4 | B | 2 | 0.52 | -1.62 | 3.5 | -1.73 | 7.1 | -1.28 | 0.87 | -1.7 | 0.0082 | 0.08501 | 0.22 | 6 | 160
M4 | B | 3 | 0.52 | -1.59 | 3.54 | -1.74 | 6.97 | -1.31 | 0.91 | -1.65 | 0.00804 | 0.0845 | 0.22 | 6 | 160
M4 | B | 4 | 0.6 | -1.62 | 4.25 | -1.69 | 7.66 | -1.32 | 1.29 | -1.98 | 0.00926 | 0.09422 | 0.22 | 6 | 160
M4 | B | 5 | 0.62 | -1.71 | 4.99 | -1.64 | 8.24 | -1.33 | 1.02 | -1.71 | 0.00971 | 0.10204 | 0.22 | 6 | 160
M4 | B | 6 | 0.66 | -1.63 | 5.09 | -1.63 | 8.28 | -1.33 | 0.65 | -1.09 | 0.01003 | 0.1032 | 0.22 | 6 | 160
M4 | B | 7 | 0.63 | -1.69 | 5.01 | -1.64 | 8.16 | -1.34 | 0.65 | -1.27 | 0.00966 | 0.10258 | 0.22 | 6 | 160
M4 | B | 8 | 0.64 | -1.67 | 4.91 | -1.65 | 8.1 | -1.35 | 0.8 | -1.47 | 0.00978 | 0.10268 | 0.22 | 6 | 160
M4 | B | 9 | 0.69 | -1.63 | 4.98 | -1.66 | 8.34 | -1.35 | 1.3 | -1.94 | 0.01032 | 0.10619 | 0.22 | 6 | 160
M4 | B | 10 | 0.69 | -1.67 | 4.81 | -1.67 | 8.32 | -1.33 | 0.94 | -1.61 | 0.01046 | 0.10647 | 0.22 | 6 | 160
M4 | B | 11 | 0.67 | -1.67 | 4.45 | -1.71 | 8.21 | -1.33 | 1.19 | -1.85 | 0.01013 | 0.10591 | 0.22 | 6 | 160
M4 | B | 12 | 0.69 | -1.63 | 4.24 | -1.72 | 8.15 | -1.32 | 1.45 | -2.05 | 0.01027 | 0.1057 | 0.22 | 6 | 160
M4 | B | 13 | 0.7 | -1.59 | 4.01 | -1.74 | 8.02 | -1.32 | 0.96 | -1.51 | 0.01033 | 0.10475 | 0.22 | 6 | 160
M4 | B | 14 | 0.71 | -1.63 | 3.76 | -1.77 | 8.14 | -1.3 | 0.9 | -1.52 | 0.01055 | 0.10659 | 0.22 | 6 | 160
M4 | B | 15 | 0.73 | -1.58 | 3.47 | -1.82 | 8.35 | -1.28 | 0.85 | -1.4 | 0.01067 | 0.11 | 0.22 | 6 | 160
M4 | B | 16 | 0.74 | -1.62 | 3.05 | -1.86 | 8.51 | -1.25 | 1.37 | -1.81 | 0.01095 | 0.11225 | 0.22 | 6 | 160
M5 | A | 1 | 0.6 | -1.53 | 4.83 | -1.5 | 6.13 | -1.32 | 0.94 | -1.48 | 0.0113 | 0.05693 | 0.1 | 6 | 115
M5 | A | 2 | 0.48 | -1.64 | 4.06 | -1.57 | 5.19 | -1.36 | 0.63 | -1.3 | 0.00932 | 0.04926 | 0.1 | 6 | 115
M5 | A | 3 | 0.49 | -1.71 | 4.21 | -1.58 | 5.22 | -1.39 | 0.71 | -1.4 | 0.00957 | 0.05024 | 0.1 | 6 | 115
M5 | A | 4 | 0.54 | -1.62 | 5.04 | -1.54 | 5.87 | -1.4 | 0.76 | -1.41 | 0.01069 | 0.05761 | 0.1 | 6 | 115
M5 | A | 5 | 0.61 | -1.7 | 5.73 | -1.51 | 6.41 | -1.4 | 0.82 | -1.41 | 0.01201 | 0.06336 | 0.1 | 6 | 115
M5 | A | 6 | 0.63 | -1.62 | 5.88 | -1.5 | 6.46 | -1.4 | 0.85 | -1.33 | 0.01218 | 0.06434 | 0.1 | 6 | 115
M5 | A | 7 | 0.61 | -1.69 | 5.9 | -1.5 | 6.43 | -1.41 | 0.81 | -1.39 | 0.01192 | 0.06474 | 0.1 | 6 | 115
M5 | A | 8 | 0.62 | -1.74 | 5.89 | -1.5 | 6.41 | -1.41 | 0.82 | -1.47 | 0.01191 | 0.0652 | 0.1 | 6 | 115
M5 | A | 9 | 0.66 | -1.73 | 5.98 | -1.5 | 6.55 | -1.4 | 0.85 | -1.43 | 0.01245 | 0.06693 | 0.1 | 6 | 115
M5 | A | 10 | 0.63 | -1.72 | 5.89 | -1.51 | 6.52 | -1.4 | 0.75 | -1.26 | 0.01197 | 0.06708 | 0.1 | 6 | 115
M5 | A | 11 | 0.66 | -1.76 | 5.72 | -1.52 | 6.43 | -1.39 | 0.81 | -1.43 | 0.01209 | 0.06661 | 0.1 | 6 | 115
M5 | A | 12 | 0.62 | -1.73 | 5.49 | -1.53 | 6.32 | -1.39 | 0.73 | -1.34 | 0.01145 | 0.06586 | 0.1 | 6 | 115
M5 | A | 13 | 0.63 | -1.75 | 5.27 | -1.54 | 6.22 | -1.38 | 0.77 | -1.39 | 0.01131 | 0.06532 | 0.1 | 6 | 115
M5 | A | 14 | 0.65 | -1.61 | 5.07 | -1.55 | 6.18 | -1.36 | 0.86 | -1.36 | 0.0112 | 0.06523 | 0.1 | 6 | 115
M5 | A | 15 | 0.62 | -1.6 | 5.44 | -1.45 | 6.16 | -1.33 | 0.77 | -1.28 | 0.01093 | 0.06528 | 0.1 | 6 | 115
M5 | A | 16 | 0.63 | -1.58 | 5.24 | -1.43 | 6.12 | -1.28 | 0.95 | -1.46 | 0.01076 | 0.06493 | 0.1 | 6 | 115
M5 | B | 1 | 0.62 | -1.49 | 4.81 | -1.5 | 6.13 | -1.32 | 0.97 | -1.45 | 0.01171 | 0.05425 | 0.1 | 6 | 115
M5 | B | 2 | 0.49 | -1.67 | 4.02 | -1.57 | 5.14 | -1.38 | 0.72 | -1.42 | 0.00966 | 0.0466 | 0.1 | 6 | 115
M5 | B | 3 | 0.49 | -1.7 | 4.13 | -1.58 | 5.14 | -1.4 | 0.74 | -1.4 | 0.00982 | 0.04735 | 0.1 | 6 | 115
M5 | B | 4 | 0.55 | -1.62 | 4.91 | -1.54 | 5.75 | -1.41 | 0.79 | -1.42 | 0.01096 | 0.05402 | 0.1 | 6 | 115
M5 | B | 5 | 0.62 | -1.75 | 5.57 | -1.51 | 6.24 | -1.42 | 0.83 | -1.36 | 0.01239 | 0.05925 | 0.1 | 6 | 115
M5 | B | 6 | 0.64 | -1.63 | 5.7 | -1.5 | 6.27 | -1.42 | 0.87 | -1.39 | 0.01247 | 0.06008 | 0.1 | 6 | 115
M5 | B | 7 | 0.62 | -1.71 | 5.71 | -1.5 | 6.23 | -1.42 | 0.79 | -1.32 | 0.01226 | 0.06035 | 0.1 | 6 | 115
M5 | B | 8 | 0.6 | -1.73 | 5.68 | -1.5 | 6.2 | -1.42 | 0.73 | -1.31 | 0.01196 | 0.06071 | 0.1 | 6 | 115
M5 | B | 9 | 0.65 | -1.7 | 5.78 | -1.5 | 6.32 | -1.42 | 0.83 | -1.38 | 0.01249 | 0.06232 | 0.1 | 6 | 115
M5 | B | 10 | 0.64 | -1.69 | 5.69 | -1.51 | 6.27 | -1.42 | 0.72 | -1.23 | 0.01227 | 0.06225 | 0.1 | 6 | 115
M5 | B | 11 | 0.64 | -1.73 | 5.5 | -1.52 | 6.18 | -1.41 | 0.75 | -1.35 | 0.01212 | 0.06188 | 0.1 | 6 | 115
M5 | B | 12 | 0.62 | -1.74 | 5.3 | -1.53 | 6.07 | -1.41 | 0.95 | -1.63 | 0.01167 | 0.06112 | 0.1 | 6 | 115
M5 | B | 13 | 0.62 | -1.67 | 5.08 | -1.54 | 5.97 | -1.4 | 0.76 | -1.35 | 0.01134 | 0.06061 | 0.1 | 6 | 115
M5 | B | 14 | 0.63 | -1.55 | 4.9 | -1.54 | 5.93 | -1.38 | 0.75 | -1.2 | 0.01122 | 0.06056 | 0.1 | 6 | 115
M5 | B | 15 | 0.63 | -1.68 | 4.71 | -1.55 | 5.89 | -1.35 | 1.01 | -1.58 | 0.01106 | 0.0605 | 0.1 | 6 | 115
M5 | B | 16 | 0.63 | -1.57 | 5.12 | -1.42 | 5.86 | -1.31 | 0.9 | -1.39 | 0.011 | 0.06021 | 0.1 | 6 | 115
M6 | A | 1 | 0.6 | -1.68 | 5.04 | -1.4 | 5.54 | -1.33 | 0.97 | -1.59 | 0.00194 | 0.01244 | 0.096 | 6 | 147
M6 | A | 2 | 0.49 | -1.61 | 4.29 | -1.47 | 4.72 | -1.38 | 0.78 | -1.42 | 0.0016 | 0.01063 | 0.096 | 6 | 147
M6 | A | 3 | 0.48 | -1.67 | 4.39 | -1.48 | 4.71 | -1.42 | 0.75 | -1.39 | 0.00153 | 0.01058 | 0.096 | 6 | 147
M6 | A | 4 | 0.57 | -1.61 | 5.23 | -1.46 | 5.37 | -1.43 | 0.8 | -1.3 | 0.00181 | 0.01222 | 0.096 | 6 | 147
M6 | A | 5 | 0.63 | -1.69 | 5.92 | -1.44 | 5.87 | -1.43 | 0.86 | -1.37 | 0.00193 | 0.01339 | 0.096 | 6 | 147
M6 | A | 6 | 0.64 | -1.72 | 6.06 | -1.43 | 5.89 | -1.43 | 0.89 | -1.38 | 0.00195 | 0.01356 | 0.096 | 6 | 147
M6 | A | 7 | 0.63 | -1.66 | 6.1 | -1.42 | 5.87 | -1.44 | 0.85 | -1.39 | 0.00193 | 0.0136 | 0.096 | 6 | 147
M6 | A | 8 | 0.63 | -1.71 | 6.14 | -1.42 | 5.88 | -1.44 | 0.87 | -1.45 | 0.00189 | 0.01371 | 0.096 | 6 | 147
M6 | A | 9 | 0.67 | -1.67 | 6.29 | -1.41 | 6.03 | -1.44 | 0.86 | -1.34 | 0.00202 | 0.0142 | 0.096 | 6 | 147
M6 | A | 10 | 0.66 | -1.7 | 6.23 | -1.42 | 6.02 | -1.43 | 0.83 | -1.33 | 0.00196 | 0.01428 | 0.096 | 6 | 147
M6 | A | 11 | 0.66 | -1.63 | 6.1 | -1.42 | 5.97 | -1.42 | 0.81 | -1.29 | 0.002 | 0.01432 | 0.096 | 6 | 147
M6 | A | 12 | 0.68 | -1.69 | 5.96 | -1.43 | 5.88 | -1.42 | 0.81 | -1.31 | 0.00199 | 0.01423 | 0.096 | 6 | 147
M6 | A | 13 | 0.65 | -1.66 | 5.79 | -1.43 | 5.81 | -1.4 | 0.81 | -1.37 | 0.00193 | 0.01422 | 0.096 | 6 | 147
M6 | A | 14 | 0.66 | -1.67 | 5.75 | -1.41 | 5.79 | -1.38 | 0.83 | -1.38 | 0.00195 | 0.01437 | 0.096 | 6 | 147
M6 | A | 15 | 0.66 | -1.55 | 5.6 | -1.41 | 5.83 | -1.35 | 0.91 | -1.41 | 0.00204 | 0.01485 | 0.096 | 6 | 147
M6 | A | 16 | 0.67 | -1.59 | 5.49 | -1.37 | 5.79 | -1.31 | 0.86 | -1.31 | 0.00202 | 0.01501 | 0.096 | 6 | 147
M6 | B | 1 | 0.62 | -1.68 | 5 | -1.4 | 5.56 | -1.34 | 0.94 | -1.5 | 0.00207 | 0.01169 | 0.096 | 6 | 147
M6 | B | 2 | 0.5 | -1.62 | 4.23 | -1.47 | 4.71 | -1.37 | 0.79 | -1.38 | 0.00169 | 0.01003 | 0.096 | 6 | 147
M6 | B | 3 | 0.49 | -1.66 | 4.3 | -1.49 | 4.68 | -1.41 | 0.76 | -1.36 | 0.00162 | 0.00998 | 0.096 | 6 | 147
M6 | B | 4 | 0.58 | -1.62 | 5.09 | -1.47 | 5.26 | -1.45 | 0.8 | -1.28 | 0.00191 | 0.01128 | 0.096 | 6 | 147
M6 | B | 5 | 0.64 | -1.71 | 5.72 | -1.45 | 5.74 | -1.45 | 0.89 | -1.37 | 0.00201 | 0.01234 | 0.096 | 6 | 147
M6 | B | 6 | 0.66 | -1.71 | 5.85 | -1.44 | 5.75 | -1.45 | 0.89 | -1.36 | 0.00207 | 0.01248 | 0.096 | 6 | 147
M6 | B | 7 | 0.63 | -1.66 | 5.86 | -1.44 | 5.7 | -1.46 | 0.82 | -1.34 | 0.00201 | 0.01246 | 0.096 | 6 | 147
M6 | B | 8 | 0.64 | -1.69 | 5.89 | -1.43 | 5.69 | -1.46 | 0.87 | -1.43 | 0.00199 | 0.01259 | 0.096 | 6 | 147
M6 | B | 9 | 0.67 | -1.68 | 6.03 | -1.43 | 5.84 | -1.45 | 0.84 | -1.3 | 0.0021 | 0.01303 | 0.096 | 6 | 147
M6 | B | 10 | 0.67 | -1.71 | 5.96 | -1.43 | 5.81 | -1.45 | 0.83 | -1.34 | 0.00206 | 0.01307 | 0.096 | 6 | 147
M6 | B | 11 | 0.67 | -1.64 | 5.83 | -1.43 | 5.72 | -1.45 | 0.82 | -1.3 | 0.00212 | 0.01299 | 0.096 | 6 | 147
M6 | B | 12 | 0.67 | -1.62 | 5.71 | -1.43 | 5.64 | -1.44 | 0.85 | -1.35 | 0.00211 | 0.01297 | 0.096 | 6 | 147
M6 | B | 13 | 0.65 | -1.65 | 5.56 | -1.43 | 5.56 | -1.43 | 0.74 | -1.27 | 0.00201 | 0.01287 | 0.096 | 6 | 147
M6 | B | 14 | 0.66 | -1.69 | 5.46 | -1.43 | 5.53 | -1.41 | 0.79 | -1.34 | 0.002 | 0.01301 | 0.096 | 6 | 147
M6 | B | 15 | 0.66 | -1.56 | 5.39 | -1.41 | 5.56 | -1.38 | 0.88 | -1.37 | 0.00213 | 0.01343 | 0.096 | 6 | 147
M6 | B | 16 | 0.67 | -1.61 | 5.3 | -1.37 | 5.55 | -1.33 | 0.86 | -1.32 | 0.0021 | 0.01369 | 0.096 | 6 | 147
M7 | A | 1 | 0.48 | -1.49 | 3.87 | -1.44 | 4.56 | -1.23 | 0.85 | -1.5 | 0.00613 | 0.04807 | 0.064 | 6 | 124
M7 | A | 2 | 0.4 | -1.56 | 3.66 | -1.43 | 3.93 | -1.28 | 0.74 | -1.55 | 0.00517 | 0.04231 | 0.064 | 6 | 124
M7 | A | 3 | 0.41 | -1.56 | 3.6 | -1.49 | 3.94 | -1.32 | 0.73 | -1.54 | 0.00519 | 0.04305 | 0.064 | 6 | 124
M7 | A | 4 | 0.48 | -1.59 | 4.45 | -1.45 | 4.56 | -1.36 | 0.82 | -1.55 | 0.0061 | 0.05084 | 0.064 | 6 | 124
M7 | A | 5 | 0.54 | -1.63 | 5.32 | -1.38 | 5.03 | -1.36 | 0.9 | -1.57 | 0.00683 | 0.05646 | 0.064 | 6 | 124
M7 | A | 6 | 0.55 | -1.61 | 5.34 | -1.39 | 5.04 | -1.38 | 0.88 | -1.54 | 0.00685 | 0.05701 | 0.064 | 6 | 124
M7 | A | 7 | 0.56 | -1.57 | 5.36 | -1.39 | 5.01 | -1.39 | 0.91 | -1.56 | 0.00682 | 0.05722 | 0.064 | 6 | 124
M7 | A | 8 | 0.55 | -1.63 | 5.32 | -1.39 | 4.97 | -1.39 | 0.93 | -1.61 | 0.00678 | 0.05732 | 0.064 | 6 | 124
M7 | A | 9 | 0.57 | -1.65 | 5.4 | -1.4 | 5.07 | -1.39 | 0.91 | -1.57 | 0.00699 | 0.05882 | 0.064 | 6 | 124
M7 | A | 10 | 0.58 | -1.59 | 5.28 | -1.41 | 5.07 | -1.38 | 0.92 | -1.54 | 0.00701 | 0.05921 | 0.064 | 6 | 124
M7 | A | 11 | 0.57 | -1.63 | 5.17 | -1.41 | 5.01 | -1.38 | 0.91 | -1.56 | 0.00687 | 0.05888 | 0.064 | 6 | 124
M7 | A | 12 | 0.56 | -1.63 | 5 | -1.42 | 4.95 | -1.37 | 0.93 | -1.58 | 0.0068 | 0.05858 | 0.064 | 6 | 124
M7 | A | 13 | 0.56 | -1.61 | 4.74 | -1.44 | 4.9 | -1.36 | 0.95 | -1.58 | 0.00671 | 0.05817 | 0.064 | 6 | 124
M7 | A | 14 | 0.57 | -1.54 | 4.59 | -1.44 | 4.88 | -1.33 | 0.94 | -1.5 | 0.00669 | 0.05816 | 0.064 | 6 | 124
M7 | A | 15 | 0.57 | -1.57 | 4.2 | -1.49 | 4.88 | -1.29 | 0.95 | -1.5 | 0.00667 | 0.0584 | 0.064 | 6 | 124
M7 | A | 16 | 0.57 | -1.54 | 4.23 | -1.42 | 4.85 | -1.26 | 1.02 | -1.55 | 0.00659 | 0.05808 | 0.064 | 6 | 124
M7 | B | 1 | 0.49 | -1.49 | 3.81 | -1.45 | 4.54 | -1.24 | 0.86 | -1.47 | 0.0065 | 0.04521 | 0.064 | 6 | 124
M7 | B | 2 | 0.41 | -1.55 | 3.46 | -1.47 | 3.89 | -1.28 | 0.77 | -1.56 | 0.00545 | 0.03961 | 0.064 | 6 | 124
M7 | B | 3 | 0.41 | -1.55 | 3.51 | -1.49 | 3.88 | -1.32 | 0.73 | -1.52 | 0.00546 | 0.04014 | 0.064 | 6 | 124
M7 | B | 4 | 0.49 | -1.57 | 4.44 | -1.42 | 4.46 | -1.36 | 0.9 | -1.63 | 0.00638 | 0.04718 | 0.064 | 6 | 124
M7 | B | 5 | 0.54 | -1.62 | 5.12 | -1.39 | 4.91 | -1.38 | 0.89 | -1.53 | 0.00714 | 0.05238 | 0.064 | 6 | 124
M7 | B | 6 | 0.55 | -1.6 | 5.25 | -1.38 | 4.9 | -1.4 | 0.9 | -1.55 | 0.00719 | 0.05283 | 0.064 | 6 | 124
M7 | B | 7 | 0.56 | -1.58 | 5.27 | -1.37 | 4.85 | -1.4 | 0.88 | -1.51 | 0.00713 | 0.05283 | 0.064 | 6 | 124
M7 | B | 8 | 0.55 | -1.62 | 5.2 | -1.38 | 4.8 | -1.41 | 1 | -1.68 | 0.00705 | 0.05284 | 0.064 | 6 | 124
M7 | B | 9 | 0.57 | -1.62 | 5.15 | -1.4 | 4.88 | -1.41 | 0.91 | -1.56 | 0.00725 | 0.05416 | 0.064 | 6 | 124
M7 | B | 10 | 0.58 | -1.61 | 5.21 | -1.38 | 4.87 | -1.4 | 0.93 | -1.55 | 0.00732 | 0.05443 | 0.064 | 6 | 124
M7 | B | 11 | 0.56 | -1.64 | 4.92 | -1.42 | 4.8 | -1.4 | 0.95 | -1.6 | 0.00715 | 0.0541 | 0.064 | 6 | 124
M7 | B | 12 | 0.56 | -1.65 | 4.88 | -1.4 | 4.75 | -1.39 | 0.9 | -1.53 | 0.00711 | 0.05375 | 0.064 | 6 | 124
M7 | B | 13 | 0.56 | -1.61 | 4.55 | -1.43 | 4.69 | -1.37 | 0.94 | -1.58 | 0.00697 | 0.05346 | 0.064 | 6 | 124
M7 | B | 14 | 0.57 | -1.54 | 4.54 | -1.41 | 4.67 | -1.34 | 0.92 | -1.49 | 0.00696 | 0.05343 | 0.064 | 6 | 124
M7 | B | 15 | 0.56 | -1.56 | 3.97 | -1.5 | 4.65 | -1.31 | 0.94 | -1.49 | 0.00692 | 0.05343 | 0.064 | 6 | 124
M7 | B | 16 | 0.56 | -1.53 | 3.71 | -1.51 | 4.63 | -1.27 | 0.99 | -1.53 | 0.00679 | 0.05321 | 0.064 | 6 | 124
M8 | A | 1 | 0.58 | -1.52 | 5.28 | -1.47 | 5.97 | -1.36 | 0.91 | -1.76 | 0.02645 | 0.02747 | 0.054 | 6 | 57
M8 | A | 2 | 0.59 | -1.59 | 4.55 | -1.54 | 5.21 | -1.38 | 0.58 | -1.27 | 0.02533 | 0.02209 | 0.054 | 6 | 57
M8 | A | 3 | 0.46 | -1.46 | 4.63 | -1.58 | 5.28 | -1.45 | 0.57 | -1.46 | 0.02328 | 0.02529 | 0.054 | 6 | 57
M8 | A | 4 | 0.6 | -1.64 | 5.46 | -1.59 | 6.3 | -1.42 | 0.61 | -1.37 | 0.03008 | 0.02762 | 0.054 | 6 | 57
M8 | A | 5 | 0.61 | -1.39 | 6.42 | -1.57 | 7.22 | -1.45 | 0.74 | -1.33 | 0.03156 | 0.03545 | 0.054 | 6 | 57
M8 | A | 6 | 0.65 | -1.58 | 6.4 | -1.58 | 7.06 | -1.51 | 0.65 | -1.3 | 0.03471 | 0.03191 | 0.054 | 6 | 57
M8 | A | 7 | 0.7 | -1.73 | 6.58 | -1.55 | 7.4 | -1.43 | 0.73 | -1.43 | 0.03167 | 0.03683 | 0.054 | 6 | 57
M8 | A | 8 | 0.74 | -1.94 | 6.54 | -1.56 | 7.55 | -1.37 | 0.68 | -1.46 | 0.0346 | 0.03391 | 0.054 | 6 | 57
M8 | A | 9 | 0.59 | -1.33 | 6.61 | -1.54 | 7.45 | -1.42 | 0.79 | -1.41 | 0.03135 | 0.0377 | 0.054 | 6 | 57
M8 | A | 10 | 0.74 | -1.62 | 6.42 | -1.57 | 7.41 | -1.37 | 0.7 | -1.25 | 0.03371 | 0.0338 | 0.054 | 6 | 57
M8 | A | 11 | 0.59 | -1.43 | 6.53 | -1.53 | 7.35 | -1.4 | 0.85 | -1.56 | 0.03046 | 0.03761 | 0.054 | 6 | 57
M8 | A | 12 | 0.65 | -1.5 | 6.32 | -1.56 | 7.31 | -1.37 | 0.65 | -1.22 | 0.03291 | 0.03368 | 0.054 | 6 | 57
M8 | A | 13 | 0.69 | -1.55 | 6.31 | -1.52 | 7.18 | -1.38 | 0.93 | -1.6 | 0.02921 | 0.03719 | 0.054 | 6 | 57
M8 | A | 14 | 0.73 | -1.69 | 6.05 | -1.55 | 7.06 | -1.34 | 0.73 | -1.37 | 0.03123 | 0.03295 | 0.054 | 6 | 57
M8 | A | 15 | 0.59 | -1.44 | 5.74 | -1.53 | 6.69 | -1.35 | 0.81 | -1.53 | 0.02645 | 0.03504 | 0.054 | 6 | 57
M8 | A | 16 | 0.76 | -1.87 | 5.54 | -1.5 | 6.58 | -1.28 | 0.79 | -1.51 | 0.02802 | 0.03098 | 0.054 | 6 | 57
M8 | B | 1 | 0.58 | -1.46 | 5.34 | -1.47 | 6.09 | -1.37 | 0.89 | -1.69 | 0.02758 | 0.02632 | 0.054 | 6 | 57
M8 | B | 2 | 0.63 | -1.65 | 4.56 | -1.54 | 5.26 | -1.4 | 0.62 | -1.33 | 0.02625 | 0.02105 | 0.054 | 6 | 57
M8 | B | 3 | 0.48 | -1.52 | 4.63 | -1.58 | 5.31 | -1.46 | 0.58 | -1.45 | 0.02407 | 0.024 | 0.054 | 6 | 57
M8 | B | 4 | 0.6 | -1.6 | 5.4 | -1.6 | 6.01 | -1.55 | 0.62 | -1.35 | 0.03087 | 0.02542 | 0.054 | 6 | 57
M8 | B | 5 | 0.63 | -1.41 | 6.36 | -1.56 | 7.2 | -1.46 | 0.76 | -1.34 | 0.03246 | 0.03354 | 0.054 | 6 | 57
M8 | B | 6 | 0.69 | -1.61 | 6.37 | -1.58 | 7.39 | -1.41 | 0.67 | -1.33 | 0.03585 | 0.03107 | 0.054 | 6 | 57
M8 | B | 7 | 0.67 | -1.65 | 6.5 | -1.55 | 7.35 | -1.45 | 0.77 | -1.53 | 0.0325 | 0.03479 | 0.054 | 6 | 57
M8 | B | 8 | 0.77 | -2.04 | 6.43 | -1.56 | 7.48 | -1.38 | 0.7 | -1.45 | 0.03541 | 0.03203 | 0.054 | 6 | 57
M8 | B | 9 | 0.59 | -1.33 | 6.57 | -1.53 | 7.37 | -1.43 | 0.85 | -1.51 | 0.0323 | 0.03557 | 0.054 | 6 | 57
M8 | B | 10 | 0.74 | -1.63 | 6.33 | -1.57 | 7.34 | -1.39 | 0.71 | -1.28 | 0.03465 | 0.03196 | 0.054 | 6 | 57
M8 | B | 11 | 0.57 | -1.36 | 6.41 | -1.52 | 7.23 | -1.42 | 0.84 | -1.56 | 0.03113 | 0.03534 | 0.054 | 6 | 57
M8 | B | 12 | 0.56 | -1.28 | 6.2 | -1.56 | 7.19 | -1.38 | 0.62 | -1.15 | 0.03362 | 0.03168 | 0.054 | 6 | 57
M8 | B | 13 | 0.67 | -1.5 | 6.18 | -1.52 | 7.05 | -1.4 | 0.99 | -1.69 | 0.02984 | 0.03495 | 0.054 | 6 | 57
M8 | B | 14 | 0.74 | -1.69 | 5.91 | -1.55 | 6.93 | -1.36 | 0.73 | -1.35 | 0.03184 | 0.03097 | 0.054 | 6 | 57
M8 | B | 15 | 0.63 | -1.56 | 5.61 | -1.52 | 6.55 | -1.37 | 0.76 | -1.43 | 0.02701 | 0.03289 | 0.054 | 6 | 57
M8 | B | 16 | 0.76 | -1.83 | 5.41 | -1.5 | 6.43 | -1.29 | 0.78 | -1.5 | 0.02861 | 0.02908 | 0.054 | 6 | 57
M9 | A | 1 | 0.38 | -1.61 | 3.82 | -1.55 | 3.42 | -1.7 | 0.38 | -1.34 | ~ | ~ | ~ | ~ | ~
M9 | A | 2 | 0.36 | -1.74 | 3.78 | -1.61 | 3.82 | -1.66 | 0.29 | -1.27 | ~ | ~ | ~ | ~ | ~
M9 | A | 3 | 0.41 | -1.77 | 4.4 | -1.63 | 4.18 | -1.68 | 0.38 | -1.47 | ~ | ~ | ~ | ~ | ~
M9 | A | 4 | 0.55 | -2.19 | 4.82 | -1.66 | 5.08 | -1.6 | 0.36 | -1.43 | ~ | ~ | ~ | ~ | ~
M9 | A | 5 | 0.54 | -1.92 | 5.16 | -1.64 | 4.97 | -1.65 | 0.44 | -1.54 | ~ | ~ | ~ | ~ | ~
M9 | A | 6 | 0.5 | -1.92 | 4.94 | -1.67 | 5.29 | -1.57 | 0.35 | -1.28 | ~ | ~ | ~ | ~ | ~
M9 | A | 7 | 0.5 | -1.96 | 5.12 | -1.64 | 4.95 | -1.63 | 0.37 | -1.33 | ~ | ~ | ~ | ~ | ~
M9 | A | 8 | 0.55 | -1.94 | 5.19 | -1.64 | 5.59 | -1.54 | 0.37 | -1.23 | ~ | ~ | ~ | ~ | ~
M9 | A | 9 | 0.47 | -1.81 | 5.27 | -1.63 | 5.08 | -1.62 | 0.37 | -1.29 | ~ | ~ | ~ | ~ | ~
M9 | A | 10 | 0.51 | -1.87 | 5.18 | -1.64 | 5.4 | -1.59 | 0.38 | -1.35 | ~ | ~ | ~ | ~ | ~
M9 | A | 11 | 0.47 | -1.81 | 5.35 | -1.61 | 5.14 | -1.6 | 0.43 | -1.48 | ~ | ~ | ~ | ~ | ~
M9 | A | 12 | 0.57 | -1.86 | 5.3 | -1.61 | 5.42 | -1.59 | 0.41 | -1.27 | ~ | ~ | ~ | ~ | ~
M9 | A | 13 | 0.5 | -1.76 | 5.42 | -1.59 | 5.13 | -1.59 | 0.44 | -1.36 | ~ | ~ | ~ | ~ | ~
M9 | A | 14 | 0.58 | -1.94 | 5.22 | -1.57 | 5.18 | -1.6 | 0.4 | -1.28 | ~ | ~ | ~ | ~ | ~
M9 | A | 15 | 0.58 | -1.96 | 5.14 | -1.57 | 4.76 | -1.6 | 0.52 | -1.55 | ~ | ~ | ~ | ~ | ~
M9 | A | 16 | 0.51 | -1.92 | 4.94 | -1.52 | 4.82 | -1.55 | 0.43 | -1.52 | ~ | ~ | ~ | ~ | ~
M9 | B | 1 | 0.4 | -1.66 | 3.83 | -1.56 | 3.43 | -1.71 | 0.4 | -1.38 | ~ | ~ | ~ | ~ | ~
M9 | B | 2 | 0.39 | -1.84 | 3.76 | -1.64 | 3.78 | -1.71 | 0.31 | -1.32 | ~ | ~ | ~ | ~ | ~
M9 | B | 3 | 0.44 | -1.83 | 4.41 | -1.63 | 4.19 | -1.7 | 0.38 | -1.37 | ~ | ~ | ~ | ~ | ~
M9 | B | 4 | 0.53 | -2.06 | 4.69 | -1.78 | 5.09 | -1.62 | 0.38 | -1.54 | ~ | ~ | ~ | ~ | ~
M9 | B | 5 | 0.58 | -2 | 5.14 | -1.63 | 4.96 | -1.66 | 0.43 | -1.44 | ~ | ~ | ~ | ~ | ~
M9 | B | 6 | 0.49 | -1.8 | 4.91 | -1.67 | 5.27 | -1.59 | 0.36 | -1.28 | ~ | ~ | ~ | ~ | ~
M9 | B | 7 | 0.46 | -1.81 | 5.06 | -1.63 | 4.93 | -1.64 | 0.37 | -1.36 | ~ | ~ | ~ | ~ | ~
M9 | B | 8 | 0.54 | -1.88 | 5.13 | -1.64 | 5.55 | -1.55 | 0.39 | -1.3 | ~ | ~ | ~ | ~ | ~
M9 | B | 9 | 0.49 | -1.84 | 5.2 | -1.63 | 5.02 | -1.63 | 0.39 | -1.33 | ~ | ~ | ~ | ~ | ~
M9 | B | 10 | 0.5 | -1.83 | 5.11 | -1.64 | 5.35 | -1.61 | 0.37 | -1.28 | ~ | ~ | ~ | ~ | ~
M9 | B | 11 | 0.49 | -1.84 | 5.27 | -1.61 | 5.06 | -1.62 | 0.41 | -1.36 | ~ | ~ | ~ | ~ | ~
M9 | B | 12 | 0.6 | -1.96 | 5.21 | -1.6 | 5.36 | -1.6 | 0.42 | -1.26 | ~ | ~ | ~ | ~ | ~
M9 | B | 13 | 0.51 | -1.79 | 5.33 | -1.58 | 5.02 | -1.62 | 0.43 | -1.38 | ~ | ~ | ~ | ~ | ~
M9 | B | 14 | 0.58 | -2.02 | 5.13 | -1.57 | 5.1 | -1.62 | 0.39 | -1.25 | ~ | ~ | ~ | ~ | ~
M9 | B | 15 | 0.58 | -1.94 | 5.05 | -1.57 | 4.68 | -1.62 | 0.5 | -1.57 | ~ | ~ | ~ | ~ | ~
M9 | B | 16 | 0.5 | -1.82 | 4.84 | -1.52 | 4.64 | -1.61 | 0.4 | -1.35 | ~ | ~ | ~ | ~ | ~
M10 | A | 1 | 0.22 | -1.31 | 3.18 | -1.63 | 3.45 | -1.54 | 4.63 | -4.12 | 0.00852 | 0.0251 | 0.073 | 6 | 86.1
M10 | A | 2 | 0.22 | -1.34 | 3.51 | -1.63 | 3.8 | -1.5 | 3.3 | -3.79 | 0.01507 | 0.02784 | 0.073 | 6 | 86.1
M10 | A | 3 | 0.27 | -1.25 | 4.08 | -1.63 | 4.34 | -1.51 | 5.44 | -3.91 | 0.01066 | 0.03252 | 0.073 | 6 | 86.1
M10 | A | 4 | 0.27 | -1.26 | 4.98 | -1.59 | 5.09 | -1.54 | 3.61 | -3.51 | 0.02047 | 0.03865 | 0.073 | 6 | 86.1
M10 | A | 5 | 0.29 | -1.18 | 5.11 | -1.56 | 5.08 | -1.51 | 6.15 | -3.94 | 0.01241 | 0.03881 | 0.073 | 6 | 86.1
M10 | A | 6 | 0.28 | -1.3 | 5.08 | -1.57 | 5.13 | -1.55 | 3.07 | -3.27 | 0.02038 | 0.03975 | 0.073 | 6 | 86.1
M10 | A | 7 | 0.27 | -1.2 | 4.97 | -1.56 | 4.89 | -1.64 | 4.93 | -3.71 | 0.01177 | 0.03877 | 0.073 | 6 | 86.1
M10 | A | 8 | 0.29 | -1.38 | 5.18 | -1.55 | 5.25 | -1.53 | 2.78 | -3.19 | 0.02029 | 0.04137 | 0.073 | 6 | 86.1
M10 | A | 9 | 0.3 | -1.48 | 5.05 | -1.56 | 5.07 | -1.62 | 5.96 | -3.96 | 0.01197 | 0.04068 | 0.073 | 6 | 86.1
M10 | A | 10 | 0.28 | -1.39 | 5.09 | -1.57 | 5.23 | -1.51 | 2.75 | -3.22 | 0.01973 | 0.04171 | 0.073 | 6 | 86.1
M10 | A | 11 | 0.28 | -1.27 | 4.8 | -1.59 | 4.96 | -1.59 | 5.2 | -3.84 | 0.01131 | 0.04031 | 0.073 | 6 | 86.1
M10 | A | 12 | 0.26 | -1.14 | 4.84 | -1.6 | 5.14 | -1.48 | 2.47 | -3.07 | 0.01871 | 0.04142 | 0.073 | 6 | 86.1
M10 | A | 13 | 0.28 | -1.26 | 4.52 | -1.62 | 4.89 | -1.54 | 5.3 | -3.87 | 0.0107 | 0.04001 | 0.073 | 6 | 86.1
M10 | A | 14 | 0.26 | -1.17 | 4.67 | -1.6 | 5.21 | -1.41 | 2.79 | -3.27 | 0.01794 | 0.04218 | 0.073 | 6 | 86.1
M10 | A | 15 | 0.28 | -1.08 | 4.15 | -1.63 | 4.85 | -1.48 | 3.95 | -3.38 | 0.00991 | 0.03994 | 0.073 | 6 | 86.1
M10 | A | 16 | 0.29 | -1.08 | 4.24 | -1.59 | 5.04 | -1.37 | 2.12 | -2.71 | 0.01631 | 0.04106 | 0.073 | 6 | 86.1
M10 | B | 1 | 0.22 | -1.29 | 3.25 | -1.62 | 3.6 | -1.41 | 4.7 | -4.07 | 0.00895 | 0.02409 | 0.073 | 6 | 86.1
M10 | B | 2 | 0.22 | -1.29 | 3.55 | -1.62 | 3.85 | -1.51 | 3 | -3.63 | 0.01571 | 0.02656 | 0.073 | 6 | 86.1
M10 | B | 3 | 0.28 | -1.28 | 4.12 | -1.62 | 4.38 | -1.5 | 5.54 | -3.91 | 0.01113 | 0.03089 | 0.073 | 6 | 86.1
M10 | B | 4 | 0.28 | -1.29 | 5.02 | -1.57 | 5.12 | -1.55 | 3.89 | -3.56 | 0.02131 | 0.03678 | 0.073 | 6 | 86.1
M10 | B | 5 | 0.29 | -1.19 | 5.13 | -1.55 | 5.08 | -1.51 | 6.15 | -3.92 | 0.01288 | 0.03678 | 0.073 | 6 | 86.1
M10 | B | 6 | 0.29 | -1.34 | 5.07 | -1.56 | 5.15 | -1.49 | 3.26 | -3.36 | 0.02112 | 0.03748 | 0.073 | 6 | 86.1
M10 | B | 7 | 0.27 | -1.15 | 4.97 | -1.54 | 4.93 | -1.51 | 4.87 | -3.68 | 0.01216 | 0.03649 | 0.073 | 6 | 86.1
M10 | B | 8 | 0.3 | -1.43 | 5.14 | -1.55 | 5.24 | -1.46 | 2.71 | -3.13 | 0.02094 | 0.0389 | 0.073 | 6 | 86.1
M10 | B | 9 | 0.3 | -1.5 | 5.03 | -1.55 | 5.03 | -1.62 | 5.84 | -3.94 | 0.01238 | 0.03842 | 0.073 | 6 | 86.1
M10 | B | 10 | 0.28 | -1.38 | 5.03 | -1.56 | 5.21 | -1.44 | 2.66 | -3.16 | 0.0203 | 0.03917 | 0.073 | 6 | 86.1
M10 | B | 11 | 0.28 | -1.26 | 4.75 | -1.58 | 4.9 | -1.6 | 5.24 | -3.84 | 0.01165 | 0.03797 | 0.073 | 6 | 86.1
M10 | B | 12 | 0.26 | -1.15 | 4.72 | -1.61 | 5.07 | -1.49 | 2.42 | -3.03 | 0.01915 | 0.03908 | 0.073 | 6 | 86.1
M10 | B | 13 | 0.28 | -1.23 | 4.46 | -1.61 | 4.81 | -1.55 | 5.23 | -3.86 | 0.01103 | 0.03764 | 0.073 | 6 | 86.1
M10 | B | 14 | 0.25 | -1.15 | 4.58 | -1.6 | 5.12 | -1.43 | 2.75 | -3.26 | 0.01838 | 0.03975 | 0.073 | 6 | 86.1
M10 | B | 15 | 0.27 | -1.07 | 4.09 | -1.62 | 4.78 | -1.47 | 4.11 | -3.44 | 0.01018 | 0.03767 | 0.073 | 6 | 86.1
M10 | B | 16 | 0.28 | -1.07 | 4.14 | -1.59 | 4.94 | -1.39 | 2.08 | -2.71 | 0.0167 | 0.03869 | 0.073 | 6 | 86.1
M11 | A | 1 | 0.22 | -1.41 | 2.15 | -1.91 | 2.36 | -2.03 | 1.66 | -2.99 | 0.00019 | 0.00024 | 0.0012 | 6 | 1.2
M11 | A | 2 | 0.3 | -1.4 | 2.62 | -1.9 | 2.66 | -1.83 | 0.84 | -2.25 | 0.00027 | 0.00027 | 0.0012 | 6 | 1.2
M11 | A | 3 | 0.21 | -1.12 | 3.18 | -1.87 | 3.23 | -1.92 | 2.39 | -3.37 | 0.00027 | 0.00034 | 0.0012 | 6 | 1.2
M11 | A | 4 | 0.25 | -1.03 | 3.64 | -1.81 | 3.52 | -1.79 | 1.25 | -2.55 | 0.00036 | 0.00037 | 0.0012 | 6 | 1.2
M11 | A | 5 | 0.24 | -0.99 | 3.65 | -1.78 | 3.51 | -1.93 | 2.28 | -3.25 | 0.0003 | 0.00038 | 0.0012 | 6 | 1.2
M11 | A | 6 | 0.15 | -0.62 | 3.52 | -1.76 | 3.37 | -1.76 | 0.83 | -2.22 | 0.00034 | 0.00036 | 0.0012 | 6 | 1.2
M11 | A | 7 | 0.19 | -0.79 | 3.54 | -1.74 | 3.51 | -1.79 | 1.55 | -2.36 | 0.00028 | 0.00038 | 0.0012 | 6 | 1.2
M11 | A | 8 | 0.35 | -1.37 | 3.93 | -1.72 | 3.74 | -1.71 | 1.28 | -2.75 | 0.00037 | 0.0004 | 0.0012 | 6 | 1.2
M11 | A | 9 | 0.3 | -1.1 | 3.82 | -1.72 | 3.82 | -1.73 | 2.12 | -2.97 | 0.0003 | 0.00042 | 0.0012 | 6 | 1.2
M11 | A | 10 | 0.21 | -0.81 | 3.75 | -1.71 | 3.51 | -1.72 | 1.21 | -2.51 | 0.00034 | 0.00038 | 0.0012 | 6 | 1.2
M11 | A | 11 | 0.23 | -0.38 | 3.62 | -1.73 | 3.55 | -1.76 | 2.24 | -2.8 | 0.00027 | 0.0004 | 0.0012 | 6 | 1.2
M11 | A | 12 | 0.22 | -0.53 | 3.77 | -1.68 | 3.66 | -1.61 | 0.78 | -1.65 | 0.00034 | 0.0004 | 0.0012 | 6 | 1.2
M11 | A | 13 | 0.25 | -0.55 | 3.64 | -1.72 | 3.83 | -1.69 | 2.03 | -2.78 | 0.00028 | 0.00043 | 0.0012 | 6 | 1.2
M11 | A | 14 | 0.31 | -0.96 | 3.59 | -1.7 | 3.52 | -1.62 | 1.74 | -3.17 | 0.00033 | 0.00039 | 0.0012 | 6 | 1.2
M11 | A | 15 | 0.2 | -0.74 | 3.28 | -1.74 | 3.39 | -1.78 | 0.98 | -2.18 | 0.00025 | 0.00039 | 0.0012 | 6 | 1.2
M11 | A | 16 | 0.27 | -0.64 | 3.78 | -1.54 | 3.67 | -1.54 | 1.21 | -2.26 | 0.00033 | 0.00041 | 0.0012 | 6 | 1.2
M11 | B | 1 | 0.18 | -1.04 | 2.21 | -1.89 | 2.42 | -2.03 | 2.61 | -3.99 | 0.0002 | 0.00023 | 0.0012 | 6 | 1.2
M11 | B | 2 | 0.26 | -1.14 | 2.65 | -1.89 | 2.64 | -1.97 | 1.09 | -2.62 | 0.00028 | 0.00026 | 0.0012 | 6 | 1.2
M11 | B | 3 | 0.19 | -0.64 | 3.17 | -1.87 | 3.27 | -1.97 | 2.53 | -3.71 | 0.00028 | 0.00033 | 0.0012 | 6 | 1.2
M11 | B | 4 | 0.24 | -0.82 | 3.67 | -1.8 | 3.49 | -1.84 | 1.14 | -2.31 | 0.00037 | 0.00035 | 0.0012 | 6 | 1.2
M11 | B | 5 | 0.18 | -0.63 | 3.66 | -1.77 | 3.54 | -1.93 | 1.94 | -3.23 | 0.0003 | 0.00037 | 0.0012 | 6 | 1.2
M11 | B | 6 | 0.19 | -0.95 | 3.55 | -1.75 | 3.35 | -1.79 | 0.84 | -2.4 | 0.00035 | 0.00034 | 0.0012 | 6 | 1.2
M11 | B | 7 | 0.23 | -1.16 | 3.56 | -1.73 | 3.48 | -1.89 | 1.3 | -2.67 | 0.00029 | 0.00036 | 0.0012 | 6 | 1.2
M11 | B | 8 | 0.34 | -1.19 | 3.91 | -1.71 | 3.74 | -1.73 | 1.09 | -2.26 | 0.00038 | 0.00038 | 0.0012 | 6 | 1.2
M11 | B | 9 | 0.34 | -1.11 | 3.83 | -1.7 | 3.75 | -1.86 | 1.43 | -2.38 | 0.00031 | 0.0004 | 0.0012 | 6 | 1.2
M11 | B | 10 | 0.21 | -0.84 | 3.72 | -1.7 | 3.52 | -1.72 | 1.18 | -2.79 | 0.00035 | 0.00037 | 0.0012 | 6 | 1.2
M11 | B | 11 | 0.21 | -0.49 | 3.62 | -1.71 | 3.49 | -1.79 | 1.5 | -2.86 | 0.00028 | 0.00037 | 0.0012 | 6 | 1.2
M11 | B | 12 | 0.25 | -0.62 | 3.71 | -1.68 | 3.62 | -1.64 | 1.14 | -2.43 | 0.00035 | 0.00038 | 0.0012 | 6 | 1.2
M11 | B | 13 | 0.33 | -1 | 3.62 | -1.71 | 3.75 | -1.72 | 1.76 | -2.44 | 0.00029 | 0.0004 | 0.0012 | 6 | 1.2
M11 | B | 14 | 0.32 | -1.09 | 3.56 | -1.69 | 3.53 | -1.6 | 1.26 | -2.49 | 0.00034 | 0.00037 | 0.0012 | 6 | 1.2
M11 | B | 15 | 0.17 | -0.35 | 3.25 | -1.72 | 3.33 | -1.81 | 0.84 | -1.7 | 0.00026 | 0.00036 | 0.0012 | 6 | 1.2
M11 | B | 16 | 0.34 | -1.13 | 3.76 | -1.52 | 3.56 | -1.56 | 1.33 | -2.74 | 0.00034 | 0.00038 | 0.0012 | 6 | 1.2

## DNB NFR
Band | HAM | Band Pass Filter | Baffle locations |  | Mean | [7 empty] | Median | [7 empty] | Sensor Performance Compared with Requirement
[5 empty] | Harvy Shack Fit Parameters | [7 empty] | Harvy Shack Fit Parameters | [7 empty] | Mean |  | Median |  | Max Allowed Lscat | Angle (mrad) from Target | L Bright Target
|  |  | Left | Right | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Scattered Radiance
[5 empty] | A | M | A | M | A | M | A | M | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright | Left of Lbright | Right of Lbright
DNBLGS | A | M6 | -71 | 3 | 13.5312 | -1.3506 | 10.2794 | -2.0319 | 0.77687 | -3.2481 | 1.2812 | -0.9 | 6.415 | -1.545 | 8.355 | -2.14 | 0.235 | -3.31 | 1.19 | -0.85 | 0.00087875 | 0.0001475 | 0.00075 | 0.00013 | 0.002 | 6 | 2
DNBLGS | B | M6 | -71 | 3 | 4.7106 | -0.16187 | 9.115 | -2.0919 | 2.4169 | -2.8044 | 1.1475 | -0.85187 | 1.145 | -0.015 | 6.945 | -2.225 | 0.745 | -2.87 | 1.125 | -0.81 | 0.0008575 | 0.00012938 | 0.0007 | 0.000125 | 0.002 | 6 | 2
DNBLGS | A | M5 | -71 | 3 | 8.9675 | 0.14938 | 9.0644 | -1.8169 | 0.6425 | -3.3237 | 1.0694 | -1.0244 | 1.115 | -0.21 | 7.595 | -1.89 | 0.195 | -3.395 | 0.965 | -0.995 | 0.00075625 | 0.00012375 | 0.00065 | 0.00011 | 0.002 | 6 | 2
DNBLGS | B | M5 | -71 | 3 | 10.2469 | -0.58063 | 8.2044 | -1.7994 | 1.9106 | -2.895 | 0.965 | -1.065 | 0.83 | 0.135 | 6.74 | -1.875 | 0.575 | -2.97 | 0.935 | -1.06 | 0.00075 | 0.00011 | 0.00065 | 0.000105 | 0.002 | 6 | 2
DNBMGS | A | M6 | -67 | 7 | 1.1663 | -2.1306 | 8.7956 | -1.6331 | 0.63625 | -2.4325 | 0.53687 | -1.8069 | 1.21 | -2.145 | 8.935 | -1.635 | 0.65 | -2.41 | 0.55 | -1.81 | 0.00066438 | 7.125e-05 | 0.00067 | 7e-05 | 0.002 | 6 | 2
DNBMGS | B | M6 | -67 | 7 | 1.0413 | -2.1119 | 7.6025 | -1.6481 | 0.69375 | -2.3781 | 0.49125 | -1.79 | 1.015 | -2.11 | 7.715 | -1.67 | 0.745 | -2.325 | 0.5 | -1.79 | 0.0006425 | 6.5e-05 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | A | M5 | -67 | 7 | 0.9825 | -1.9944 | 9.4494 | -1.5844 | 0.64563 | -2.4412 | 0.51688 | -1.8675 | 1.02 | -2 | 9.695 | -1.57 | 0.7 | -2.405 | 0.54 | -1.87 | 0.00071375 | 6.6875e-05 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBHGA | B | M5 | -67 | 7 | 0.88875 | -2.0019 | 8.1594 | -1.62 | 0.68688 | -2.4 | 0.46563 | -1.8656 | 0.895 | -2.005 | 8.355 | -1.625 | 0.76 | -2.345 | 0.485 | -1.87 | 0.00069437 | 6.3125e-05 | 0.000695 | 6e-05 | 0.002 | 6 | 2
DNBHGA | A | M6 | -45 | 34 | 1.2881 | -0.805 | 6.335 | -1.8519 | 9.2644 | -1.4594 | 0.62687 | -0.49063 | 1.17 | -0.95 | 6.36 | -1.845 | 9.19 | -1.46 | 0.59 | -0.435 | 0.00059813 | 0.00094562 | 0.000605 | 0.000935 | 0.002 | 6 | 2
DNBHGA | B | M6 | -45 | 34 | 1.3125 | -1.0344 | 5.665 | -1.8475 | 8.3287 | -1.4787 | 0.69937 | -0.74062 | 1.175 | -0.915 | 5.56 | -1.845 | 8.445 | -1.48 | 0.67 | -0.71 | 0.00058375 | 0.0008475 | 0.00058 | 0.000865 | 0.002 | 6 | 2
DNBHGB | A | M6 | -59 | 20 | 3.3644 | -1.3462 | 8.4106 | -1.6294 | 11.8413 | -1.4262 | 0.94 | -1.1294 | 2.51 | -1.285 | 8.22 | -1.63 | 11.735 | -1.42 | 0.905 | -1.145 | 0.00075187 | 0.0001125 | 0.000745 | 0.00011 | 0.002 | 6 | 2
DNBHGB | B | M6 | -59 | 20 | 2.05 | -1.0425 | 7.4869 | -1.6406 | 10.5281 | -1.4575 | 0.83188 | -1.2019 | 1.34 | -0.8 | 7.455 | -1.65 | 10.565 | -1.455 | 0.82 | -1.165 | 0.0007325 | 9.75e-05 | 0.00073 | 0.0001 | 0.002 | 6 | 2
DNBHGA | A | M5 | -45 | 34 | 1.0537 | -1.7775 | 8.2744 | -1.6994 | 11.2931 | -1.285 | 1.4781 | -2.5206 | 1.065 | -1.745 | 8.365 | -1.7 | 11.325 | -1.28 | 1.43 | -2.51 | 0.00076187 | 0.0011237 | 0.00076 | 0.001125 | 0.002 | 6 | 2
DNBHGA | B | M5 | -45 | 34 | 1.26 | -1.8987 | 7.4731 | -1.7044 | 10.255 | -1.3094 | 1.58 | -2.4125 | 1.08 | -1.77 | 7.425 | -1.71 | 10.25 | -1.31 | 1.405 | -2.36 | 0.00074813 | 0.0010181 | 0.00075 | 0.001015 | 0.002 | 6 | 2
DNBHGB | A | M5 | -59 | 20 | 4.1212 | -2.7744 | 9.6881 | -1.4381 | 12.775 | -1.3213 | 1.2587 | -2.5475 | 2.5 | -2.535 | 9.645 | -1.44 | 12.735 | -1.315 | 1.215 | -2.525 | 0.00083937 | 0.000185 | 0.00084 | 0.000185 | 0.002 | 6 | 2
DNBHGB | B | M5 | -59 | 20 | 3.38 | -2.59 | 8.8019 | -1.4344 | 11.5619 | -1.3494 | 1.2794 | -2.5319 | 2.285 | -2.29 | 8.76 | -1.44 | 11.54 | -1.34 | 1.245 | -2.5 | 0.00082875 | 0.000185 | 0.00083 | 0.00018 | 0.002 | 6 | 2
Band | Detector | HAM | Band Pass Filter | Harvy Shack Fit Parameters | [7 empty] | Sensor Performance Compared to Requirement3 | [5 empty] | Notes:
|  |  |  | Outside Left Baffle |  | Inside Left Baffle |  | Inside Right Baffle |  | Outside Right Baffle |  | Scattered Radiance |  | Max Allowed Lscat | Angle (mrad) from Target | Bright Target4 |  | 1. This table provides VIIRS F2 DNB NFR.
|  |  |  | A | M | A | M | A | M | A | M | Left of Lbright | Right of Lbright |  |  |  |  | 2. The results were derived from FP-14 using the ScMA.
DNBLGS | 1 | A | M6 | 17.18 | -3.09 | 6.6 | -2.19 | 2.31 | -2.39 | 0.55 | -1.19 | 0.00069 | 7e-05 | 0.002 | 6 | 2 |  | 3. Linear fits were peformed of log(dn/(s*t))= log(A) + M*Log(sin(theta)/0.01))
DNBLGS | 2 | A | M6 | 2.15 | -0.3 | 6.91 | -2.14 | 2.48 | -2.38 | 1.04 | -0.83 | 0.0007 | 0.00012 | 0.002 | 6 | 2 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
DNBLGS | 3 | A | M6 | 10.38 | -2.54 | 7.36 | -2.21 | 1.47 | -2.59 | 1.03 | -0.79 | 0.00074 | 0.00012 | 0.002 | 6 | 2 |  | 5. The test was conducted in diagnostic mode.
DNBLGS | 4 | A | M6 | 0.94 | 0.15 | 6.97 | -2.18 | 1.04 | -2.73 | 1.2 | -0.69 | 0.00069 | 0.00013 | 0.002 | 6 | 2 |  | 6. Radiance units are W m-2 sr-1
DNBLGS | 5 | A | M6 | 0.18 | 1.97 | 7.98 | -2.21 | 1.75 | -2.62 | 1.01 | -0.98 | 0.00077 | 0.00012 | 0.002 | 6 | 2 |  | 7. The results using M5 and M6 band pass filters are reported, as they are closest to the DNB center wavelenth
DNBLGS | 6 | A | M6 | 41.99 | -3.56 | 23.94 | -1.12 | 2.19 | -2.48 | 2.1 | -1.14 | 0.00189 | 0.00025 | 0.002 | 6 | 2 |  | 8. Assume bright target = 2, DNB NFR requirement is half of L_max.
DNBLGS | 7 | A | M6 | 2.45 | -0.65 | 8.29 | -2.14 | 0.32 | -3.16 | 1.52 | -0.83 | 0.00077 | 0.00017 | 0.002 | 6 | 2
DNBLGS | 8 | A | M6 | 0.23 | 1.58 | 7.76 | -2.23 | 0.32 | -3.17 | 1.32 | -0.84 | 0.00071 | 0.00015 | 0.002 | 6 | 2
DNBLGS | 9 | A | M6 | 10.55 | -2.61 | 8.63 | -2.02 | 0.15 | -3.45 | 1.13 | -0.7 | 0.00075 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 10 | A | M6 | 29.6 | -3.16 | 7.95 | -2.29 | 0.07 | -3.75 | 1.25 | -0.79 | 0.00071 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 11 | A | M6 | 30.87 | -2.92 | 26.41 | -1.07 | 0.07 | -3.72 | 2.22 | -1.21 | 0.0019 | 0.00026 | 0.002 | 6 | 2
DNBLGS | 12 | A | M6 | 2.14 | -0.97 | 8.76 | -2.26 | 0.09 | -3.67 | 1.41 | -0.88 | 0.00076 | 0.00016 | 0.002 | 6 | 2
DNBLGS | 13 | A | M6 | 50.73 | -3.71 | 9.09 | -2.09 | 0.08 | -3.7 | 1.18 | -0.82 | 0.00075 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 14 | A | M6 | 0.6 | 0.66 | 9.49 | -2.14 | 0.04 | -3.92 | 1.38 | -0.86 | 0.00077 | 0.00016 | 0.002 | 6 | 2
DNBLGS | 15 | A | M6 | 15.02 | -2.12 | 8.42 | -2.14 | 0.02 | -4.17 | 1.12 | -0.94 | 0.00068 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 16 | A | M6 | 1.49 | -0.34 | 9.91 | -2.08 | 0.03 | -4.07 | 1.04 | -0.91 | 0.00078 | 0.00012 | 0.002 | 6 | 2
DNBLGS | 1 | B | M6 | 0.23 | 1.48 | 5.88 | -2.15 | 6.99 | -1.96 | 0.93 | -0.87 | 0.00065 | 0.0001 | 0.002 | 6 | 2
DNBLGS | 2 | B | M6 | 3.79 | -1.2 | 6.12 | -2.23 | 7.83 | -1.94 | 1.24 | -0.57 | 0.00068 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 3 | B | M6 | 3.96 | -1.29 | 6.63 | -2.18 | 4.7 | -2.14 | 0.78 | -0.85 | 0.00071 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 4 | B | M6 | 0.38 | 0.94 | 6.47 | -2.25 | 3.32 | -2.28 | 0.79 | -0.9 | 0.00069 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 5 | B | M6 | 9.08 | -2.23 | 8.04 | -2.06 | 5.69 | -2.16 | 0.78 | -0.78 | 0.00082 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 6 | B | M6 | 1.78 | -0.19 | 21.76 | -1.1 | 6.29 | -2.07 | 1.9 | -1.11 | 0.00186 | 0.00022 | 0.002 | 6 | 2
DNBLGS | 7 | B | M6 | 0.25 | 1.19 | 6.9 | -2.26 | 1.03 | -2.71 | 0.97 | -0.8 | 0.00071 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 8 | B | M6 | 1.04 | -0.11 | 6.6 | -2.36 | 1.06 | -2.71 | 1.25 | -0.73 | 0.00068 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 9 | B | M6 | 41.89 | -3.86 | 6.32 | -2.39 | 0.46 | -3.03 | 0.82 | -0.82 | 0.00064 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 10 | B | M6 | 0.29 | 1.02 | 6.99 | -2.34 | 0.21 | -3.31 | 0.87 | -0.8 | 0.00069 | 0.0001 | 0.002 | 6 | 2
DNBLGS | 11 | B | M6 | 1.01 | 0.21 | 24.61 | -1.12 | 0.23 | -3.29 | 1.7 | -1.32 | 0.00195 | 0.0002 | 0.002 | 6 | 2
DNBLGS | 12 | B | M6 | 0.05 | 3.22 | 9.95 | -2 | 0.29 | -3.22 | 1.58 | -0.79 | 0.00091 | 0.00018 | 0.002 | 6 | 2
DNBLGS | 13 | B | M6 | 0.94 | 0.08 | 8.01 | -2.22 | 0.27 | -3.25 | 1.24 | -0.87 | 0.00075 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 14 | B | M6 | 1.46 | -0.21 | 7.79 | -2.22 | 0.14 | -3.47 | 1.06 | -0.93 | 0.00072 | 0.00012 | 0.002 | 6 | 2
DNBLGS | 15 | B | M6 | 1.25 | 0.22 | 6.34 | -2.36 | 0.07 | -3.71 | 1.26 | -0.77 | 0.00059 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 16 | B | M6 | 7.97 | -1.86 | 7.43 | -2.23 | 0.09 | -3.62 | 1.19 | -0.72 | 0.00067 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 1 | A | M5 | 0.05 | 2.44 | 6.08 | -1.84 | 2.01 | -2.45 | 0.6 | -1.15 | 0.0006 | 7e-05 | 0.002 | 6 | 2
DNBLGS | 2 | A | M5 | 0.75 | -0.28 | 6.71 | -1.94 | 2.12 | -2.45 | 0.77 | -1.16 | 0.00066 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 3 | A | M5 | 0.13 | 2.16 | 7.16 | -1.89 | 1.28 | -2.64 | 0.66 | -0.91 | 0.00069 | 8e-05 | 0.002 | 6 | 2
DNBLGS | 4 | A | M5 | 2.08 | -1.04 | 6.69 | -1.97 | 0.86 | -2.81 | 0.93 | -0.93 | 0.00064 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 5 | A | M5 | 11.02 | -2.62 | 7.5 | -1.89 | 1.44 | -2.69 | 0.93 | -1.01 | 0.00069 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 6 | A | M5 | 3.31 | -0.61 | 20.2 | -0.96 | 1.57 | -2.61 | 1.5 | -1.29 | 0.00156 | 0.00018 | 0.002 | 6 | 2
DNBLGS | 7 | A | M5 | 0.02 | 3.69 | 7.69 | -1.75 | 0.28 | -3.21 | 0.62 | -1.13 | 0.00067 | 7e-05 | 0.002 | 6 | 2
DNBLGS | 8 | A | M5 | 0.43 | 0.95 | 6.34 | -2.2 | 0.27 | -3.23 | 0.75 | -1.17 | 0.00058 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 9 | A | M5 | 5.21 | -1.55 | 7.06 | -1.94 | 0.12 | -3.56 | 0.79 | -1.17 | 0.00061 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 10 | A | M5 | 0 | 5.67 | 7.72 | -1.89 | 0.06 | -3.82 | 1 | -0.98 | 0.00065 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 11 | A | M5 | 4.95 | -1.21 | 22.72 | -0.91 | 0.06 | -3.82 | 1.83 | -1.31 | 0.00159 | 0.00022 | 0.002 | 6 | 2
DNBLGS | 12 | A | M5 | 108.57 | -4.51 | 9.5 | -1.76 | 0.07 | -3.75 | 1.4 | -0.82 | 0.00076 | 0.00016 | 0.002 | 6 | 2
DNBLGS | 13 | A | M5 | 1.21 | -0.14 | 7.73 | -1.94 | 0.06 | -3.8 | 1.65 | -0.74 | 0.00063 | 0.00018 | 0.002 | 6 | 2
DNBLGS | 14 | A | M5 | 1.02 | 0.24 | 8.15 | -1.89 | 0.04 | -4 | 1.41 | -0.8 | 0.00065 | 0.00016 | 0.002 | 6 | 2
DNBLGS | 15 | A | M5 | 0.66 | 0.48 | 7.87 | -1.95 | 0.02 | -4.2 | 1.08 | -0.85 | 0.00062 | 0.00012 | 0.002 | 6 | 2
DNBLGS | 16 | A | M5 | 4.07 | -1.28 | 5.91 | -2.35 | 0.02 | -4.14 | 1.19 | -0.97 | 0.0005 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 1 | B | M5 | 54.27 | -4.51 | 5.07 | -2 | 5.82 | -2.04 | 0.72 | -1.14 | 0.00055 | 8e-05 | 0.002 | 6 | 2
DNBLGS | 2 | B | M5 | 0.06 | 2.15 | 6.25 | -1.8 | 6.21 | -2.04 | 0.79 | -1.08 | 0.00065 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 3 | B | M5 | 1.15 | -0.19 | 6.15 | -2.03 | 3.6 | -2.24 | 0.74 | -1.2 | 0.00065 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 4 | B | M5 | 0.09 | 2.16 | 5.86 | -2.08 | 2.75 | -2.36 | 0.54 | -1.33 | 0.00062 | 6e-05 | 0.002 | 6 | 2
DNBLGS | 5 | B | M5 | 24.13 | -3.57 | 7.32 | -1.79 | 4.6 | -2.23 | 0.76 | -1.11 | 0.00072 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 6 | B | M5 | 0.36 | 1.36 | 18.48 | -0.95 | 4.57 | -2.19 | 1.16 | -1.5 | 0.00155 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 7 | B | M5 | 0.51 | 0.46 | 5.64 | -2.19 | 0.88 | -2.78 | 0.98 | -0.9 | 0.00058 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 8 | B | M5 | 0.49 | 0.61 | 7 | -1.84 | 0.79 | -2.82 | 0.7 | -1.04 | 0.00066 | 8e-05 | 0.002 | 6 | 2
DNBLGS | 9 | B | M5 | 1.94 | -0.71 | 6.6 | -1.92 | 0.36 | -3.12 | 1.02 | -0.75 | 0.00063 | 0.00011 | 0.002 | 6 | 2
DNBLGS | 10 | B | M5 | 0.05 | 2.66 | 5.91 | -2.09 | 0.17 | -3.39 | 1.1 | -0.87 | 0.00057 | 0.00012 | 0.002 | 6 | 2
DNBLGS | 11 | B | M5 | 45.21 | -4.12 | 19.89 | -1.02 | 0.16 | -3.41 | 1.79 | -1.28 | 0.00157 | 0.00021 | 0.002 | 6 | 2
DNBLGS | 12 | B | M5 | 17.43 | -3.15 | 8.1 | -1.77 | 0.22 | -3.32 | 1.18 | -0.88 | 0.00072 | 0.00013 | 0.002 | 6 | 2
DNBLGS | 13 | B | M5 | 8.71 | -1.74 | 7.65 | -1.74 | 0.19 | -3.37 | 1.03 | -0.98 | 0.00067 | 0.00012 | 0.002 | 6 | 2
DNBLGS | 14 | B | M5 | 0.46 | 0.89 | 7.87 | -1.73 | 0.11 | -3.55 | 1.26 | -0.74 | 0.00068 | 0.00014 | 0.002 | 6 | 2
DNBLGS | 15 | B | M5 | 8.62 | -2.1 | 6.64 | -1.91 | 0.06 | -3.77 | 0.78 | -1.24 | 0.00058 | 9e-05 | 0.002 | 6 | 2
DNBLGS | 16 | B | M5 | 0.47 | 0.51 | 6.84 | -1.93 | 0.08 | -3.69 | 0.89 | -1 | 0.0006 | 0.0001 | 0.002 | 6 | 2
DNBMGS | 1 | A | M6 | 0.59 | -1.7 | 6.91 | -1.71 | 0.64 | -2.41 | 0.43 | -1.8 | 0.00062 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 2 | A | M6 | 0.92 | -2.06 | 7.42 | -1.69 | 0.77 | -2.32 | 0.47 | -1.81 | 0.00065 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 3 | A | M6 | 1.01 | -2.15 | 7.63 | -1.69 | 0.76 | -2.33 | 0.48 | -1.81 | 0.00065 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 4 | A | M6 | 1.35 | -2.43 | 7.76 | -1.66 | 0.72 | -2.36 | 0.49 | -1.81 | 0.00064 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 5 | A | M6 | 0.9 | -1.98 | 7.97 | -1.65 | 0.69 | -2.37 | 0.5 | -1.8 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 6 | A | M6 | 1 | -2.06 | 8.24 | -1.65 | 0.72 | -2.36 | 0.52 | -1.78 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 7 | A | M6 | 0.88 | -1.89 | 8.55 | -1.65 | 0.75 | -2.36 | 0.54 | -1.79 | 0.00066 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 8 | A | M6 | 1.38 | -2.31 | 8.94 | -1.64 | 0.73 | -2.38 | 0.55 | -1.78 | 0.00068 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 9 | A | M6 | 1.25 | -2.21 | 8.93 | -1.63 | 0.66 | -2.41 | 0.55 | -1.82 | 0.00067 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 10 | A | M6 | 1.17 | -2.13 | 9.2 | -1.62 | 0.62 | -2.44 | 0.56 | -1.82 | 0.00067 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 11 | A | M6 | 1.58 | -2.42 | 9.52 | -1.61 | 0.59 | -2.47 | 0.58 | -1.8 | 0.00068 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 12 | A | M6 | 1.27 | -2.14 | 9.71 | -1.61 | 0.56 | -2.5 | 0.58 | -1.85 | 0.00068 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 13 | A | M6 | 1.13 | -1.98 | 9.86 | -1.6 | 0.54 | -2.52 | 0.59 | -1.82 | 0.00068 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 14 | A | M6 | 1.39 | -2.19 | 10.04 | -1.59 | 0.51 | -2.54 | 0.59 | -1.81 | 0.00068 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 15 | A | M6 | 1.48 | -2.26 | 10.05 | -1.57 | 0.48 | -2.56 | 0.59 | -1.8 | 0.00068 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 16 | A | M6 | 1.36 | -2.18 | 10 | -1.56 | 0.44 | -2.59 | 0.57 | -1.81 | 0.00069 | 8e-05 | 0.002 | 6 | 2
DNBMGS | 1 | B | M6 | 1.01 | -2.35 | 6.11 | -1.7 | 0.87 | -2.27 | 0.41 | -1.79 | 0.0006 | 5e-05 | 0.002 | 6 | 2
DNBMGS | 2 | B | M6 | 0.82 | -2.01 | 6.52 | -1.69 | 0.95 | -2.22 | 0.44 | -1.78 | 0.00062 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 3 | B | M6 | 1.02 | -2.23 | 6.67 | -1.69 | 0.92 | -2.23 | 0.45 | -1.8 | 0.00063 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 4 | B | M6 | 0.72 | -1.86 | 6.81 | -1.69 | 0.84 | -2.27 | 0.45 | -1.77 | 0.00063 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 5 | B | M6 | 0.93 | -2.09 | 6.96 | -1.69 | 0.78 | -2.3 | 0.46 | -1.81 | 0.00063 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 6 | B | M6 | 1.07 | -2.24 | 7.16 | -1.69 | 0.82 | -2.28 | 0.47 | -1.77 | 0.00064 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 7 | B | M6 | 0.93 | -2.04 | 7.41 | -1.68 | 0.85 | -2.28 | 0.49 | -1.77 | 0.00065 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 8 | B | M6 | 0.94 | -2.01 | 7.72 | -1.69 | 0.81 | -2.3 | 0.5 | -1.78 | 0.00066 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 9 | B | M6 | 0.82 | -1.87 | 7.71 | -1.66 | 0.71 | -2.35 | 0.5 | -1.81 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 10 | B | M6 | 0.83 | -1.85 | 7.88 | -1.65 | 0.64 | -2.4 | 0.51 | -1.79 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 11 | B | M6 | 1.23 | -2.25 | 8.11 | -1.61 | 0.57 | -2.45 | 0.52 | -1.79 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 12 | B | M6 | 1.15 | -2.15 | 8.24 | -1.61 | 0.52 | -2.49 | 0.53 | -1.8 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 13 | B | M6 | 1.14 | -2.13 | 8.36 | -1.6 | 0.5 | -2.51 | 0.53 | -1.79 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 14 | B | M6 | 1.27 | -2.21 | 8.49 | -1.59 | 0.47 | -2.54 | 0.53 | -1.83 | 0.00065 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 15 | B | M6 | 1.16 | -2.07 | 8.64 | -1.57 | 0.44 | -2.56 | 0.54 | -1.77 | 0.00066 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 16 | B | M6 | 1.62 | -2.43 | 8.85 | -1.56 | 0.41 | -2.6 | 0.53 | -1.79 | 0.00066 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 1 | A | M5 | 0.64 | -1.81 | 7.6 | -1.71 | 0.78 | -2.32 | 0.42 | -1.87 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 2 | A | M5 | 0.68 | -1.8 | 8.02 | -1.71 | 0.77 | -2.33 | 0.45 | -1.85 | 0.00071 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 3 | A | M5 | 0.79 | -1.92 | 8.16 | -1.59 | 0.75 | -2.35 | 0.46 | -1.89 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 4 | A | M5 | 0.88 | -2 | 8.38 | -1.59 | 0.71 | -2.38 | 0.47 | -1.87 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 5 | A | M5 | 0.79 | -1.88 | 8.65 | -1.59 | 0.72 | -2.37 | 0.48 | -1.87 | 0.0007 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 6 | A | M5 | 0.79 | -1.86 | 8.92 | -1.58 | 0.74 | -2.37 | 0.5 | -1.86 | 0.00071 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 7 | A | M5 | 0.98 | -2.04 | 9.2 | -1.58 | 0.78 | -2.35 | 0.51 | -1.87 | 0.00071 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 8 | A | M5 | 0.98 | -2 | 9.58 | -1.58 | 0.73 | -2.39 | 0.53 | -1.85 | 0.00073 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 9 | A | M5 | 1.17 | -2.17 | 9.81 | -1.56 | 0.69 | -2.42 | 0.55 | -1.86 | 0.00073 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 10 | A | M5 | 1.06 | -2.05 | 9.99 | -1.56 | 0.67 | -2.42 | 0.55 | -1.9 | 0.00073 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 11 | A | M5 | 1.09 | -2.06 | 10.11 | -1.56 | 0.57 | -2.5 | 0.56 | -1.86 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 12 | A | M5 | 1.11 | -2.05 | 10.22 | -1.56 | 0.53 | -2.53 | 0.55 | -1.87 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 13 | A | M5 | 1.28 | -2.18 | 10.37 | -1.56 | 0.49 | -2.57 | 0.56 | -1.87 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 14 | A | M5 | 1.06 | -1.96 | 10.56 | -1.55 | 0.5 | -2.55 | 0.56 | -1.87 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 15 | A | M5 | 1.11 | -1.96 | 10.82 | -1.54 | 0.47 | -2.59 | 0.57 | -1.85 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 16 | A | M5 | 1.31 | -2.17 | 10.8 | -1.53 | 0.43 | -2.62 | 0.55 | -1.87 | 0.00073 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 1 | B | M5 | 0.69 | -1.96 | 6.72 | -1.69 | 0.99 | -2.21 | 0.39 | -1.88 | 0.00066 | 5e-05 | 0.002 | 6 | 2
DNBMGS | 2 | B | M5 | 0.7 | -1.9 | 7.06 | -1.69 | 0.92 | -2.25 | 0.42 | -1.86 | 0.00068 | 5e-05 | 0.002 | 6 | 2
DNBMGS | 3 | B | M5 | 0.79 | -2.02 | 7.22 | -1.69 | 0.86 | -2.28 | 0.42 | -1.88 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 4 | B | M5 | 0.9 | -2.14 | 7.38 | -1.7 | 0.83 | -2.29 | 0.43 | -1.86 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 5 | B | M5 | 0.85 | -2.06 | 7.57 | -1.7 | 0.82 | -2.29 | 0.44 | -1.85 | 0.00069 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 6 | B | M5 | 0.95 | -2.14 | 7.78 | -1.7 | 0.78 | -2.33 | 0.45 | -1.87 | 0.0007 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 7 | B | M5 | 0.73 | -1.81 | 8.02 | -1.7 | 0.82 | -2.31 | 0.46 | -1.85 | 0.00071 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 8 | B | M5 | 0.9 | -2.01 | 8.29 | -1.69 | 0.82 | -2.31 | 0.48 | -1.85 | 0.00072 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 9 | B | M5 | 0.98 | -2.1 | 8.42 | -1.56 | 0.74 | -2.36 | 0.49 | -1.88 | 0.0007 | 6e-05 | 0.002 | 6 | 2
DNBMGS | 10 | B | M5 | 0.88 | -1.94 | 8.54 | -1.56 | 0.63 | -2.43 | 0.49 | -1.88 | 0.0007 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 11 | B | M5 | 0.96 | -2.03 | 8.62 | -1.55 | 0.56 | -2.47 | 0.49 | -1.87 | 0.0007 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 12 | B | M5 | 0.92 | -1.96 | 8.69 | -1.55 | 0.5 | -2.52 | 0.5 | -1.85 | 0.00069 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 13 | B | M5 | 0.89 | -1.92 | 8.79 | -1.55 | 0.48 | -2.53 | 0.49 | -1.88 | 0.00069 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 14 | B | M5 | 0.86 | -1.87 | 8.96 | -1.54 | 0.44 | -2.58 | 0.5 | -1.87 | 0.00069 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 15 | B | M5 | 1.02 | -2 | 9.14 | -1.53 | 0.42 | -2.6 | 0.5 | -1.85 | 0.0007 | 7e-05 | 0.002 | 6 | 2
DNBMGS | 16 | B | M5 | 1.2 | -2.17 | 9.35 | -1.52 | 0.38 | -2.64 | 0.5 | -1.87 | 0.0007 | 7e-05 | 0.002 | 6 | 2
DNBHGA | 1 | A | M6 | 1.06 | -0.96 | 5.27 | -1.9 | 8.09 | -1.52 | 1.09 | -1.29 | 0.00055 | 0.00084 | 0.002 | 6 | 2
DNBHGA | 2 | A | M6 | 0.88 | -0.63 | 5.87 | -1.82 | 8.41 | -1.51 | 0.49 | -1.1 | 0.0006 | 0.00086 | 0.002 | 6 | 2
DNBHGA | 3 | A | M6 | 2.52 | -1.96 | 5.71 | -1.89 | 8.71 | -1.49 | 0.84 | -0.95 | 0.00058 | 0.00089 | 0.002 | 6 | 2
DNBHGA | 4 | A | M6 | 0.79 | -0.94 | 5.74 | -1.88 | 8.54 | -1.5 | 0.37 | -0.06 | 0.00058 | 0.00087 | 0.002 | 6 | 2
DNBHGA | 5 | A | M6 | 0.71 | -0.45 | 5.6 | -1.93 | 9 | -1.47 | 0.66 | -0.8 | 0.00056 | 0.00092 | 0.002 | 6 | 2
DNBHGA | 6 | A | M6 | 0.61 | -0.41 | 6.01 | -1.89 | 8.65 | -1.49 | 0.59 | -0.9 | 0.00059 | 0.00089 | 0.002 | 6 | 2
DNBHGA | 7 | A | M6 | 0.59 | 0.02 | 6.39 | -1.82 | 9.16 | -1.45 | 0.57 | -0.22 | 0.00061 | 0.00093 | 0.002 | 6 | 2
DNBHGA | 8 | A | M6 | 0.26 | 1.11 | 6.33 | -1.83 | 9.24 | -1.44 | 0.56 | -0.38 | 0.0006 | 0.00094 | 0.002 | 6 | 2
DNBHGA | 9 | A | M6 | 1.51 | -1.43 | 6.2 | -1.86 | 9.48 | -1.43 | 0.55 | -0.79 | 0.00058 | 0.00096 | 0.002 | 6 | 2
DNBHGA | 10 | A | M6 | 1.28 | -0.97 | 6.62 | -1.83 | 8.99 | -1.48 | 0.59 | -0.52 | 0.00061 | 0.00092 | 0.002 | 6 | 2
DNBHGA | 11 | A | M6 | 1.67 | -1.15 | 6.71 | -1.85 | 9.74 | -1.44 | 0.48 | 0.1 | 0.00061 | 0.00099 | 0.002 | 6 | 2
DNBHGA | 12 | A | M6 | 0.75 | -0.31 | 6.9 | -1.82 | 9.22 | -1.47 | 0.64 | -0.23 | 0.00062 | 0.00095 | 0.002 | 6 | 2
DNBHGA | 13 | A | M6 | 2.91 | -1.63 | 7.09 | -1.8 | 10.24 | -1.4 | 0.67 | -0.22 | 0.00063 | 0.00104 | 0.002 | 6 | 2
DNBHGA | 14 | A | M6 | 2.26 | -1.39 | 6.83 | -1.84 | 10.22 | -1.42 | 0.77 | -0.49 | 0.00061 | 0.00104 | 0.002 | 6 | 2
DNBHGA | 15 | A | M6 | 1.43 | -0.97 | 7.03 | -1.82 | 10.52 | -1.39 | 0.71 | -0.19 | 0.00062 | 0.00107 | 0.002 | 6 | 2
DNBHGA | 16 | A | M6 | 1.38 | -0.81 | 7.06 | -1.85 | 10.02 | -1.45 | 0.45 | 0.19 | 0.00062 | 0.00102 | 0.002 | 6 | 2
DNBHGA | 1 | B | M6 | 1.13 | -1.45 | 5.31 | -1.81 | 7.68 | -1.52 | 1.09 | -1.95 | 0.00059 | 0.00078 | 0.002 | 6 | 2
DNBHGA | 2 | B | M6 | 1.55 | -1.75 | 5 | -1.92 | 7.5 | -1.56 | 0.36 | -0.08 | 0.00056 | 0.00076 | 0.002 | 6 | 2
DNBHGA | 3 | B | M6 | 2.57 | -1.97 | 5.11 | -1.91 | 7.77 | -1.54 | 0.53 | -0.6 | 0.00056 | 0.00079 | 0.002 | 6 | 2
DNBHGA | 4 | B | M6 | 0.72 | -0.28 | 5.18 | -1.91 | 7.98 | -1.51 | 0.76 | -1.72 | 0.00056 | 0.00081 | 0.002 | 6 | 2
DNBHGA | 5 | B | M6 | 2.18 | -1.7 | 5.39 | -1.86 | 7.71 | -1.52 | 0.19 | 0.88 | 0.00057 | 0.00078 | 0.002 | 6 | 2
DNBHGA | 6 | B | M6 | 0.47 | -0.09 | 5.66 | -1.81 | 7.71 | -1.53 | 0.77 | -1.08 | 0.00059 | 0.00079 | 0.002 | 6 | 2
DNBHGA | 7 | B | M6 | 1.09 | -0.71 | 5.54 | -1.87 | 8.57 | -1.44 | 0.68 | -0.56 | 0.00058 | 0.00086 | 0.002 | 6 | 2
DNBHGA | 8 | B | M6 | 1.11 | -0.83 | 5.58 | -1.83 | 8 | -1.49 | 0.55 | -0.52 | 0.00058 | 0.00081 | 0.002 | 6 | 2
DNBHGA | 9 | B | M6 | 0.44 | -0.21 | 5.51 | -1.86 | 9.38 | -1.36 | 0.66 | -0.7 | 0.00057 | 0.00093 | 0.002 | 6 | 2
DNBHGA | 10 | B | M6 | 1.63 | -1.4 | 6.05 | -1.8 | 8.71 | -1.44 | 0.51 | -0.68 | 0.00061 | 0.00088 | 0.002 | 6 | 2
DNBHGA | 11 | B | M6 | 1.65 | -1.98 | 5.49 | -1.92 | 8.64 | -1.45 | 0.64 | -0.72 | 0.00056 | 0.00088 | 0.002 | 6 | 2
DNBHGA | 12 | B | M6 | 0.77 | -0.41 | 5.81 | -1.85 | 8.56 | -1.46 | 0.75 | -0.88 | 0.00058 | 0.00087 | 0.002 | 6 | 2
DNBHGA | 13 | B | M6 | 1.57 | -1 | 5.99 | -1.84 | 8.33 | -1.49 | 0.96 | -0.84 | 0.00059 | 0.00087 | 0.002 | 6 | 2
DNBHGA | 14 | B | M6 | 0.96 | -0.64 | 5.95 | -1.84 | 8.75 | -1.46 | 1.17 | -0.96 | 0.00059 | 0.00091 | 0.002 | 6 | 2
DNBHGA | 15 | B | M6 | 1.22 | -0.72 | 6.31 | -1.8 | 8.65 | -1.47 | 1 | -1.33 | 0.00061 | 0.0009 | 0.002 | 6 | 2
DNBHGA | 16 | B | M6 | 1.94 | -1.41 | 6.76 | -1.73 | 9.32 | -1.42 | 0.57 | -0.11 | 0.00064 | 0.00094 | 0.002 | 6 | 2
DNBHGB | 1 | A | M6 | 0.59 | -0.06 | 7.25 | -1.63 | 10.09 | -1.49 | 1.04 | -1.55 | 0.00072 | 0.00013 | 0.002 | 6 | 2
DNBHGB | 2 | A | M6 | 1.43 | -1.1 | 7.62 | -1.65 | 10.82 | -1.46 | 0.8 | -0.77 | 0.00075 | 9e-05 | 0.002 | 6 | 2
DNBHGB | 3 | A | M6 | 0.57 | 0.24 | 7.19 | -1.7 | 10.87 | -1.45 | 0.88 | -0.9 | 0.00071 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 4 | A | M6 | 0.72 | -0.27 | 7.55 | -1.69 | 10.41 | -1.48 | 0.79 | -1.43 | 0.00073 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 5 | A | M6 | 1.5 | -0.93 | 7.63 | -1.68 | 10.62 | -1.47 | 0.69 | -1.33 | 0.00072 | 8e-05 | 0.002 | 6 | 2
DNBHGB | 6 | A | M6 | 2.55 | -1.44 | 8.03 | -1.61 | 11.65 | -1.43 | 0.63 | -0.49 | 0.00074 | 7e-05 | 0.002 | 6 | 2
DNBHGB | 7 | A | M6 | 7.32 | -2.95 | 8.2 | -1.63 | 11.09 | -1.45 | 0.92 | -1.23 | 0.00074 | 0.00011 | 0.002 | 6 | 2
DNBHGB | 8 | A | M6 | 5.18 | -2.1 | 8.19 | -1.62 | 11.82 | -1.41 | 0.9 | -1.16 | 0.00074 | 0.00011 | 0.002 | 6 | 2
DNBHGB | 9 | A | M6 | 12.84 | -3.29 | 8.45 | -1.62 | 11.63 | -1.42 | 1.26 | -1.56 | 0.00075 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 10 | A | M6 | 2.77 | -1.53 | 8.24 | -1.68 | 12.79 | -1.38 | 0.62 | -0.64 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBHGB | 11 | A | M6 | 6.64 | -2.42 | 8.9 | -1.59 | 12.32 | -1.42 | 0.91 | -1.39 | 0.00076 | 0.00011 | 0.002 | 6 | 2
DNBHGB | 12 | A | M6 | 2.03 | -1.34 | 9.26 | -1.57 | 12.46 | -1.41 | 1.11 | -0.88 | 0.00078 | 0.00013 | 0.002 | 6 | 2
DNBHGB | 13 | A | M6 | 2.92 | -1.23 | 9.7 | -1.57 | 12.5 | -1.42 | 0.76 | -1.11 | 0.00081 | 9e-05 | 0.002 | 6 | 2
DNBHGB | 14 | A | M6 | 1.05 | -0.41 | 9.25 | -1.66 | 13.08 | -1.39 | 1.31 | -1.38 | 0.00078 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 15 | A | M6 | 2.47 | -1.49 | 9.07 | -1.63 | 13.8 | -1.36 | 1 | -1.12 | 0.00076 | 0.00012 | 0.002 | 6 | 2
DNBHGB | 16 | A | M6 | 3.25 | -1.22 | 10.04 | -1.54 | 13.51 | -1.38 | 1.42 | -1.13 | 0.00082 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 1 | B | M6 | 0.31 | 0.57 | 6.48 | -1.69 | 9.98 | -1.48 | 0.79 | -1.72 | 0.0007 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 2 | B | M6 | 0.91 | -0.32 | 6.82 | -1.67 | 9.88 | -1.49 | 0.67 | -0.79 | 0.00072 | 7e-05 | 0.002 | 6 | 2
DNBHGB | 3 | B | M6 | 4.13 | -2.49 | 6.7 | -1.69 | 9.94 | -1.49 | 0.51 | -1.1 | 0.0007 | 6e-05 | 0.002 | 6 | 2
DNBHGB | 4 | B | M6 | 6.32 | -2.82 | 6.99 | -1.67 | 9.76 | -1.49 | 0.82 | -1.14 | 0.00072 | 9e-05 | 0.002 | 6 | 2
DNBHGB | 5 | B | M6 | 4.88 | -2.64 | 7.05 | -1.66 | 9.57 | -1.5 | 0.72 | -1.32 | 0.00072 | 9e-05 | 0.002 | 6 | 2
DNBHGB | 6 | B | M6 | 2.37 | -1.6 | 7.19 | -1.66 | 10.33 | -1.47 | 0.53 | -1.29 | 0.00073 | 6e-05 | 0.002 | 6 | 2
DNBHGB | 7 | B | M6 | 2.41 | -1.79 | 7.27 | -1.65 | 10 | -1.47 | 0.74 | -0.62 | 0.00072 | 8e-05 | 0.002 | 6 | 2
DNBHGB | 8 | B | M6 | 2.57 | -1.67 | 7.46 | -1.63 | 11 | -1.41 | 0.88 | -1.71 | 0.00073 | 0.00011 | 0.002 | 6 | 2
DNBHGB | 9 | B | M6 | 1.26 | -0.85 | 7.52 | -1.63 | 10.3 | -1.46 | 0.82 | -1.25 | 0.00073 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 10 | B | M6 | 1.14 | -0.68 | 7.45 | -1.65 | 10.97 | -1.43 | 0.81 | -0.94 | 0.00072 | 9e-05 | 0.002 | 6 | 2
DNBHGB | 11 | B | M6 | 1.32 | -1.15 | 7.57 | -1.71 | 10.91 | -1.44 | 0.84 | -1.19 | 0.00073 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 12 | B | M6 | 0.26 | 0.94 | 8.03 | -1.58 | 10.8 | -1.45 | 1.07 | -1.84 | 0.00075 | 0.00014 | 0.002 | 6 | 2
DNBHGB | 13 | B | M6 | 1.88 | -0.75 | 8.25 | -1.58 | 11.1 | -1.44 | 1.05 | -1.64 | 0.00076 | 0.00013 | 0.002 | 6 | 2
DNBHGB | 14 | B | M6 | 0.76 | -0.23 | 8.45 | -1.57 | 10.95 | -1.45 | 0.9 | -0.98 | 0.00077 | 0.0001 | 0.002 | 6 | 2
DNBHGB | 15 | B | M6 | 0.92 | -0.45 | 8.36 | -1.59 | 11.34 | -1.43 | 1.18 | -0.66 | 0.00077 | 0.00013 | 0.002 | 6 | 2
DNBHGB | 16 | B | M6 | 1.36 | -0.75 | 8.2 | -1.62 | 11.62 | -1.42 | 0.98 | -1.04 | 0.00075 | 0.00011 | 0.002 | 6 | 2
DNBHGA | 1 | A | M5 | 1.17 | -2.1 | 7.25 | -1.73 | 10.26 | -1.33 | 2.05 | -3.49 | 0.00074 | 0.00103 | 0.002 | 6 | 2
DNBHGA | 2 | A | M5 | 1.03 | -1.62 | 7.56 | -1.71 | 10.66 | -1.31 | 1.87 | -2.8 | 0.00076 | 0.00107 | 0.002 | 6 | 2
DNBHGA | 3 | A | M5 | 0.93 | -1.77 | 7.68 | -1.7 | 10.83 | -1.31 | 1.72 | -2.62 | 0.00076 | 0.00108 | 0.002 | 6 | 2
DNBHGA | 4 | A | M5 | 1.28 | -2.18 | 7.74 | -1.72 | 10.87 | -1.3 | 1.76 | -2.92 | 0.00075 | 0.00109 | 0.002 | 6 | 2
DNBHGA | 5 | A | M5 | 0.88 | -1.74 | 7.75 | -1.73 | 10.85 | -1.3 | 1.65 | -2.71 | 0.00075 | 0.00108 | 0.002 | 6 | 2
DNBHGA | 6 | A | M5 | 1.27 | -2.16 | 7.8 | -1.74 | 11.02 | -1.28 | 1.24 | -2.36 | 0.00074 | 0.00109 | 0.002 | 6 | 2
DNBHGA | 7 | A | M5 | 0.52 | -0.94 | 8.14 | -1.7 | 11.16 | -1.28 | 1.62 | -2.53 | 0.00076 | 0.00111 | 0.002 | 6 | 2
DNBHGA | 8 | A | M5 | 0.72 | -1.08 | 8.33 | -1.68 | 11.32 | -1.27 | 1.53 | -2.85 | 0.00077 | 0.00113 | 0.002 | 6 | 2
DNBHGA | 9 | A | M5 | 0.77 | -1.54 | 8.48 | -1.67 | 11.33 | -1.28 | 1.25 | -2.49 | 0.00077 | 0.00112 | 0.002 | 6 | 2
DNBHGA | 10 | A | M5 | 1.42 | -2.25 | 8.4 | -1.71 | 11.54 | -1.27 | 1.12 | -2.25 | 0.00076 | 0.00114 | 0.002 | 6 | 2
DNBHGA | 11 | A | M5 | 1.17 | -1.69 | 8.56 | -1.72 | 11.48 | -1.28 | 1.33 | -2.31 | 0.00077 | 0.00114 | 0.002 | 6 | 2
DNBHGA | 12 | A | M5 | 1.47 | -2.35 | 8.83 | -1.68 | 11.59 | -1.28 | 1.1 | -2.17 | 0.00078 | 0.00115 | 0.002 | 6 | 2
DNBHGA | 13 | A | M5 | 0.81 | -1.55 | 8.87 | -1.68 | 11.68 | -1.28 | 1.73 | -2.99 | 0.00077 | 0.00117 | 0.002 | 6 | 2
DNBHGA | 14 | A | M5 | 1.1 | -1.62 | 8.93 | -1.68 | 12.02 | -1.27 | 1.11 | -1.97 | 0.00077 | 0.00119 | 0.002 | 6 | 2
DNBHGA | 15 | A | M5 | 1.33 | -2.1 | 8.88 | -1.69 | 12.01 | -1.26 | 1.28 | -1.8 | 0.00076 | 0.00119 | 0.002 | 6 | 2
DNBHGA | 16 | A | M5 | 0.99 | -1.75 | 9.19 | -1.65 | 12.07 | -1.26 | 1.29 | -2.07 | 0.00078 | 0.0012 | 0.002 | 6 | 2
DNBHGA | 1 | B | M5 | 0.89 | -1.5 | 6.98 | -1.72 | 9.9 | -1.34 | 1.57 | -2.21 | 0.00075 | 0.00097 | 0.002 | 6 | 2
DNBHGA | 2 | B | M5 | 0.83 | -1.73 | 6.99 | -1.72 | 9.99 | -1.34 | 1.35 | -2.07 | 0.00075 | 0.00098 | 0.002 | 6 | 2
DNBHGA | 3 | B | M5 | 0.95 | -1.8 | 7.1 | -1.72 | 10.28 | -1.32 | 2.19 | -2.99 | 0.00075 | 0.00102 | 0.002 | 6 | 2
DNBHGA | 4 | B | M5 | 0.7 | -1.27 | 7.23 | -1.72 | 10.14 | -1.33 | 1.81 | -2.79 | 0.00075 | 0.001 | 0.002 | 6 | 2
DNBHGA | 5 | B | M5 | 1.17 | -1.85 | 7.22 | -1.73 | 10.09 | -1.32 | 1.39 | -2.35 | 0.00075 | 0.00099 | 0.002 | 6 | 2
DNBHGA | 6 | B | M5 | 0.87 | -1.48 | 7.2 | -1.75 | 10.15 | -1.31 | 1.41 | -2.39 | 0.00074 | 0.001 | 0.002 | 6 | 2
DNBHGA | 7 | B | M5 | 1.23 | -2.05 | 7.37 | -1.72 | 10.28 | -1.3 | 1.3 | -2.07 | 0.00075 | 0.00101 | 0.002 | 6 | 2
DNBHGA | 8 | B | M5 | 2.25 | -2.94 | 7.38 | -1.71 | 10.13 | -1.31 | 2.12 | -2.96 | 0.00074 | 0.00101 | 0.002 | 6 | 2
DNBHGA | 9 | B | M5 | 0.99 | -1.7 | 7.47 | -1.71 | 10.34 | -1.29 | 2.37 | -3.13 | 0.00074 | 0.00104 | 0.002 | 6 | 2
DNBHGA | 10 | B | M5 | 1.41 | -2.09 | 7.54 | -1.71 | 10.22 | -1.31 | 1.14 | -1.8 | 0.00074 | 0.00101 | 0.002 | 6 | 2
DNBHGA | 11 | B | M5 | 0.79 | -1.24 | 7.64 | -1.7 | 10.22 | -1.31 | 1.4 | -2.56 | 0.00075 | 0.00102 | 0.002 | 6 | 2
DNBHGA | 12 | B | M5 | 1.37 | -2.01 | 7.78 | -1.68 | 10.3 | -1.3 | 1.28 | -2.12 | 0.00075 | 0.00103 | 0.002 | 6 | 2
DNBHGA | 13 | B | M5 | 1.21 | -1.74 | 7.75 | -1.7 | 10.49 | -1.29 | 1.68 | -2.6 | 0.00075 | 0.00105 | 0.002 | 6 | 2
DNBHGA | 14 | B | M5 | 2.06 | -2.63 | 7.85 | -1.68 | 10.56 | -1.29 | 1.49 | -2.37 | 0.00075 | 0.00106 | 0.002 | 6 | 2
DNBHGA | 15 | B | M5 | 0.91 | -1.33 | 7.99 | -1.65 | 10.44 | -1.3 | 1.4 | -2.26 | 0.00075 | 0.00104 | 0.002 | 6 | 2
DNBHGA | 16 | B | M5 | 2.53 | -3.02 | 8.08 | -1.65 | 10.55 | -1.29 | 1.38 | -1.93 | 0.00076 | 0.00106 | 0.002 | 6 | 2
DNBHGB | 1 | A | M5 | 2.1 | -2.42 | 8.44 | -1.49 | 11.49 | -1.36 | 1.12 | -2.67 | 0.00082 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 2 | A | M5 | 4.03 | -3.18 | 8.61 | -1.47 | 11.58 | -1.36 | 1.35 | -2.32 | 0.00082 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 3 | A | M5 | 2.29 | -2.47 | 8.85 | -1.46 | 11.75 | -1.35 | 1.21 | -2.27 | 0.00083 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 4 | A | M5 | 2.35 | -2.49 | 8.95 | -1.46 | 11.91 | -1.34 | 1.17 | -2.65 | 0.00083 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 5 | A | M5 | 4.98 | -3.35 | 8.99 | -1.46 | 12.21 | -1.33 | 1.18 | -3 | 0.00082 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 6 | A | M5 | 4.62 | -3.27 | 9.16 | -1.45 | 12.34 | -1.33 | 1.08 | -2.59 | 0.00083 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 7 | A | M5 | 2.28 | -2.26 | 9.52 | -1.44 | 12.23 | -1.33 | 1.2 | -2.98 | 0.00084 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 8 | A | M5 | 1.77 | -2.22 | 9.6 | -1.43 | 12.84 | -1.3 | 1.14 | -2.43 | 0.00084 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 9 | A | M5 | 6.37 | -3.43 | 9.69 | -1.44 | 12.63 | -1.32 | 1.22 | -2.25 | 0.00083 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 10 | A | M5 | 16.31 | -4.61 | 9.95 | -1.44 | 13.18 | -1.31 | 1.2 | -2.53 | 0.00084 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 11 | A | M5 | 2.03 | -1.95 | 10.23 | -1.42 | 13.53 | -1.3 | 1.28 | -2.52 | 0.00085 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 12 | A | M5 | 5.58 | -3.21 | 10.29 | -1.43 | 13.7 | -1.3 | 1.31 | -2.77 | 0.00085 | 0.0002 | 0.002 | 6 | 2
DNBHGB | 13 | A | M5 | 1.1 | -1.72 | 10.54 | -1.42 | 13.91 | -1.3 | 1.59 | -2.28 | 0.00086 | 0.00022 | 0.002 | 6 | 2
DNBHGB | 14 | A | M5 | 1.27 | -1.82 | 10.51 | -1.41 | 13.67 | -1.3 | 1.29 | -2.38 | 0.00085 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 15 | A | M5 | 2.65 | -2.58 | 10.77 | -1.4 | 13.65 | -1.31 | 1.4 | -2.82 | 0.00086 | 0.00022 | 0.002 | 6 | 2
DNBHGB | 16 | A | M5 | 6.21 | -3.41 | 10.91 | -1.39 | 13.78 | -1.3 | 1.4 | -2.3 | 0.00086 | 0.0002 | 0.002 | 6 | 2
DNBHGB | 1 | B | M5 | 3.5 | -3.09 | 8.19 | -1.44 | 10.91 | -1.38 | 1.33 | -2.66 | 0.00084 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 2 | B | M5 | 1.62 | -2.27 | 8.12 | -1.49 | 10.98 | -1.38 | 1.33 | -2.26 | 0.00083 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 3 | B | M5 | 2.24 | -2.39 | 8.3 | -1.45 | 11.02 | -1.38 | 1.19 | -2.25 | 0.00083 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 4 | B | M5 | 1.81 | -2.2 | 8.36 | -1.45 | 11.25 | -1.37 | 1.1 | -2.47 | 0.00083 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 5 | B | M5 | 1.56 | -1.91 | 8.34 | -1.49 | 11.14 | -1.37 | 1.2 | -2.39 | 0.00082 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 6 | B | M5 | 4.45 | -3.37 | 8.51 | -1.44 | 11.38 | -1.36 | 1.2 | -2.66 | 0.00082 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 7 | B | M5 | 13.97 | -4.48 | 8.55 | -1.45 | 11.38 | -1.35 | 1.2 | -2.17 | 0.00082 | 0.00016 | 0.002 | 6 | 2
DNBHGB | 8 | B | M5 | 1.58 | -1.98 | 8.71 | -1.43 | 11.48 | -1.34 | 1.12 | -2.88 | 0.00082 | 0.00017 | 0.002 | 6 | 2
DNBHGB | 9 | B | M5 | 5.8 | -3.56 | 8.81 | -1.44 | 11.6 | -1.34 | 1.23 | -2.49 | 0.00082 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 10 | B | M5 | 2 | -2.21 | 9 | -1.43 | 11.67 | -1.34 | 1.3 | -2.69 | 0.00083 | 0.00019 | 0.002 | 6 | 2
DNBHGB | 11 | B | M5 | 0.93 | -1.64 | 8.96 | -1.44 | 11.86 | -1.33 | 1.29 | -2.25 | 0.00082 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 12 | B | M5 | 1.31 | -1.65 | 9.15 | -1.42 | 12.02 | -1.33 | 1.48 | -2.78 | 0.00083 | 0.00023 | 0.002 | 6 | 2
DNBHGB | 13 | B | M5 | 3.19 | -2.79 | 9.37 | -1.4 | 12.12 | -1.33 | 1.65 | -2.92 | 0.00084 | 0.00026 | 0.002 | 6 | 2
DNBHGB | 14 | B | M5 | 5.1 | -3.41 | 9.35 | -1.4 | 12.15 | -1.33 | 1.21 | -2.67 | 0.00083 | 0.00018 | 0.002 | 6 | 2
DNBHGB | 15 | B | M5 | 2.69 | -2.18 | 9.42 | -1.39 | 12.01 | -1.33 | 1.38 | -2.46 | 0.00083 | 0.0002 | 0.002 | 6 | 2
DNBHGB | 16 | B | M5 | 2.33 | -2.31 | 9.69 | -1.39 | 12.02 | -1.33 | 1.26 | -2.51 | 0.00085 | 0.00018 | 0.002 | 6 | 2

## Polarization Factor
Band | HAM | Band Median | [10 empty] | Band Average
|  | Polarization Factor | [10 empty] | Polarization Factor
|  | Scan Angle | [10 empty] | Scan Angle
|  | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55 | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55
I1 | A | 0.77383 | 0.66656 | 0.6594 | 0.62853 | 0.63077 | 0.68407 | 0.69988 | 0.73676 | 0.75247 | 0.76062 | 0.76111 | 0.71219 | 0.61016 | 0.60647 | 0.57448 | 0.58714 | 0.64031 | 0.65008 | 0.69248 | 0.72969 | 0.7459 | 0.74699
I1 | B | 0.40078 | 0.5128 | 0.64277 | 0.69125 | 0.7524 | 0.81391 | 0.85221 | 0.92197 | 0.93656 | 0.94226 | 0.94127 | 0.44733 | 0.55035 | 0.66024 | 0.695 | 0.74199 | 0.80675 | 0.82362 | 0.8823 | 0.91281 | 0.92625 | 0.92501
I2 | A | 0.67681 | 0.56403 | 0.49567 | 0.40926 | 0.3251 | 0.34 | 0.2813 | 0.2403 | 0.26192 | 0.3618 | 0.41006 | 0.65725 | 0.54492 | 0.47638 | 0.38564 | 0.28733 | 0.28381 | 0.23626 | 0.23559 | 0.27893 | 0.38372 | 0.43395
I2 | B | 1.1381 | 0.8646 | 0.70663 | 0.56672 | 0.41026 | 0.38571 | 0.31718 | 0.2816 | 0.30159 | 0.3706 | 0.40095 | 1.104 | 0.83098 | 0.66553 | 0.51654 | 0.3704 | 0.34486 | 0.28654 | 0.28208 | 0.31739 | 0.3887 | 0.42431
M1 | A | 3.165 | 3.5271 | 3.7437 | 4.0108 | 4.2405 | 4.3624 | 4.4783 | 4.6809 | 4.9171 | 5.0316 | 5.0108 | 3.2575 | 3.6209 | 3.8274 | 4.079 | 4.3154 | 4.4187 | 4.5425 | 4.731 | 4.9593 | 5.0683 | 5.0466
M1 | B | 3.4346 | 3.8171 | 4.0916 | 4.4007 | 4.7284 | 4.8716 | 5.0769 | 5.3839 | 5.7538 | 5.9562 | 5.9345 | 3.4872 | 3.8734 | 4.1388 | 4.4571 | 4.7727 | 4.9133 | 5.1109 | 5.4101 | 5.7594 | 5.9429 | 5.9146
M2 | A | 2.0204 | 2.3051 | 2.4906 | 2.6859 | 2.9003 | 3.0152 | 3.1009 | 3.2549 | 3.4912 | 3.6219 | 3.6062 | 2.2952 | 2.5537 | 2.7158 | 2.8795 | 3.0596 | 3.1465 | 3.2223 | 3.3554 | 3.5389 | 3.6498 | 3.652
M2 | B | 2.0914 | 2.4299 | 2.6504 | 2.8727 | 3.11 | 3.2325 | 3.3466 | 3.5532 | 3.85 | 4.0157 | 3.9998 | 2.335 | 2.634 | 2.8312 | 3.0317 | 3.2522 | 3.3562 | 3.466 | 3.6486 | 3.8894 | 4.0421 | 4.0501
M3 | A | 1.5242 | 1.7112 | 1.8679 | 2.0174 | 2.1887 | 2.2765 | 2.3378 | 2.4615 | 2.5438 | 2.7065 | 2.7152 | 1.7075 | 1.8569 | 1.9598 | 2.0635 | 2.1943 | 2.2595 | 2.3084 | 2.4102 | 2.544 | 2.6285 | 2.6152
M3 | B | 1.5736 | 1.837 | 2.0117 | 2.1734 | 2.3535 | 2.4461 | 2.5199 | 2.6229 | 2.7749 | 2.9496 | 2.9603 | 1.6792 | 1.8895 | 2.0321 | 2.1668 | 2.3236 | 2.4013 | 2.4695 | 2.5984 | 2.7601 | 2.869 | 2.8624
M4 | A | 1.5688 | 1.8008 | 2.0229 | 2.1246 | 2.2862 | 2.4177 | 2.4689 | 2.5895 | 2.6802 | 2.7275 | 2.7228 | 1.7902 | 1.9764 | 2.1533 | 2.2302 | 2.3272 | 2.433 | 2.4698 | 2.5582 | 2.6223 | 2.6585 | 2.6555
M4 | B | 1.7184 | 2.0083 | 2.2407 | 2.3391 | 2.4882 | 2.6097 | 2.6548 | 2.7674 | 2.8448 | 2.88 | 2.8698 | 1.8742 | 2.0908 | 2.2777 | 2.359 | 2.4642 | 2.5695 | 2.6064 | 2.6989 | 2.7652 | 2.8037 | 2.7993
M5 | A | 1.7031 | 1.5505 | 1.5247 | 1.481 | 1.4654 | 1.5387 | 1.5448 | 1.6113 | 1.6354 | 1.5774 | 1.5646 | 1.4777 | 1.3231 | 1.2786 | 1.2374 | 1.246 | 1.3017 | 1.3145 | 1.3728 | 1.4102 | 1.4312 | 1.4397
M5 | B | 1.0324 | 1.1496 | 1.3147 | 1.3874 | 1.4968 | 1.6106 | 1.6608 | 1.7737 | 1.8277 | 1.7407 | 1.7055 | 1.0791 | 1.1244 | 1.2386 | 1.3002 | 1.3777 | 1.4566 | 1.4883 | 1.5672 | 1.6016 | 1.6041 | 1.5972
M6 | A | 1.4522 | 1.1496 | 0.97619 | 0.8235 | 0.68736 | 0.67245 | 0.61781 | 0.58907 | 0.57672 | 0.59655 | 0.61347 | 1.4117 | 1.1173 | 0.94196 | 0.79168 | 0.65865 | 0.64101 | 0.58339 | 0.53507 | 0.5219 | 0.54578 | 0.5652
M6 | B | 0.87602 | 0.76536 | 0.75839 | 0.7335 | 0.73402 | 0.77018 | 0.76462 | 0.79588 | 0.7995 | 0.77981 | 0.77271 | 0.83615 | 0.71183 | 0.68233 | 0.63879 | 0.63926 | 0.68587 | 0.68886 | 0.72867 | 0.73944 | 0.73348 | 0.72768
M7 | A | 0.67453 | 0.56689 | 0.49261 | 0.40633 | 0.31963 | 0.3269 | 0.28302 | 0.24533 | 0.25842 | 0.35 | 0.3988 | 0.65713 | 0.54672 | 0.47407 | 0.38553 | 0.28748 | 0.28319 | 0.23225 | 0.22902 | 0.26601 | 0.36894 | 0.41822
M7 | B | 1.1233 | 0.85966 | 0.6972 | 0.55924 | 0.41523 | 0.39146 | 0.32044 | 0.28748 | 0.29006 | 0.35771 | 0.39273 | 1.0972 | 0.8286 | 0.66092 | 0.51452 | 0.36815 | 0.34207 | 0.28058 | 0.27415 | 0.30582 | 0.37626 | 0.41144
Band | HAM | Detector | Detector Dependent | [11 empty] | Notes:
|  |  | Polarization Factor | [11 empty] | 1. This table provides VIIRS F2 Polarization factors [in %].
|  |  | Scan Angle | [11 empty] | 2. The results were derived from FP-11 and FP-11' using the SIS-100-2 source.
|  |  | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55 |  | 3. The bvonir polarization sheet was used with (M1-M3) and without (M4-M7 and I1-I2) blocking filter.
I1 | A | 1 | 0.793 | 0.73928 | 0.76512 | 0.72933 | 0.73104 | 0.7866 | 0.75823 | 0.77901 | 0.74584 | 0.69397 | 0.66295 |  | 4. Fourier analysis was performed to determine the two-cycle polarization factor.
I1 | A | 2 | 0.7822 | 0.73262 | 0.7514 | 0.73153 | 0.7219 | 0.76829 | 0.74853 | 0.76343 | 0.74432 | 0.69951 | 0.67618 |  | 5. Specification states that polarization sensitivity should be less then 3% for M1, M7, and I2 and 2.5% for M2-M6 and I1.
I1 | A | 3 | 0.77826 | 0.7214 | 0.73979 | 0.70735 | 0.71162 | 0.74915 | 0.74401 | 0.75607 | 0.74865 | 0.71164 | 0.69465 |  | 6. The testing was conducted at Raytheon El Segundo, December 2013 / January 2014 (FP-11) and December 2014 (FP-11') .
I1 | A | 4 | 0.77656 | 0.71727 | 0.74979 | 0.71627 | 0.70396 | 0.75388 | 0.74977 | 0.78335 | 0.7563 | 0.72565 | 0.71955 |  | 7. The test was set to diagnostic mode.
I1 | A | 5 | 0.77875 | 0.71222 | 0.74822 | 0.71478 | 0.69978 | 0.74756 | 0.74928 | 0.78535 | 0.76498 | 0.74022 | 0.73764
I1 | A | 6 | 0.77816 | 0.70531 | 0.72302 | 0.69356 | 0.6912 | 0.74018 | 0.74278 | 0.76979 | 0.77174 | 0.75264 | 0.73631
I1 | A | 7 | 0.79613 | 0.71322 | 0.7438 | 0.7131 | 0.69713 | 0.74578 | 0.74829 | 0.77481 | 0.78259 | 0.76846 | 0.76323
I1 | A | 8 | 0.7838 | 0.712 | 0.72406 | 0.68925 | 0.69315 | 0.7392 | 0.74347 | 0.76751 | 0.78634 | 0.77909 | 0.77563
I1 | A | 9 | 0.79076 | 0.70995 | 0.71992 | 0.68536 | 0.68951 | 0.75446 | 0.74035 | 0.7712 | 0.79625 | 0.7933 | 0.79018
I1 | A | 10 | 0.79587 | 0.70964 | 0.73093 | 0.68838 | 0.69063 | 0.7504 | 0.75358 | 0.80122 | 0.80875 | 0.81197 | 0.80895
I1 | A | 11 | 0.78974 | 0.70636 | 0.71091 | 0.67337 | 0.67904 | 0.73018 | 0.74105 | 0.77587 | 0.80094 | 0.81184 | 0.81966
I1 | A | 12 | 0.80815 | 0.70938 | 0.70961 | 0.68186 | 0.68811 | 0.72519 | 0.74327 | 0.76945 | 0.82401 | 0.83459 | 0.83397
I1 | A | 13 | 0.78889 | 0.68212 | 0.69908 | 0.64834 | 0.65379 | 0.74251 | 0.72757 | 0.78204 | 0.7997 | 0.81883 | 0.82307
I1 | A | 14 | 0.7889 | 0.69096 | 0.68461 | 0.66468 | 0.66327 | 0.71494 | 0.72413 | 0.76234 | 0.80516 | 0.82616 | 0.82752
I1 | A | 15 | 0.79664 | 0.67771 | 0.67274 | 0.63286 | 0.64844 | 0.70721 | 0.71565 | 0.75223 | 0.81004 | 0.83712 | 0.84113
I1 | A | 16 | 0.79191 | 0.66912 | 0.66278 | 0.63028 | 0.63455 | 0.69916 | 0.70952 | 0.74856 | 0.80623 | 0.84338 | 0.85206
I1 | A | 17 | 0.7711 | 0.664 | 0.65602 | 0.62679 | 0.62699 | 0.66898 | 0.69025 | 0.72496 | 0.79211 | 0.82381 | 0.82317
I1 | A | 18 | 0.75685 | 0.64253 | 0.62471 | 0.59779 | 0.60597 | 0.6603 | 0.67381 | 0.71608 | 0.7759 | 0.80923 | 0.81543
I1 | A | 19 | 0.74588 | 0.63456 | 0.62944 | 0.58743 | 0.58304 | 0.63971 | 0.65574 | 0.70598 | 0.7568 | 0.79734 | 0.80627
I1 | A | 20 | 0.72159 | 0.61435 | 0.59496 | 0.553 | 0.56796 | 0.64261 | 0.64362 | 0.70702 | 0.74659 | 0.78938 | 0.80485
I1 | A | 21 | 0.71134 | 0.59645 | 0.57273 | 0.55527 | 0.55877 | 0.61013 | 0.62987 | 0.6747 | 0.73908 | 0.78096 | 0.79229
I1 | A | 22 | 0.70948 | 0.5701 | 0.57104 | 0.5267 | 0.53537 | 0.58973 | 0.61216 | 0.66634 | 0.72149 | 0.77913 | 0.77983
I1 | A | 23 | 0.6765 | 0.55853 | 0.55083 | 0.49261 | 0.52073 | 0.56564 | 0.59514 | 0.65778 | 0.70962 | 0.75278 | 0.75898
I1 | A | 24 | 0.6606 | 0.53063 | 0.50682 | 0.47568 | 0.49715 | 0.54941 | 0.57082 | 0.62186 | 0.69355 | 0.73867 | 0.74422
I1 | A | 25 | 0.63338 | 0.51152 | 0.47788 | 0.4483 | 0.47967 | 0.53615 | 0.55106 | 0.60326 | 0.67217 | 0.71359 | 0.72698
I1 | A | 26 | 0.6079 | 0.48164 | 0.45745 | 0.42774 | 0.45418 | 0.50728 | 0.53159 | 0.58625 | 0.65018 | 0.68978 | 0.7042
I1 | A | 27 | 0.61005 | 0.47987 | 0.43369 | 0.40806 | 0.45631 | 0.51445 | 0.52403 | 0.57813 | 0.64975 | 0.69053 | 0.70259
I1 | A | 28 | 0.55787 | 0.43131 | 0.40758 | 0.37246 | 0.40947 | 0.46987 | 0.49203 | 0.56343 | 0.61233 | 0.65262 | 0.65886
I1 | A | 29 | 0.5495 | 0.42334 | 0.38951 | 0.38122 | 0.41309 | 0.44776 | 0.48155 | 0.52819 | 0.60445 | 0.64041 | 0.64555
I1 | A | 30 | 0.5082 | 0.37527 | 0.35723 | 0.32845 | 0.36449 | 0.43071 | 0.4518 | 0.51523 | 0.57144 | 0.6043 | 0.6135
I1 | A | 31 | 0.48775 | 0.35441 | 0.32703 | 0.31094 | 0.3579 | 0.41623 | 0.43304 | 0.49015 | 0.55456 | 0.5875 | 0.59539
I1 | A | 32 | 0.46427 | 0.3479 | 0.31442 | 0.29072 | 0.36022 | 0.38619 | 0.42669 | 0.4777 | 0.54825 | 0.57049 | 0.56879
I1 | B | 1 | 0.85477 | 0.89466 | 0.94954 | 0.94688 | 0.9355 | 0.9801 | 0.94989 | 0.96418 | 0.91657 | 0.85175 | 0.81604
I1 | B | 2 | 0.80632 | 0.85453 | 0.92474 | 0.92292 | 0.91388 | 0.95637 | 0.93755 | 0.94518 | 0.91525 | 0.86121 | 0.8298
I1 | B | 3 | 0.7786 | 0.82452 | 0.89432 | 0.89463 | 0.9001 | 0.93209 | 0.93398 | 0.94402 | 0.92898 | 0.88148 | 0.85704
I1 | B | 4 | 0.7456 | 0.79376 | 0.88352 | 0.88045 | 0.89136 | 0.94599 | 0.93275 | 0.96432 | 0.93665 | 0.90224 | 0.87861
I1 | B | 5 | 0.71344 | 0.77586 | 0.86469 | 0.85951 | 0.88003 | 0.94064 | 0.92738 | 0.96623 | 0.94563 | 0.91692 | 0.89373
I1 | B | 6 | 0.6645 | 0.73783 | 0.83021 | 0.84325 | 0.85947 | 0.91453 | 0.91564 | 0.94677 | 0.94836 | 0.92701 | 0.90926
I1 | B | 7 | 0.64085 | 0.72 | 0.82755 | 0.84393 | 0.85437 | 0.91473 | 0.91973 | 0.96951 | 0.96112 | 0.94544 | 0.93529
I1 | B | 8 | 0.60799 | 0.68966 | 0.79667 | 0.81487 | 0.84025 | 0.89602 | 0.91144 | 0.95265 | 0.96105 | 0.95468 | 0.94724
I1 | B | 9 | 0.57725 | 0.66802 | 0.77267 | 0.79516 | 0.82966 | 0.89312 | 0.90468 | 0.95015 | 0.97216 | 0.96954 | 0.964
I1 | B | 10 | 0.55923 | 0.64968 | 0.76696 | 0.80187 | 0.82753 | 0.9099 | 0.91241 | 0.97469 | 0.99313 | 0.99389 | 0.987
I1 | B | 11 | 0.53207 | 0.62433 | 0.7385 | 0.76708 | 0.81122 | 0.8812 | 0.89968 | 0.95316 | 0.98402 | 0.99334 | 0.99325
I1 | B | 12 | 0.50295 | 0.62111 | 0.7303 | 0.76027 | 0.81333 | 0.87077 | 0.89779 | 0.95787 | 0.99801 | 1.015 | 1.0057
I1 | B | 13 | 0.47802 | 0.58691 | 0.69321 | 0.74056 | 0.79078 | 0.87339 | 0.87838 | 0.95302 | 0.98144 | 1.0003 | 0.99823
I1 | B | 14 | 0.44896 | 0.55829 | 0.69887 | 0.7333 | 0.76921 | 0.84807 | 0.87533 | 0.94137 | 0.98094 | 1.0056 | 1.0089
I1 | B | 15 | 0.43607 | 0.55418 | 0.66367 | 0.72098 | 0.77439 | 0.83472 | 0.86978 | 0.92441 | 0.98842 | 1.0153 | 1.0222
I1 | B | 16 | 0.42573 | 0.5341 | 0.64835 | 0.70959 | 0.77325 | 0.82282 | 0.86461 | 0.92321 | 0.99203 | 1.0262 | 1.0343
I1 | B | 17 | 0.37582 | 0.49149 | 0.6372 | 0.67292 | 0.73559 | 0.80499 | 0.83981 | 0.92073 | 0.96653 | 1.003 | 1.0045
I1 | B | 18 | 0.34756 | 0.47537 | 0.615 | 0.66707 | 0.71467 | 0.78865 | 0.82234 | 0.89419 | 0.95502 | 0.9931 | 0.99946
I1 | B | 19 | 0.33324 | 0.45609 | 0.59137 | 0.63143 | 0.69728 | 0.78332 | 0.81069 | 0.90061 | 0.94359 | 0.98504 | 0.99098
I1 | B | 20 | 0.30868 | 0.43878 | 0.56612 | 0.61981 | 0.69627 | 0.77963 | 0.80062 | 0.88315 | 0.93647 | 0.97508 | 0.98732
I1 | B | 21 | 0.28744 | 0.4211 | 0.57273 | 0.61715 | 0.67842 | 0.74757 | 0.78531 | 0.86324 | 0.92386 | 0.96878 | 0.98353
I1 | B | 22 | 0.27101 | 0.41064 | 0.53945 | 0.58329 | 0.66315 | 0.74474 | 0.77374 | 0.8608 | 0.90896 | 0.95914 | 0.96332
I1 | B | 23 | 0.26699 | 0.38706 | 0.52375 | 0.56683 | 0.65093 | 0.74149 | 0.75828 | 0.84921 | 0.89558 | 0.93909 | 0.94814
I1 | B | 24 | 0.25511 | 0.39019 | 0.51351 | 0.56275 | 0.64112 | 0.70791 | 0.74462 | 0.81554 | 0.87875 | 0.91953 | 0.93107
I1 | B | 25 | 0.24193 | 0.37876 | 0.49906 | 0.55032 | 0.62733 | 0.69443 | 0.7263 | 0.79517 | 0.85976 | 0.90042 | 0.91885
I1 | B | 26 | 0.2436 | 0.3679 | 0.4915 | 0.54267 | 0.61319 | 0.68223 | 0.7174 | 0.78718 | 0.84377 | 0.88512 | 0.91113
I1 | B | 27 | 0.2374 | 0.37463 | 0.48561 | 0.54156 | 0.6189 | 0.68184 | 0.70967 | 0.77314 | 0.84378 | 0.87951 | 0.88292
I1 | B | 28 | 0.25264 | 0.3713 | 0.47553 | 0.53295 | 0.59171 | 0.68067 | 0.68562 | 0.76116 | 0.79979 | 0.83481 | 0.85039
I1 | B | 29 | 0.25434 | 0.37934 | 0.49505 | 0.53322 | 0.6011 | 0.64948 | 0.68315 | 0.7554 | 0.7989 | 0.82641 | 0.82199
I1 | B | 30 | 0.2714 | 0.38317 | 0.47255 | 0.52266 | 0.57769 | 0.64592 | 0.66303 | 0.72518 | 0.76274 | 0.78882 | 0.80398
I1 | B | 31 | 0.29136 | 0.39914 | 0.47952 | 0.53011 | 0.58832 | 0.63495 | 0.65385 | 0.70475 | 0.75056 | 0.77099 | 0.77499
I1 | B | 32 | 0.30353 | 0.39883 | 0.48588 | 0.53008 | 0.58367 | 0.63363 | 0.65024 | 0.71338 | 0.73796 | 0.75138 | 0.7472
I2 | A | 1 | 0.5109 | 0.42248 | 0.38204 | 0.33941 | 0.32472 | 0.3562 | 0.3651 | 0.43155 | 0.50344 | 0.61172 | 0.65807
I2 | A | 2 | 0.54597 | 0.44857 | 0.40409 | 0.344 | 0.32077 | 0.35048 | 0.34884 | 0.41081 | 0.46963 | 0.57603 | 0.62448
I2 | A | 3 | 0.57708 | 0.47186 | 0.42486 | 0.36921 | 0.32549 | 0.35098 | 0.34452 | 0.39483 | 0.44867 | 0.55535 | 0.60366
I2 | A | 4 | 0.6033 | 0.50066 | 0.44752 | 0.38519 | 0.33181 | 0.35149 | 0.33914 | 0.37924 | 0.42713 | 0.53047 | 0.57813
I2 | A | 5 | 0.62058 | 0.52341 | 0.46605 | 0.39444 | 0.3378 | 0.35718 | 0.33283 | 0.36446 | 0.41301 | 0.51362 | 0.56044
I2 | A | 6 | 0.64497 | 0.54242 | 0.48297 | 0.41421 | 0.3442 | 0.36266 | 0.33287 | 0.35466 | 0.3947 | 0.49221 | 0.54208
I2 | A | 7 | 0.65767 | 0.5584 | 0.49871 | 0.41758 | 0.35011 | 0.36258 | 0.32938 | 0.34536 | 0.38089 | 0.47679 | 0.52234
I2 | A | 8 | 0.67361 | 0.56966 | 0.50768 | 0.42266 | 0.34825 | 0.36009 | 0.32334 | 0.33584 | 0.35959 | 0.44988 | 0.50551
I2 | A | 9 | 0.69709 | 0.58734 | 0.53192 | 0.44312 | 0.35985 | 0.37135 | 0.33074 | 0.33418 | 0.34576 | 0.4366 | 0.48529
I2 | A | 10 | 0.70306 | 0.59558 | 0.53378 | 0.44325 | 0.36374 | 0.36277 | 0.32437 | 0.31683 | 0.33529 | 0.42551 | 0.47142
I2 | A | 11 | 0.71315 | 0.60104 | 0.53598 | 0.4547 | 0.35902 | 0.3631 | 0.31781 | 0.30341 | 0.31373 | 0.40047 | 0.45122
I2 | A | 12 | 0.71829 | 0.6135 | 0.54814 | 0.45475 | 0.36407 | 0.36961 | 0.31453 | 0.29546 | 0.30667 | 0.39351 | 0.44451
I2 | A | 13 | 0.72705 | 0.61913 | 0.5525 | 0.45635 | 0.364 | 0.36124 | 0.31109 | 0.28996 | 0.28876 | 0.38 | 0.43484
I2 | A | 14 | 0.72898 | 0.6215 | 0.55759 | 0.46731 | 0.36286 | 0.36295 | 0.30647 | 0.27672 | 0.28063 | 0.36577 | 0.41931
I2 | A | 15 | 0.72916 | 0.62191 | 0.55632 | 0.46114 | 0.35644 | 0.35411 | 0.29466 | 0.26163 | 0.26074 | 0.34817 | 0.39825
I2 | A | 16 | 0.73257 | 0.62264 | 0.55472 | 0.45753 | 0.35273 | 0.34505 | 0.28838 | 0.2516 | 0.2474 | 0.33511 | 0.38714
I2 | A | 17 | 0.72955 | 0.62103 | 0.55114 | 0.45193 | 0.34652 | 0.33495 | 0.27423 | 0.229 | 0.23432 | 0.32624 | 0.37611
I2 | A | 18 | 0.73221 | 0.61659 | 0.54654 | 0.4543 | 0.33761 | 0.3269 | 0.26442 | 0.219 | 0.21671 | 0.30818 | 0.36418
I2 | A | 19 | 0.72086 | 0.60741 | 0.53719 | 0.43477 | 0.31933 | 0.31407 | 0.24565 | 0.19925 | 0.19686 | 0.29447 | 0.34891
I2 | A | 20 | 0.72298 | 0.60672 | 0.53608 | 0.4405 | 0.31833 | 0.30582 | 0.24104 | 0.19263 | 0.19078 | 0.29108 | 0.34404
I2 | A | 21 | 0.71981 | 0.60348 | 0.52929 | 0.42666 | 0.3034 | 0.29062 | 0.22384 | 0.17097 | 0.17531 | 0.28159 | 0.34009
I2 | A | 22 | 0.69471 | 0.57774 | 0.50041 | 0.4043 | 0.28009 | 0.26582 | 0.1948 | 0.13737 | 0.1609 | 0.27463 | 0.32991
I2 | A | 23 | 0.68866 | 0.57247 | 0.49263 | 0.39047 | 0.27275 | 0.25014 | 0.18304 | 0.1253 | 0.15738 | 0.27583 | 0.32888
I2 | A | 24 | 0.68001 | 0.55674 | 0.48014 | 0.38126 | 0.25028 | 0.23344 | 0.15946 | 0.10471 | 0.14876 | 0.2715 | 0.32683
I2 | A | 25 | 0.66145 | 0.54535 | 0.46901 | 0.35877 | 0.23242 | 0.21404 | 0.13831 | 0.085657 | 0.14636 | 0.2771 | 0.32698
I2 | A | 26 | 0.6476 | 0.52412 | 0.44614 | 0.34209 | 0.21096 | 0.19178 | 0.11186 | 0.066151 | 0.15757 | 0.28529 | 0.33757
I2 | A | 27 | 0.62654 | 0.50053 | 0.41685 | 0.31108 | 0.17999 | 0.15803 | 0.078853 | 0.061142 | 0.16959 | 0.29702 | 0.34863
I2 | A | 28 | 0.60207 | 0.47758 | 0.39733 | 0.29298 | 0.15823 | 0.13388 | 0.054164 | 0.070722 | 0.18883 | 0.31363 | 0.36205
I2 | A | 29 | 0.5888 | 0.46184 | 0.37454 | 0.26398 | 0.1351 | 0.11038 | 0.034136 | 0.093823 | 0.21232 | 0.33941 | 0.38778
I2 | A | 30 | 0.56006 | 0.43144 | 0.34465 | 0.24022 | 0.11148 | 0.084365 | 0.032552 | 0.11781 | 0.23742 | 0.35783 | 0.40081
I2 | A | 31 | 0.55149 | 0.42236 | 0.3283 | 0.22154 | 0.094775 | 0.070967 | 0.046722 | 0.14529 | 0.26311 | 0.38337 | 0.42912
I2 | A | 32 | 0.5219 | 0.39181 | 0.30919 | 0.20086 | 0.077554 | 0.054715 | 0.073171 | 0.1734 | 0.29337 | 0.41053 | 0.44771
I2 | B | 1 | 1.0728 | 0.83045 | 0.69847 | 0.58307 | 0.5009 | 0.50501 | 0.48027 | 0.50204 | 0.52566 | 0.57965 | 0.60875
I2 | B | 2 | 1.091 | 0.84045 | 0.70587 | 0.57844 | 0.49001 | 0.49335 | 0.4635 | 0.48176 | 0.49741 | 0.55022 | 0.58136
I2 | B | 3 | 1.1161 | 0.86282 | 0.72554 | 0.60204 | 0.49566 | 0.49412 | 0.46072 | 0.47176 | 0.48529 | 0.53617 | 0.56594
I2 | B | 4 | 1.1362 | 0.88086 | 0.73687 | 0.6058 | 0.49479 | 0.49581 | 0.45311 | 0.45641 | 0.46833 | 0.52065 | 0.5489
I2 | B | 5 | 1.1461 | 0.88992 | 0.74135 | 0.6052 | 0.49613 | 0.49051 | 0.44688 | 0.44599 | 0.45691 | 0.50818 | 0.53597
I2 | B | 6 | 1.1642 | 0.90334 | 0.75318 | 0.61689 | 0.49469 | 0.4922 | 0.44188 | 0.43677 | 0.44348 | 0.49305 | 0.52411
I2 | B | 7 | 1.1678 | 0.90714 | 0.75483 | 0.61066 | 0.49153 | 0.48342 | 0.43341 | 0.42571 | 0.43138 | 0.48196 | 0.51065
I2 | B | 8 | 1.174 | 0.90673 | 0.75244 | 0.61114 | 0.48144 | 0.47138 | 0.41975 | 0.41075 | 0.4112 | 0.46464 | 0.49576
I2 | B | 9 | 1.1879 | 0.91851 | 0.76484 | 0.62081 | 0.48557 | 0.47269 | 0.41904 | 0.40622 | 0.39962 | 0.45201 | 0.48264
I2 | B | 10 | 1.1868 | 0.91786 | 0.75879 | 0.61535 | 0.47601 | 0.45881 | 0.40563 | 0.38551 | 0.38931 | 0.44249 | 0.47422
I2 | B | 11 | 1.1899 | 0.91956 | 0.7573 | 0.6097 | 0.46876 | 0.45396 | 0.39195 | 0.36933 | 0.36751 | 0.42292 | 0.45556
I2 | B | 12 | 1.1894 | 0.92018 | 0.75712 | 0.60719 | 0.46296 | 0.44956 | 0.38468 | 0.36221 | 0.36215 | 0.41916 | 0.45137
I2 | B | 13 | 1.1896 | 0.91673 | 0.75036 | 0.60247 | 0.44709 | 0.43197 | 0.36949 | 0.34538 | 0.34329 | 0.40579 | 0.44429
I2 | B | 14 | 1.1858 | 0.91105 | 0.74708 | 0.59406 | 0.44342 | 0.42718 | 0.36095 | 0.33219 | 0.3352 | 0.39548 | 0.43272
I2 | B | 15 | 1.1821 | 0.90672 | 0.73807 | 0.58041 | 0.43354 | 0.40956 | 0.34314 | 0.31242 | 0.31635 | 0.37984 | 0.41318
I2 | B | 16 | 1.1773 | 0.90026 | 0.73009 | 0.57562 | 0.41576 | 0.39221 | 0.32629 | 0.29426 | 0.29974 | 0.36666 | 0.4048
I2 | B | 17 | 1.1649 | 0.8876 | 0.71854 | 0.55782 | 0.40475 | 0.3792 | 0.30806 | 0.26895 | 0.28746 | 0.35631 | 0.39273
I2 | B | 18 | 1.1586 | 0.87709 | 0.7074 | 0.55213 | 0.38623 | 0.36182 | 0.28843 | 0.25282 | 0.26548 | 0.33976 | 0.38046
I2 | B | 19 | 1.1435 | 0.86637 | 0.6888 | 0.52796 | 0.37045 | 0.33627 | 0.2649 | 0.23136 | 0.24488 | 0.32301 | 0.36547
I2 | B | 20 | 1.14 | 0.85624 | 0.68093 | 0.52668 | 0.35432 | 0.32407 | 0.2515 | 0.21468 | 0.23645 | 0.31975 | 0.3628
I2 | B | 21 | 1.1305 | 0.8424 | 0.66825 | 0.50714 | 0.3334 | 0.30485 | 0.22746 | 0.19185 | 0.22082 | 0.30816 | 0.35173
I2 | B | 22 | 1.1031 | 0.81697 | 0.64193 | 0.47819 | 0.30703 | 0.27141 | 0.19241 | 0.16007 | 0.19956 | 0.2941 | 0.33751
I2 | B | 23 | 1.086 | 0.80502 | 0.62618 | 0.46833 | 0.29176 | 0.25326 | 0.173 | 0.13754 | 0.19554 | 0.2919 | 0.33352
I2 | B | 24 | 1.0701 | 0.78717 | 0.60689 | 0.45209 | 0.27031 | 0.22792 | 0.14638 | 0.1165 | 0.18148 | 0.27948 | 0.32503
I2 | B | 25 | 1.0545 | 0.77219 | 0.59302 | 0.42613 | 0.25371 | 0.20466 | 0.12149 | 0.096495 | 0.17958 | 0.28266 | 0.32346
I2 | B | 26 | 1.0305 | 0.74798 | 0.5647 | 0.40505 | 0.2266 | 0.18472 | 0.09111 | 0.081127 | 0.18093 | 0.28248 | 0.32519
I2 | B | 27 | 1.0081 | 0.71986 | 0.54233 | 0.38182 | 0.20448 | 0.1513 | 0.062707 | 0.078378 | 0.18804 | 0.28349 | 0.3242
I2 | B | 28 | 0.98364 | 0.70288 | 0.52035 | 0.35921 | 0.18811 | 0.13246 | 0.048955 | 0.094462 | 0.20285 | 0.29663 | 0.33395
I2 | B | 29 | 0.95904 | 0.67727 | 0.49416 | 0.33773 | 0.174 | 0.11926 | 0.058722 | 0.12164 | 0.22066 | 0.31232 | 0.35201
I2 | B | 30 | 0.93363 | 0.65375 | 0.47504 | 0.32389 | 0.16795 | 0.11496 | 0.083176 | 0.14832 | 0.24727 | 0.32831 | 0.35987
I2 | B | 31 | 0.91342 | 0.63132 | 0.45771 | 0.30893 | 0.16562 | 0.11817 | 0.10784 | 0.18129 | 0.26907 | 0.34661 | 0.38283
I2 | B | 32 | 0.89016 | 0.61476 | 0.43839 | 0.29737 | 0.17578 | 0.12949 | 0.14246 | 0.21241 | 0.30344 | 0.37454 | 0.3971
M1 | A | 1 | 5.1239 | 5.2645 | 5.3483 | 5.5168 | 5.5414 | 5.5608 | 5.6332 | 5.697 | 5.6539 | 5.5037 | 5.3671
M1 | A | 2 | 4.8319 | 5.0031 | 5.1014 | 5.298 | 5.3519 | 5.3763 | 5.4577 | 5.5229 | 5.5343 | 5.4181 | 5.3002
M1 | A | 3 | 4.5455 | 4.7432 | 4.8554 | 5.0544 | 5.1543 | 5.1829 | 5.275 | 5.3728 | 5.4112 | 5.3308 | 5.2271
M1 | A | 4 | 4.2837 | 4.5056 | 4.623 | 4.8382 | 4.9685 | 5.0125 | 5.1141 | 5.2276 | 5.3224 | 5.2699 | 5.1825
M1 | A | 5 | 4.0119 | 4.2812 | 4.4227 | 4.6511 | 4.8006 | 4.8471 | 4.9631 | 5.0891 | 5.2189 | 5.209 | 5.1429
M1 | A | 6 | 3.7478 | 4.0412 | 4.218 | 4.4598 | 4.6202 | 4.7007 | 4.8123 | 4.9646 | 5.1274 | 5.153 | 5.1087
M1 | A | 7 | 3.5248 | 3.8479 | 4.0241 | 4.2762 | 4.4829 | 4.5631 | 4.6887 | 4.8555 | 5.0458 | 5.106 | 5.069
M1 | A | 8 | 3.2793 | 3.6208 | 3.8319 | 4.0991 | 4.3125 | 4.4283 | 4.5432 | 4.7346 | 4.9609 | 5.0563 | 5.0244
M1 | A | 9 | 3.0507 | 3.4333 | 3.6554 | 3.9226 | 4.1685 | 4.2965 | 4.4135 | 4.6271 | 4.8733 | 5.007 | 4.9973
M1 | A | 10 | 2.8285 | 3.2552 | 3.4826 | 3.747 | 4.0409 | 4.1557 | 4.2985 | 4.5191 | 4.8042 | 4.9661 | 4.9681
M1 | A | 11 | 2.6063 | 3.0576 | 3.3206 | 3.5889 | 3.8971 | 4.0411 | 4.1801 | 4.4119 | 4.7294 | 4.9273 | 4.9442
M1 | A | 12 | 2.4071 | 2.8767 | 3.1622 | 3.4501 | 3.7649 | 3.9218 | 4.063 | 4.3114 | 4.6586 | 4.8857 | 4.9141
M1 | A | 13 | 2.2163 | 2.7178 | 2.9865 | 3.3043 | 3.6457 | 3.8079 | 3.9528 | 4.2214 | 4.575 | 4.8456 | 4.8841
M1 | A | 14 | 2.0416 | 2.5678 | 2.843 | 3.1688 | 3.5227 | 3.6919 | 3.8424 | 4.1299 | 4.5162 | 4.8133 | 4.871
M1 | A | 15 | 1.8876 | 2.4297 | 2.7333 | 3.0001 | 3.4337 | 3.5981 | 3.7611 | 4.0389 | 4.4808 | 4.806 | 4.871
M1 | A | 16 | 1.733 | 2.2894 | 2.6302 | 2.8888 | 3.3398 | 3.5142 | 3.6817 | 3.9729 | 4.4358 | 4.7946 | 4.8739
M1 | B | 1 | 5.5677 | 5.7307 | 5.8619 | 6.0592 | 6.168 | 6.1906 | 6.3128 | 6.4041 | 6.4137 | 6.1632 | 5.9547
M1 | B | 2 | 5.2439 | 5.4443 | 5.6035 | 5.8251 | 5.96 | 6.001 | 6.1344 | 6.2642 | 6.3127 | 6.1307 | 5.9427
M1 | B | 3 | 4.9463 | 5.1639 | 5.3173 | 5.5898 | 5.7376 | 5.7814 | 5.9333 | 6.0885 | 6.2085 | 6.0962 | 5.9391
M1 | B | 4 | 4.6316 | 4.8865 | 5.0954 | 5.3486 | 5.544 | 5.6207 | 5.7821 | 5.9711 | 6.1352 | 6.0761 | 5.9391
M1 | B | 5 | 4.3535 | 4.6409 | 4.8669 | 5.1368 | 5.3408 | 5.4631 | 5.6223 | 5.8442 | 6.0456 | 6.0455 | 5.9415
M1 | B | 6 | 4.0859 | 4.4017 | 4.6245 | 4.9175 | 5.171 | 5.2803 | 5.4606 | 5.707 | 5.9592 | 6.0173 | 5.9361
M1 | B | 7 | 3.8138 | 4.1561 | 4.4081 | 4.7058 | 4.9825 | 5.1241 | 5.3091 | 5.5892 | 5.8823 | 6.0024 | 5.9513
M1 | B | 8 | 3.5543 | 3.9231 | 4.1976 | 4.4995 | 4.8072 | 4.9508 | 5.1525 | 5.4491 | 5.7939 | 5.9675 | 5.9376
M1 | B | 9 | 3.3149 | 3.711 | 3.9855 | 4.3019 | 4.6495 | 4.7924 | 5.0012 | 5.3187 | 5.7137 | 5.9448 | 5.933
M1 | B | 10 | 3.0563 | 3.4753 | 3.7709 | 4.1072 | 4.481 | 4.6568 | 4.8655 | 5.2209 | 5.6299 | 5.9011 | 5.9128
M1 | B | 11 | 2.7982 | 3.2523 | 3.5669 | 3.9173 | 4.3135 | 4.4942 | 4.7171 | 5.0943 | 5.546 | 5.8704 | 5.9092
M1 | B | 12 | 2.5474 | 3.0428 | 3.3783 | 3.7354 | 4.1496 | 4.3411 | 4.5753 | 4.9595 | 5.4555 | 5.8343 | 5.8859
M1 | B | 13 | 2.3037 | 2.8372 | 3.1655 | 3.5469 | 3.9945 | 4.1727 | 4.4206 | 4.8219 | 5.3824 | 5.8041 | 5.8735
M1 | B | 14 | 2.0675 | 2.6219 | 2.9811 | 3.3745 | 3.8427 | 4.0426 | 4.2889 | 4.7041 | 5.3001 | 5.7653 | 5.8552
M1 | B | 15 | 1.8617 | 2.4388 | 2.7849 | 3.2007 | 3.6847 | 3.914 | 4.1587 | 4.6137 | 5.2206 | 5.7452 | 5.8618
M1 | B | 16 | 1.6489 | 2.2474 | 2.612 | 3.0469 | 3.5361 | 3.7866 | 4.0395 | 4.5115 | 5.1517 | 5.7224 | 5.8595
M2 | A | 1 | 3.7216 | 3.7886 | 3.8529 | 3.9507 | 3.9007 | 3.8904 | 3.936 | 3.9508 | 3.8947 | 3.7856 | 3.689
M2 | A | 2 | 3.3907 | 3.4833 | 3.5552 | 3.6717 | 3.6557 | 3.6694 | 3.7092 | 3.7526 | 3.7255 | 3.6543 | 3.5806
M2 | A | 3 | 3.0889 | 3.2112 | 3.3094 | 3.4374 | 3.452 | 3.4851 | 3.5316 | 3.5949 | 3.6044 | 3.5662 | 3.5078
M2 | A | 4 | 2.8199 | 2.9854 | 3.0991 | 3.2474 | 3.2992 | 3.3446 | 3.3953 | 3.4718 | 3.5216 | 3.5153 | 3.4754
M2 | A | 5 | 2.5813 | 2.7831 | 2.9225 | 3.0799 | 3.1687 | 3.2327 | 3.2932 | 3.392 | 3.474 | 3.4955 | 3.4677
M2 | A | 6 | 2.3693 | 2.6109 | 2.7659 | 2.9315 | 3.0623 | 3.1433 | 3.2068 | 3.3261 | 3.4419 | 3.4895 | 3.4821
M2 | A | 7 | 2.1844 | 2.4605 | 2.6465 | 2.8228 | 2.9745 | 3.0749 | 3.1436 | 3.268 | 3.4252 | 3.5067 | 3.5032
M2 | A | 8 | 2.0379 | 2.3419 | 2.5353 | 2.7304 | 2.9102 | 3.0198 | 3.092 | 3.2386 | 3.4231 | 3.5334 | 3.5439
M2 | A | 9 | 1.9085 | 2.2437 | 2.4458 | 2.6414 | 2.8623 | 2.9706 | 3.0541 | 3.2102 | 3.4269 | 3.5636 | 3.5823
M2 | A | 10 | 1.8069 | 2.1627 | 2.3747 | 2.569 | 2.8213 | 2.9398 | 3.0272 | 3.194 | 3.4391 | 3.6022 | 3.63
M2 | A | 11 | 1.7405 | 2.1099 | 2.3108 | 2.5169 | 2.7893 | 2.9004 | 3.0009 | 3.1737 | 3.4547 | 3.6415 | 3.6731
M2 | A | 12 | 1.7138 | 2.0763 | 2.2961 | 2.4743 | 2.7846 | 2.9047 | 3.0017 | 3.1844 | 3.4792 | 3.687 | 3.7288
M2 | A | 13 | 1.7237 | 2.0744 | 2.2816 | 2.4653 | 2.7726 | 2.9043 | 2.9993 | 3.19 | 3.5032 | 3.734 | 3.7804
M2 | A | 14 | 1.7639 | 2.0928 | 2.2902 | 2.452 | 2.7831 | 2.9081 | 3.0079 | 3.199 | 3.5385 | 3.7872 | 3.8398
M2 | A | 15 | 1.8683 | 2.1663 | 2.3352 | 2.5062 | 2.8263 | 2.9448 | 3.0487 | 3.2419 | 3.5991 | 3.8709 | 3.9234
M2 | A | 16 | 2.0028 | 2.2683 | 2.4319 | 2.5755 | 2.8903 | 3.0107 | 3.1098 | 3.2984 | 3.6713 | 3.9642 | 4.0241
M2 | B | 1 | 4.0751 | 4.0812 | 4.1296 | 4.2273 | 4.1806 | 4.1796 | 4.2272 | 4.2527 | 4.1912 | 4.0418 | 3.9169
M2 | B | 2 | 3.6723 | 3.7264 | 3.7963 | 3.9154 | 3.921 | 3.9439 | 4.0002 | 4.064 | 4.0565 | 3.9654 | 3.8687
M2 | B | 3 | 3.3355 | 3.4281 | 3.5254 | 3.6664 | 3.6986 | 3.7418 | 3.8133 | 3.9032 | 3.9478 | 3.905 | 3.8338
M2 | B | 4 | 3.0525 | 3.1896 | 3.3039 | 3.4665 | 3.5417 | 3.6014 | 3.6801 | 3.7951 | 3.8893 | 3.885 | 3.8357
M2 | B | 5 | 2.8093 | 2.9885 | 3.1331 | 3.3143 | 3.4175 | 3.4971 | 3.5826 | 3.7139 | 3.8478 | 3.8856 | 3.8517
M2 | B | 6 | 2.5785 | 2.8049 | 2.9696 | 3.1585 | 3.3103 | 3.399 | 3.4955 | 3.6519 | 3.8239 | 3.9038 | 3.8889
M2 | B | 7 | 2.3665 | 2.6403 | 2.8224 | 3.0367 | 3.2208 | 3.3244 | 3.428 | 3.6031 | 3.8107 | 3.9243 | 3.921
M2 | B | 8 | 2.1735 | 2.4943 | 2.7086 | 2.9257 | 3.147 | 3.259 | 3.3722 | 3.5684 | 3.8076 | 3.9593 | 3.9755
M2 | B | 9 | 2.0093 | 2.3655 | 2.5922 | 2.8197 | 3.0731 | 3.2061 | 3.321 | 3.538 | 3.8065 | 3.9954 | 4.0241
M2 | B | 10 | 1.8498 | 2.2476 | 2.507 | 2.7327 | 3.0271 | 3.1543 | 3.2833 | 3.5022 | 3.8149 | 4.036 | 4.0777
M2 | B | 11 | 1.7185 | 2.1504 | 2.4101 | 2.6526 | 2.9778 | 3.1152 | 3.2461 | 3.4745 | 3.8243 | 4.0749 | 4.1223
M2 | B | 12 | 1.6185 | 2.0785 | 2.3467 | 2.5856 | 2.9398 | 3.079 | 3.2213 | 3.4692 | 3.8332 | 4.1187 | 4.1788
M2 | B | 13 | 1.5401 | 2.0103 | 2.3016 | 2.5407 | 2.9123 | 3.0603 | 3.2052 | 3.4545 | 3.8522 | 4.1572 | 4.2271
M2 | B | 14 | 1.4962 | 1.9734 | 2.2495 | 2.4959 | 2.8909 | 3.0399 | 3.1878 | 3.4514 | 3.8723 | 4.2075 | 4.2776
M2 | B | 15 | 1.5013 | 1.9675 | 2.2481 | 2.4896 | 2.8864 | 3.0472 | 3.1949 | 3.4637 | 3.9075 | 4.2721 | 4.3575
M2 | B | 16 | 1.5634 | 1.9984 | 2.2555 | 2.4791 | 2.8897 | 3.0504 | 3.1968 | 3.4717 | 3.9453 | 4.3419 | 4.4451
M3 | A | 1 | 2.8883 | 2.8527 | 2.8323 | 2.8467 | 2.7281 | 2.6944 | 2.6732 | 2.6306 | 2.53 | 2.3947 | 2.2745
M3 | A | 2 | 2.5672 | 2.5663 | 2.5633 | 2.5953 | 2.5271 | 2.509 | 2.4993 | 2.4846 | 2.4274 | 2.3319 | 2.2365
M3 | A | 3 | 2.3451 | 2.3917 | 2.4223 | 2.4786 | 2.454 | 2.4592 | 2.4634 | 2.475 | 2.4549 | 2.3932 | 2.3226
M3 | A | 4 | 2.1549 | 2.2519 | 2.3136 | 2.3949 | 2.4149 | 2.4411 | 2.4575 | 2.4963 | 2.5087 | 2.4784 | 2.4283
M3 | A | 5 | 1.9767 | 2.1143 | 2.206 | 2.3089 | 2.3649 | 2.4102 | 2.4378 | 2.5004 | 2.5431 | 2.5427 | 2.5081
M3 | A | 6 | 1.8098 | 1.9867 | 2.1033 | 2.2234 | 2.3159 | 2.3743 | 2.4148 | 2.4944 | 2.5724 | 2.6014 | 2.5802
M3 | A | 7 | 1.6601 | 1.8684 | 2.0057 | 2.1391 | 2.2642 | 2.3352 | 2.3857 | 2.4832 | 2.5908 | 2.6465 | 2.6382
M3 | A | 8 | 1.5227 | 1.7584 | 1.9108 | 2.0565 | 2.2153 | 2.2978 | 2.355 | 2.4723 | 2.6068 | 2.6917 | 2.6935
M3 | A | 9 | 1.4099 | 1.6641 | 1.825 | 1.9782 | 2.1621 | 2.2552 | 2.3206 | 2.4507 | 2.6152 | 2.7254 | 2.7368
M3 | A | 10 | 1.319 | 1.58 | 1.7475 | 1.901 | 2.1101 | 2.2081 | 2.2806 | 2.425 | 2.6164 | 2.7491 | 2.7653
M3 | A | 11 | 1.2454 | 1.504 | 1.668 | 1.8208 | 2.052 | 2.1557 | 2.2328 | 2.3875 | 2.6041 | 2.76 | 2.7808
M3 | A | 12 | 1.2027 | 1.4447 | 1.6022 | 1.7512 | 1.9943 | 2.1 | 2.1802 | 2.3449 | 2.582 | 2.7615 | 2.7857
M3 | A | 13 | 1.1847 | 1.3914 | 1.5387 | 1.6727 | 1.9259 | 2.0313 | 2.1175 | 2.2884 | 2.5444 | 2.744 | 2.7709
M3 | A | 14 | 1.207 | 1.3677 | 1.494 | 1.6054 | 1.865 | 1.9708 | 2.0543 | 2.2287 | 2.5045 | 2.7255 | 2.7526
M3 | A | 15 | 1.3014 | 1.401 | 1.4959 | 1.5781 | 1.827 | 1.9282 | 2.0103 | 2.1843 | 2.4771 | 2.7214 | 2.75
M3 | A | 16 | 1.5256 | 1.5666 | 1.6284 | 1.6659 | 1.8878 | 1.9822 | 2.0518 | 2.2172 | 2.5268 | 2.7884 | 2.8199
M3 | B | 1 | 2.9179 | 2.8552 | 2.8323 | 2.8554 | 2.7623 | 2.7381 | 2.743 | 2.7407 | 2.6811 | 2.5816 | 2.4695
M3 | B | 2 | 2.5583 | 2.5402 | 2.5467 | 2.5876 | 2.5494 | 2.547 | 2.5648 | 2.5923 | 2.5885 | 2.5391 | 2.4645
M3 | B | 3 | 2.3617 | 2.3941 | 2.4318 | 2.496 | 2.4991 | 2.517 | 2.5447 | 2.5951 | 2.6255 | 2.6113 | 2.5579
M3 | B | 4 | 2.2149 | 2.2957 | 2.3618 | 2.4484 | 2.4887 | 2.5261 | 2.5615 | 2.6334 | 2.6869 | 2.7 | 2.665
M3 | B | 5 | 2.0691 | 2.1926 | 2.2881 | 2.3947 | 2.4642 | 2.5193 | 2.5625 | 2.6535 | 2.7337 | 2.7677 | 2.7474
M3 | B | 6 | 1.9236 | 2.0917 | 2.2117 | 2.3377 | 2.4389 | 2.506 | 2.5594 | 2.664 | 2.7715 | 2.8307 | 2.8204
M3 | B | 7 | 1.7781 | 1.9841 | 2.127 | 2.2705 | 2.4046 | 2.4823 | 2.5445 | 2.6649 | 2.7973 | 2.8823 | 2.8831
M3 | B | 8 | 1.6401 | 1.8852 | 2.0511 | 2.2066 | 2.3725 | 2.4612 | 2.5312 | 2.665 | 2.8237 | 2.9314 | 2.9384
M3 | B | 9 | 1.5072 | 1.7888 | 1.9722 | 2.1402 | 2.3346 | 2.431 | 2.5086 | 2.6551 | 2.8389 | 2.9677 | 2.9823
M3 | B | 10 | 1.3815 | 1.6958 | 1.893 | 2.0706 | 2.2895 | 2.3937 | 2.4779 | 2.6378 | 2.8468 | 2.9976 | 3.0182
M3 | B | 11 | 1.2636 | 1.6029 | 1.8124 | 1.996 | 2.2408 | 2.3519 | 2.4428 | 2.6124 | 2.8419 | 3.0133 | 3.036
M3 | B | 12 | 1.1575 | 1.5144 | 1.7305 | 1.9198 | 2.1848 | 2.3017 | 2.3983 | 2.579 | 2.83 | 3.0202 | 3.0438
M3 | B | 13 | 1.057 | 1.4247 | 1.6481 | 1.8363 | 2.1195 | 2.2397 | 2.3421 | 2.5323 | 2.8 | 3.0075 | 3.0326
M3 | B | 14 | 0.98594 | 1.3413 | 1.5692 | 1.7515 | 2.0542 | 2.1767 | 2.2825 | 2.4806 | 2.7705 | 2.9961 | 3.0211
M3 | B | 15 | 0.96042 | 1.2885 | 1.5075 | 1.6819 | 1.991 | 2.1166 | 2.2262 | 2.4327 | 2.7467 | 2.9991 | 3.0241
M3 | B | 16 | 1.0899 | 1.3369 | 1.5307 | 1.6751 | 1.9838 | 2.1131 | 2.2221 | 2.4357 | 2.7783 | 3.059 | 3.0938
M4 | A | 1 | 3.609 | 3.9 | 4.1023 | 4.1621 | 4.1645 | 4.2193 | 4.192 | 4.178 | 4.0425 | 3.8834 | 3.7928
M4 | A | 2 | 3.2747 | 3.568 | 3.7664 | 3.865 | 3.8941 | 3.9688 | 3.9425 | 3.9588 | 3.8585 | 3.7129 | 3.6417
M4 | A | 3 | 2.9506 | 3.2597 | 3.4895 | 3.5618 | 3.6273 | 3.7129 | 3.7106 | 3.7489 | 3.6724 | 3.5542 | 3.4933
M4 | A | 4 | 2.6543 | 2.9478 | 3.201 | 3.2894 | 3.3707 | 3.473 | 3.4804 | 3.5327 | 3.4904 | 3.4022 | 3.3519
M4 | A | 5 | 2.3663 | 2.6508 | 2.9132 | 3.0188 | 3.1251 | 3.2392 | 3.2604 | 3.3314 | 3.322 | 3.2687 | 3.2307
M4 | A | 6 | 2.1233 | 2.3999 | 2.6596 | 2.7553 | 2.8879 | 3.0115 | 3.0458 | 3.1363 | 3.1558 | 3.1322 | 3.1094
M4 | A | 7 | 1.8684 | 2.134 | 2.415 | 2.5111 | 2.6212 | 2.7825 | 2.8202 | 2.9308 | 2.9676 | 2.9671 | 2.9771
M4 | A | 8 | 1.6796 | 1.9051 | 2.1568 | 2.2276 | 2.4107 | 2.5361 | 2.5912 | 2.6984 | 2.7832 | 2.8264 | 2.8138
M4 | A | 9 | 1.458 | 1.6966 | 1.8891 | 2.0217 | 2.1618 | 2.2993 | 2.3465 | 2.4807 | 2.5772 | 2.6287 | 2.6318
M4 | A | 10 | 1.2862 | 1.481 | 1.6328 | 1.7833 | 1.9086 | 2.0218 | 2.0905 | 2.2335 | 2.3619 | 2.4406 | 2.4353
M4 | A | 11 | 1.0921 | 1.2204 | 1.437 | 1.4599 | 1.6184 | 1.7706 | 1.8413 | 1.9654 | 2.1018 | 2.2151 | 2.2468
M4 | A | 12 | 0.93436 | 1.0298 | 1.1699 | 1.2366 | 1.3837 | 1.5358 | 1.578 | 1.7233 | 1.8602 | 1.9858 | 2.0353
M4 | A | 13 | 0.81805 | 0.89242 | 0.97116 | 1.0553 | 1.1686 | 1.3001 | 1.3476 | 1.5034 | 1.639 | 1.796 | 1.8318
M4 | A | 14 | 0.78528 | 0.79613 | 0.86101 | 0.86413 | 0.98393 | 1.0832 | 1.1505 | 1.2532 | 1.4491 | 1.6326 | 1.7217
M4 | A | 15 | 0.81854 | 0.81093 | 0.84271 | 0.88415 | 0.92132 | 0.96774 | 1.0474 | 1.1256 | 1.3452 | 1.5521 | 1.5872
M4 | A | 16 | 0.92402 | 0.93009 | 0.94458 | 0.98766 | 0.98776 | 1.0062 | 1.0719 | 1.1303 | 1.3301 | 1.5385 | 1.5866
M4 | B | 1 | 4.0299 | 4.2017 | 4.3335 | 4.3505 | 4.3168 | 4.3518 | 4.3112 | 4.2868 | 4.1475 | 3.9857 | 3.9013
M4 | B | 2 | 3.6836 | 3.8746 | 4.0185 | 4.0691 | 4.0542 | 4.1096 | 4.0713 | 4.0742 | 3.9686 | 3.8262 | 3.7527
M4 | B | 3 | 3.3438 | 3.5714 | 3.7307 | 3.7651 | 3.8004 | 3.8643 | 3.8497 | 3.8674 | 3.7867 | 3.6704 | 3.6111
M4 | B | 4 | 3.0111 | 3.2517 | 3.4473 | 3.4988 | 3.5473 | 3.6379 | 3.6317 | 3.6691 | 3.6183 | 3.5252 | 3.4744
M4 | B | 5 | 2.6967 | 2.9534 | 3.1621 | 3.2317 | 3.3141 | 3.4143 | 3.4216 | 3.4796 | 3.4583 | 3.3927 | 3.3563
M4 | B | 6 | 2.4145 | 2.6844 | 2.9038 | 2.9834 | 3.0888 | 3.198 | 3.2176 | 3.2968 | 3.297 | 3.2682 | 3.2345
M4 | B | 7 | 2.1111 | 2.3828 | 2.6472 | 2.736 | 2.8328 | 2.9778 | 3.001 | 3.0997 | 3.1205 | 3.121 | 3.1141
M4 | B | 8 | 1.8714 | 2.1398 | 2.3772 | 2.4482 | 2.6146 | 2.7315 | 2.7759 | 2.8756 | 2.9491 | 2.97 | 2.9558
M4 | B | 9 | 1.5654 | 1.8768 | 2.1041 | 2.2299 | 2.3618 | 2.4879 | 2.5337 | 2.6592 | 2.7406 | 2.7899 | 2.7838
M4 | B | 10 | 1.3166 | 1.6161 | 1.7995 | 1.9632 | 2.1054 | 2.2187 | 2.276 | 2.419 | 2.5369 | 2.6019 | 2.5925
M4 | B | 11 | 1.0547 | 1.2973 | 1.5645 | 1.6071 | 1.8068 | 1.9427 | 2.0199 | 2.1384 | 2.2896 | 2.3905 | 2.4055
M4 | B | 12 | 0.79028 | 1.0313 | 1.248 | 1.3526 | 1.5334 | 1.6914 | 1.7413 | 1.893 | 2.0318 | 2.1589 | 2.2085
M4 | B | 13 | 0.55533 | 0.79239 | 0.97532 | 1.1033 | 1.2675 | 1.4349 | 1.4805 | 1.662 | 1.8022 | 1.9642 | 2.0094
M4 | B | 14 | 0.42067 | 0.58458 | 0.76422 | 0.84382 | 1.0376 | 1.1453 | 1.2494 | 1.3824 | 1.607 | 1.8122 | 1.8998
M4 | B | 15 | 0.45486 | 0.53636 | 0.6523 | 0.74805 | 0.88151 | 0.96921 | 1.0828 | 1.219 | 1.4704 | 1.7039 | 1.7567
M4 | B | 16 | 0.66656 | 0.65806 | 0.71541 | 0.81262 | 0.86481 | 0.93677 | 1.0384 | 1.1594 | 1.4182 | 1.6785 | 1.7333
M5 | A | 1 | 1.8938 | 1.862 | 1.8956 | 1.8636 | 1.8208 | 1.8467 | 1.8026 | 1.7788 | 1.6539 | 1.531 | 1.4937
M5 | A | 2 | 1.881 | 1.8311 | 1.852 | 1.8194 | 1.7907 | 1.8202 | 1.7872 | 1.7687 | 1.6862 | 1.5727 | 1.5309
M5 | A | 3 | 1.8847 | 1.8224 | 1.8383 | 1.8083 | 1.7791 | 1.8185 | 1.7967 | 1.7982 | 1.7347 | 1.6423 | 1.5983
M5 | A | 4 | 1.898 | 1.8176 | 1.825 | 1.7914 | 1.7692 | 1.8157 | 1.8033 | 1.8152 | 1.7772 | 1.7124 | 1.681
M5 | A | 5 | 1.9 | 1.8017 | 1.8075 | 1.7653 | 1.7535 | 1.8097 | 1.7985 | 1.8274 | 1.8148 | 1.7756 | 1.7552
M5 | A | 6 | 1.8759 | 1.7589 | 1.7563 | 1.7133 | 1.7047 | 1.7696 | 1.7608 | 1.7993 | 1.811 | 1.7959 | 1.793
M5 | A | 7 | 1.8298 | 1.7012 | 1.6828 | 1.6477 | 1.641 | 1.6995 | 1.7011 | 1.7377 | 1.7796 | 1.7916 | 1.7993
M5 | A | 8 | 1.7511 | 1.6088 | 1.5836 | 1.545 | 1.529 | 1.5992 | 1.6046 | 1.6669 | 1.71 | 1.7503 | 1.7662
M5 | A | 9 | 1.655 | 1.4922 | 1.4659 | 1.4171 | 1.4019 | 1.4781 | 1.4849 | 1.5556 | 1.6169 | 1.6863 | 1.7103
M5 | A | 10 | 1.541 | 1.3629 | 1.3006 | 1.2578 | 1.2557 | 1.3094 | 1.3328 | 1.4016 | 1.5006 | 1.5821 | 1.6034
M5 | A | 11 | 1.3837 | 1.1892 | 1.1144 | 1.0676 | 1.0596 | 1.1217 | 1.147 | 1.2317 | 1.3345 | 1.4344 | 1.472
M5 | A | 12 | 1.2179 | 0.99351 | 0.91376 | 0.84263 | 0.84863 | 0.91838 | 0.94785 | 1.0387 | 1.1605 | 1.2777 | 1.3378
M5 | A | 13 | 1.0271 | 0.78538 | 0.66649 | 0.59694 | 0.61921 | 0.67537 | 0.72422 | 0.8315 | 0.96896 | 1.0939 | 1.1419
M5 | A | 14 | 0.80932 | 0.54052 | 0.41363 | 0.32978 | 0.37277 | 0.4531 | 0.50669 | 0.6361 | 0.76991 | 0.89303 | 0.9418
M5 | A | 15 | 0.61487 | 0.33395 | 0.13764 | 0.058679 | 0.21485 | 0.2965 | 0.37598 | 0.51966 | 0.63139 | 0.72354 | 0.7615
M5 | A | 16 | 0.47961 | 0.26874 | 0.20474 | 0.27411 | 0.3749 | 0.39593 | 0.45844 | 0.55829 | 0.61317 | 0.63562 | 0.64849
M5 | B | 1 | 2.0999 | 2.1688 | 2.2068 | 2.1907 | 2.1334 | 2.1374 | 2.0718 | 2.0081 | 1.8385 | 1.6432 | 1.5587
M5 | B | 2 | 1.9139 | 2.0018 | 2.0852 | 2.0821 | 2.0482 | 2.078 | 2.036 | 2.0009 | 1.8717 | 1.708 | 1.625
M5 | B | 3 | 1.7638 | 1.8766 | 1.9876 | 1.9964 | 2.0005 | 2.0491 | 2.0233 | 2.0212 | 1.9308 | 1.7967 | 1.7368
M5 | B | 4 | 1.6326 | 1.758 | 1.8891 | 1.92 | 1.9442 | 2.0103 | 2.004 | 2.0273 | 1.979 | 1.8851 | 1.8349
M5 | B | 5 | 1.5164 | 1.6552 | 1.7923 | 1.8329 | 1.8902 | 1.964 | 1.9746 | 2.0221 | 2.0161 | 1.9537 | 1.919
M5 | B | 6 | 1.399 | 1.5388 | 1.6757 | 1.73 | 1.809 | 1.8899 | 1.9172 | 1.9854 | 2.0082 | 1.989 | 1.9641
M5 | B | 7 | 1.259 | 1.3954 | 1.5426 | 1.6144 | 1.6991 | 1.7933 | 1.8349 | 1.9231 | 1.9743 | 1.9851 | 1.9732
M5 | B | 8 | 1.1085 | 1.234 | 1.4 | 1.4714 | 1.5683 | 1.6822 | 1.7278 | 1.8324 | 1.9118 | 1.95 | 1.9546
M5 | B | 9 | 0.95626 | 1.0652 | 1.2295 | 1.3033 | 1.4253 | 1.539 | 1.5939 | 1.715 | 1.8169 | 1.8844 | 1.9026
M5 | B | 10 | 0.79356 | 0.87715 | 1.0287 | 1.1182 | 1.243 | 1.3564 | 1.4289 | 1.5707 | 1.6847 | 1.7735 | 1.803
M5 | B | 11 | 0.60435 | 0.66597 | 0.81117 | 0.9116 | 1.0487 | 1.1671 | 1.2421 | 1.3938 | 1.5224 | 1.637 | 1.6741
M5 | B | 12 | 0.42512 | 0.43146 | 0.58 | 0.68146 | 0.8456 | 0.96701 | 1.0514 | 1.2044 | 1.3599 | 1.4819 | 1.5277
M5 | B | 13 | 0.28413 | 0.17732 | 0.35041 | 0.46367 | 0.64452 | 0.77037 | 0.85591 | 1.0123 | 1.1655 | 1.2857 | 1.3278
M5 | B | 14 | 0.30045 | 0.092036 | 0.21486 | 0.36938 | 0.51662 | 0.61633 | 0.69754 | 0.84506 | 0.97173 | 1.0795 | 1.119
M5 | B | 15 | 0.47836 | 0.3732 | 0.37421 | 0.4358 | 0.52572 | 0.58488 | 0.63503 | 0.74759 | 0.81926 | 0.88078 | 0.90473
M5 | B | 16 | 0.7309 | 0.67969 | 0.64928 | 0.68258 | 0.70021 | 0.70103 | 0.71813 | 0.76571 | 0.75542 | 0.73215 | 0.72987
M6 | A | 1 | 1.0724 | 0.79949 | 0.65388 | 0.53142 | 0.45251 | 0.46999 | 0.44737 | 0.47381 | 0.49594 | 0.53647 | 0.55271
M6 | A | 2 | 1.2064 | 0.92773 | 0.77188 | 0.63836 | 0.5404 | 0.54772 | 0.51404 | 0.51713 | 0.52235 | 0.5526 | 0.57102
M6 | A | 3 | 1.3069 | 1.0236 | 0.86414 | 0.72855 | 0.62104 | 0.61562 | 0.57646 | 0.56098 | 0.55868 | 0.57933 | 0.59535
M6 | A | 4 | 1.3994 | 1.1139 | 0.95693 | 0.81245 | 0.69805 | 0.69483 | 0.64788 | 0.62724 | 0.61402 | 0.6314 | 0.64119
M6 | A | 5 | 1.4927 | 1.2056 | 1.0402 | 0.89775 | 0.77763 | 0.77467 | 0.71819 | 0.68489 | 0.67562 | 0.68636 | 0.70103
M6 | A | 6 | 1.5569 | 1.267 | 1.1033 | 0.95376 | 0.82659 | 0.825 | 0.76658 | 0.73332 | 0.7103 | 0.72633 | 0.74257
M6 | A | 7 | 1.5982 | 1.3014 | 1.1327 | 0.98283 | 0.86148 | 0.84826 | 0.78986 | 0.74643 | 0.73368 | 0.74382 | 0.75815
M6 | A | 8 | 1.6127 | 1.3172 | 1.142 | 0.99193 | 0.86284 | 0.84818 | 0.78828 | 0.74162 | 0.72512 | 0.73538 | 0.75201
M6 | A | 9 | 1.6141 | 1.3177 | 1.1413 | 0.98581 | 0.84783 | 0.83741 | 0.77541 | 0.72844 | 0.7093 | 0.72481 | 0.73588
M6 | A | 10 | 1.5996 | 1.2948 | 1.1189 | 0.95719 | 0.81943 | 0.80422 | 0.74025 | 0.68981 | 0.67426 | 0.6893 | 0.70682
M6 | A | 11 | 1.5497 | 1.2462 | 1.0616 | 0.90201 | 0.75033 | 0.73202 | 0.6694 | 0.61715 | 0.59476 | 0.61378 | 0.63159
M6 | A | 12 | 1.4897 | 1.1852 | 0.99546 | 0.83456 | 0.67666 | 0.65007 | 0.58775 | 0.5236 | 0.50662 | 0.52691 | 0.55266
M6 | A | 13 | 1.4147 | 1.104 | 0.91207 | 0.75218 | 0.59066 | 0.5566 | 0.48648 | 0.41559 | 0.3985 | 0.42357 | 0.44774
M6 | A | 14 | 1.3187 | 1.0148 | 0.81754 | 0.652 | 0.48824 | 0.44479 | 0.36858 | 0.28361 | 0.26572 | 0.29893 | 0.32204
M6 | A | 15 | 1.2291 | 0.92479 | 0.72094 | 0.5573 | 0.39384 | 0.34328 | 0.25879 | 0.14554 | 0.12633 | 0.17013 | 0.20628
M6 | A | 16 | 1.1267 | 0.83411 | 0.63862 | 0.48885 | 0.33093 | 0.26344 | 0.19888 | 0.071947 | 0.039261 | 0.093326 | 0.12619
M6 | B | 1 | 0.7316 | 0.67959 | 0.70414 | 0.69146 | 0.69582 | 0.72486 | 0.715 | 0.72886 | 0.68797 | 0.62123 | 0.59067
M6 | B | 2 | 0.79929 | 0.73002 | 0.74112 | 0.71917 | 0.72171 | 0.75586 | 0.74879 | 0.76824 | 0.73586 | 0.68183 | 0.6566
M6 | B | 3 | 0.85531 | 0.77195 | 0.77567 | 0.74783 | 0.74633 | 0.78451 | 0.78046 | 0.80626 | 0.78037 | 0.73781 | 0.71778
M6 | B | 4 | 0.9195 | 0.8228 | 0.82223 | 0.78976 | 0.78957 | 0.83151 | 0.82844 | 0.85362 | 0.83966 | 0.80553 | 0.7911
M6 | B | 5 | 0.97494 | 0.87725 | 0.87183 | 0.83791 | 0.83159 | 0.87676 | 0.87341 | 0.90384 | 0.90072 | 0.87246 | 0.85872
M6 | B | 6 | 1.0105 | 0.9072 | 0.8941 | 0.8603 | 0.85617 | 0.90824 | 0.90108 | 0.93714 | 0.93822 | 0.92119 | 0.91137
M6 | B | 7 | 1.0305 | 0.91787 | 0.90155 | 0.86457 | 0.86476 | 0.9092 | 0.91084 | 0.94487 | 0.95485 | 0.94654 | 0.93563
M6 | B | 8 | 1.0228 | 0.90493 | 0.88127 | 0.84367 | 0.84182 | 0.8898 | 0.89351 | 0.92792 | 0.94759 | 0.94626 | 0.94168
M6 | B | 9 | 1.0036 | 0.87663 | 0.8527 | 0.81066 | 0.81059 | 0.86461 | 0.86349 | 0.90604 | 0.92704 | 0.93472 | 0.93028
M6 | B | 10 | 0.96674 | 0.83516 | 0.80401 | 0.76043 | 0.76309 | 0.81412 | 0.81678 | 0.85787 | 0.88995 | 0.90398 | 0.90382
M6 | B | 11 | 0.89674 | 0.75876 | 0.7241 | 0.67358 | 0.68157 | 0.73237 | 0.73692 | 0.7855 | 0.81864 | 0.83603 | 0.83463
M6 | B | 12 | 0.82427 | 0.67468 | 0.62991 | 0.5788 | 0.58415 | 0.63344 | 0.6439 | 0.69158 | 0.73143 | 0.75409 | 0.75432
M6 | B | 13 | 0.72724 | 0.56822 | 0.52339 | 0.46835 | 0.46692 | 0.51729 | 0.52815 | 0.58088 | 0.62163 | 0.64652 | 0.65708
M6 | B | 14 | 0.62727 | 0.45345 | 0.39137 | 0.33303 | 0.33159 | 0.38333 | 0.39361 | 0.44992 | 0.48833 | 0.51665 | 0.52865
M6 | B | 15 | 0.53921 | 0.3506 | 0.26025 | 0.19531 | 0.18646 | 0.24005 | 0.25778 | 0.31723 | 0.35558 | 0.38252 | 0.3946
M6 | B | 16 | 0.44882 | 0.26019 | 0.1397 | 0.045857 | 0.056025 | 0.10794 | 0.12952 | 0.19895 | 0.21326 | 0.22827 | 0.23595
M7 | A | 1 | 0.5382 | 0.43894 | 0.38515 | 0.32722 | 0.29772 | 0.32293 | 0.32376 | 0.38494 | 0.4419 | 0.54856 | 0.60016
M7 | A | 2 | 0.59388 | 0.4897 | 0.4367 | 0.36619 | 0.31437 | 0.33088 | 0.31798 | 0.3556 | 0.40372 | 0.50555 | 0.55378
M7 | A | 3 | 0.63763 | 0.53423 | 0.47341 | 0.40005 | 0.329 | 0.34354 | 0.31544 | 0.3364 | 0.37326 | 0.46879 | 0.51869
M7 | A | 4 | 0.67191 | 0.56743 | 0.50117 | 0.42629 | 0.34668 | 0.35451 | 0.31885 | 0.32573 | 0.34792 | 0.44132 | 0.48984
M7 | A | 5 | 0.6999 | 0.59009 | 0.52666 | 0.44553 | 0.35651 | 0.36474 | 0.31598 | 0.30895 | 0.32107 | 0.40862 | 0.45907
M7 | A | 6 | 0.71602 | 0.61012 | 0.54305 | 0.45995 | 0.36098 | 0.36479 | 0.31465 | 0.29477 | 0.29897 | 0.38223 | 0.43535
M7 | A | 7 | 0.72611 | 0.61956 | 0.54877 | 0.45922 | 0.36282 | 0.36267 | 0.30638 | 0.2809 | 0.27951 | 0.36323 | 0.41497
M7 | A | 8 | 0.7297 | 0.62019 | 0.54978 | 0.45931 | 0.35617 | 0.3546 | 0.29307 | 0.26063 | 0.25186 | 0.33358 | 0.38263
M7 | A | 9 | 0.72666 | 0.61988 | 0.54744 | 0.45473 | 0.34548 | 0.33654 | 0.27297 | 0.23002 | 0.22321 | 0.31293 | 0.36209
M7 | A | 10 | 0.71777 | 0.60716 | 0.53218 | 0.43501 | 0.32488 | 0.31072 | 0.24835 | 0.20051 | 0.19859 | 0.29447 | 0.34447
M7 | A | 11 | 0.70073 | 0.58695 | 0.5103 | 0.41261 | 0.29582 | 0.28773 | 0.21551 | 0.16708 | 0.1681 | 0.27614 | 0.33152
M7 | A | 12 | 0.67714 | 0.56636 | 0.48404 | 0.38452 | 0.26414 | 0.24902 | 0.17789 | 0.12906 | 0.15059 | 0.27209 | 0.32426
M7 | A | 13 | 0.65017 | 0.53519 | 0.4524 | 0.35146 | 0.22747 | 0.21161 | 0.13437 | 0.088144 | 0.15067 | 0.278 | 0.32954
M7 | A | 14 | 0.62017 | 0.49953 | 0.41621 | 0.31352 | 0.18594 | 0.16351 | 0.085338 | 0.066323 | 0.16892 | 0.29827 | 0.3429
M7 | A | 15 | 0.57926 | 0.45644 | 0.36412 | 0.26211 | 0.14007 | 0.11283 | 0.032777 | 0.090798 | 0.21289 | 0.33676 | 0.38075
M7 | A | 16 | 0.52872 | 0.40568 | 0.31371 | 0.21084 | 0.091707 | 0.060363 | 0.042643 | 0.14439 | 0.26498 | 0.38259 | 0.42149
M7 | B | 1 | 1.0684 | 0.82287 | 0.68166 | 0.5621 | 0.46325 | 0.46631 | 0.43785 | 0.45492 | 0.46991 | 0.52296 | 0.55678
M7 | B | 2 | 1.1123 | 0.85952 | 0.71425 | 0.58565 | 0.47409 | 0.46885 | 0.42941 | 0.43223 | 0.44099 | 0.492 | 0.52314
M7 | B | 3 | 1.1429 | 0.88534 | 0.73207 | 0.59955 | 0.47468 | 0.47372 | 0.42183 | 0.4162 | 0.41958 | 0.47024 | 0.50209
M7 | B | 4 | 1.1634 | 0.90266 | 0.74504 | 0.6098 | 0.47552 | 0.46681 | 0.41415 | 0.40188 | 0.40118 | 0.45094 | 0.4837
M7 | B | 5 | 1.1775 | 0.91125 | 0.75173 | 0.60845 | 0.47089 | 0.46237 | 0.39977 | 0.38309 | 0.37508 | 0.42705 | 0.45698
M7 | B | 6 | 1.1819 | 0.91657 | 0.7509 | 0.60857 | 0.46125 | 0.45092 | 0.38406 | 0.35953 | 0.35517 | 0.40821 | 0.44617
M7 | B | 7 | 1.1772 | 0.90824 | 0.74439 | 0.5969 | 0.45161 | 0.42704 | 0.36441 | 0.33558 | 0.33643 | 0.3955 | 0.42756
M7 | B | 8 | 1.1708 | 0.89884 | 0.73142 | 0.57964 | 0.4279 | 0.40616 | 0.33562 | 0.30651 | 0.30292 | 0.36507 | 0.40147
M7 | B | 9 | 1.1554 | 0.88365 | 0.70851 | 0.55638 | 0.40255 | 0.37677 | 0.30527 | 0.26844 | 0.27721 | 0.34609 | 0.38398
M7 | B | 10 | 1.1344 | 0.8598 | 0.6859 | 0.52989 | 0.37071 | 0.33816 | 0.26716 | 0.2263 | 0.24782 | 0.32478 | 0.36069
M7 | B | 11 | 1.1076 | 0.83254 | 0.65344 | 0.49565 | 0.33327 | 0.298 | 0.22153 | 0.18781 | 0.21434 | 0.29915 | 0.34123
M7 | B | 12 | 1.0766 | 0.7979 | 0.61977 | 0.46092 | 0.29556 | 0.25383 | 0.17436 | 0.1413 | 0.19075 | 0.28638 | 0.32935
M7 | B | 13 | 1.0386 | 0.76007 | 0.57715 | 0.41896 | 0.25386 | 0.20788 | 0.12151 | 0.1005 | 0.18291 | 0.28225 | 0.32175
M7 | B | 14 | 0.99813 | 0.71799 | 0.53764 | 0.38219 | 0.20479 | 0.15665 | 0.066092 | 0.083628 | 0.18605 | 0.28739 | 0.32231
M7 | B | 15 | 0.9501 | 0.67415 | 0.49117 | 0.33647 | 0.17065 | 0.11612 | 0.047204 | 0.11581 | 0.22109 | 0.31178 | 0.34812
M7 | B | 16 | 0.90012 | 0.62622 | 0.44972 | 0.30117 | 0.15983 | 0.10356 | 0.099044 | 0.17261 | 0.27165 | 0.35036 | 0.37774

## Polarization Phase
Band | HAM | Band Median | [10 empty] | Band Average
|  | Polarization Phase | [10 empty] | Polarization Phase
|  | Scan Angle | [10 empty] | Scan Angle
|  | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55 | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55
I1 | A | 2.9791 | 3.1185 | 3.307 | 3.4519 | 3.5413 | 3.5556 | 3.5638 | 3.5871 | 3.4846 | 3.3521 | 3.3084 | 3.0508 | 3.2086 | 3.3747 | 3.498 | 3.5776 | 3.5869 | 3.5917 | 3.616 | 3.505 | 3.3759 | 3.3276
I1 | B | 4.2796 | 4.2025 | 4.0984 | 4.0632 | 3.951 | 3.8668 | 3.7974 | 3.7056 | 3.5151 | 3.3041 | 3.227 | 4.3561 | 4.2564 | 4.1405 | 4.1022 | 3.9806 | 3.8944 | 3.8218 | 3.7272 | 3.5306 | 3.3208 | 3.2414
I2 | A | 3.0275 | 3.0455 | 3.1111 | 3.1564 | 3.228 | 3.3207 | 3.4343 | 4.0817 | 4.4378 | 4.6207 | 4.6197 | 3.0297 | 3.0547 | 3.1288 | 3.1779 | 3.221 | 3.3002 | 3.257 | 4.3824 | 4.7326 | 4.8492 | 4.8277
I2 | B | 2.8759 | 2.875 | 2.9292 | 2.9723 | 3.0755 | 3.1965 | 3.3499 | 3.9368 | 4.2347 | 4.3913 | 4.414 | 2.8675 | 2.8588 | 2.9009 | 2.9186 | 2.9426 | 3.0274 | 3.0304 | 4.2916 | 4.5154 | 4.5921 | 4.5872
M1 | A | 3.893 | 3.7714 | 3.6966 | 3.6601 | 3.5763 | 3.5262 | 3.5115 | 3.4525 | 3.3576 | 3.256 | 3.2131 | 3.8498 | 3.744 | 3.6792 | 3.6492 | 3.5742 | 3.5244 | 3.5125 | 3.4567 | 3.3683 | 3.2708 | 3.2309
M1 | B | 4.1401 | 3.9685 | 3.8657 | 3.8018 | 3.6919 | 3.6298 | 3.5954 | 3.5056 | 3.3707 | 3.2113 | 3.1442 | 4.1294 | 3.9581 | 3.8566 | 3.7971 | 3.693 | 3.631 | 3.5993 | 3.512 | 3.3821 | 3.229 | 3.1646
M2 | A | 3.5417 | 3.4674 | 3.4235 | 3.412 | 3.3404 | 3.2998 | 3.2954 | 3.2524 | 3.1807 | 3.1015 | 3.0684 | 3.463 | 3.4291 | 3.4006 | 3.3981 | 3.3476 | 3.3129 | 3.3107 | 3.2743 | 3.214 | 3.1438 | 3.1135
M2 | B | 3.7752 | 3.6314 | 3.5521 | 3.5176 | 3.4244 | 3.3758 | 3.3578 | 3.2983 | 3.2022 | 3.0919 | 3.0442 | 3.6699 | 3.5838 | 3.5249 | 3.5021 | 3.4296 | 3.3856 | 3.3719 | 3.3171 | 3.2321 | 3.1301 | 3.086
M3 | A | 3.4861 | 3.4437 | 3.4067 | 3.4037 | 3.3409 | 3.2962 | 3.2933 | 3.2462 | 3.1803 | 3.0984 | 3.0642 | 3.3588 | 3.3672 | 3.3555 | 3.3696 | 3.3354 | 3.2984 | 3.3025 | 3.2656 | 3.2116 | 3.1384 | 3.1056
M3 | B | 3.7157 | 3.5893 | 3.5111 | 3.4826 | 3.3981 | 3.3452 | 3.3322 | 3.2721 | 3.1912 | 3.0916 | 3.0497 | 3.5996 | 3.544 | 3.4923 | 3.4789 | 3.4139 | 3.3659 | 3.3569 | 3.3012 | 3.2262 | 3.1296 | 3.0862
M4 | A | 3.4994 | 3.5647 | 3.5784 | 3.5868 | 3.5752 | 3.5517 | 3.5332 | 3.5057 | 3.4427 | 3.3643 | 3.3386 | 3.3208 | 3.3731 | 3.3975 | 3.4079 | 3.42 | 3.4116 | 3.4035 | 3.3892 | 3.3486 | 3.287 | 3.2645
M4 | B | 3.7696 | 3.7306 | 3.683 | 3.6597 | 3.6151 | 3.58 | 3.5493 | 3.5046 | 3.4297 | 3.3428 | 3.3122 | 3.4914 | 3.5183 | 3.5096 | 3.5028 | 3.4912 | 3.4684 | 3.4479 | 3.4162 | 3.3564 | 3.2792 | 3.2507
M5 | A | 3.0378 | 3.1376 | 3.2281 | 3.4379 | 3.5408 | 3.5584 | 3.5707 | 3.6146 | 3.5911 | 3.5308 | 3.5043 | 2.9868 | 3.0375 | 3.0952 | 3.6253 | 3.7305 | 3.7298 | 3.7507 | 3.767 | 3.7207 | 3.6529 | 3.6258
M5 | B | 3.3263 | 3.7597 | 3.7727 | 3.8407 | 3.8169 | 3.7798 | 3.7559 | 3.711 | 3.6082 | 3.4831 | 3.4355 | 2.9022 | 4.1175 | 4.0461 | 4.0331 | 3.9789 | 3.9228 | 3.8858 | 3.8239 | 3.7106 | 3.5818 | 3.5333
M6 | A | 2.9122 | 2.939 | 2.9983 | 3.0408 | 3.131 | 3.1893 | 3.2517 | 3.3993 | 3.5256 | 3.6834 | 3.733 | 2.8594 | 2.8755 | 2.9294 | 2.9671 | 3.0524 | 3.1198 | 3.1804 | 3.3399 | 3.7893 | 3.8511 | 3.8717
M6 | B | 3.1419 | 3.2789 | 3.405 | 3.4849 | 3.634 | 3.6408 | 3.6744 | 3.7162 | 3.6353 | 3.5295 | 3.4697 | 3.0948 | 3.2305 | 3.3824 | 3.4969 | 3.7381 | 3.7415 | 3.7702 | 3.8025 | 3.7082 | 3.591 | 3.5458
M7 | A | 3.0227 | 3.0353 | 3.1013 | 3.1412 | 3.2096 | 3.3045 | 3.4138 | 4.0181 | 4.4113 | 4.5987 | 4.6019 | 3.0162 | 3.0382 | 3.111 | 3.1572 | 3.2161 | 3.3085 | 3.2938 | 4.3021 | 4.6739 | 4.809 | 4.791
M7 | B | 2.8735 | 2.8754 | 2.9237 | 2.966 | 3.0717 | 3.1911 | 3.3322 | 3.8856 | 4.1966 | 4.3659 | 4.3938 | 2.8639 | 2.8571 | 2.8973 | 2.92 | 2.9611 | 3.0556 | 3.0702 | 4.2081 | 4.4562 | 4.5479 | 4.5475
Band | HAM | Detector | Detector Dependent | [11 empty] | Notes:
|  |  | Polarization Phase | [11 empty] | 1. This table provides VIIRS F2 Polarization phase [in radians].
|  |  | Scan Angle | [11 empty] | 2. The results were derived from FP-11 and FP-11' using the SIS-100-2 source.
|  |  | -55 | -45 | -37 | -30 | -20 | -15 | -8 | 4 | 22 | 45 | 55 |  | 3. The bvonir polarization sheet was used with (M1-M3) and without (M4-M7 and I1-I2) blocking filter.
I1 | A | 1 | 3.607 | 3.7253 | 3.8005 | 3.8686 | 3.8733 | 3.8365 | 3.8315 | 3.8194 | 3.7125 | 3.5953 | 3.5406 |  | 4. Fourier analysis was performed to determine the two-cycle polarization phase.
I1 | A | 2 | 3.5499 | 3.6702 | 3.7539 | 3.8204 | 3.8334 | 3.7955 | 3.7899 | 3.7802 | 3.6688 | 3.5544 | 3.4931 |  | 5. The testing was conducted at Raytheon El Segundo, December 2013 / January 2014 (FP-11) and December 2014 (FP-11') .
I1 | A | 3 | 3.4926 | 3.6184 | 3.7071 | 3.7707 | 3.7886 | 3.7558 | 3.7517 | 3.7416 | 3.6346 | 3.5153 | 3.4592 |  | 6. The test was set to diagnostic mode.
I1 | A | 4 | 3.4435 | 3.5675 | 3.6635 | 3.7314 | 3.7506 | 3.7202 | 3.7163 | 3.71 | 3.6039 | 3.4861 | 3.4301
I1 | A | 5 | 3.3948 | 3.5209 | 3.6223 | 3.6929 | 3.7141 | 3.6877 | 3.6847 | 3.6821 | 3.5775 | 3.4569 | 3.3942
I1 | A | 6 | 3.3454 | 3.4746 | 3.5801 | 3.6523 | 3.6772 | 3.6586 | 3.6544 | 3.6536 | 3.5519 | 3.4295 | 3.3705
I1 | A | 7 | 3.3 | 3.4317 | 3.5416 | 3.6155 | 3.6414 | 3.6281 | 3.6245 | 3.6269 | 3.5278 | 3.4059 | 3.3497
I1 | A | 8 | 3.2649 | 3.3901 | 3.5077 | 3.5758 | 3.6135 | 3.6029 | 3.6009 | 3.607 | 3.5061 | 3.3864 | 3.3324
I1 | A | 9 | 3.2202 | 3.351 | 3.4726 | 3.5501 | 3.5811 | 3.5782 | 3.5744 | 3.5832 | 3.4899 | 3.3661 | 3.3184
I1 | A | 10 | 3.1862 | 3.3167 | 3.4379 | 3.5166 | 3.5546 | 3.553 | 3.5538 | 3.5643 | 3.4704 | 3.3521 | 3.3009
I1 | A | 11 | 3.146 | 3.279 | 3.4071 | 3.4864 | 3.5281 | 3.533 | 3.5318 | 3.5488 | 3.4527 | 3.3351 | 3.2806
I1 | A | 12 | 3.1145 | 3.2459 | 3.3804 | 3.4608 | 3.5087 | 3.5154 | 3.5179 | 3.5344 | 3.438 | 3.3264 | 3.284
I1 | A | 13 | 3.0855 | 3.2143 | 3.3539 | 3.4413 | 3.4895 | 3.5009 | 3.5004 | 3.5249 | 3.4283 | 3.3136 | 3.2674
I1 | A | 14 | 3.0499 | 3.1846 | 3.3242 | 3.4119 | 3.4659 | 3.4848 | 3.4882 | 3.5138 | 3.4164 | 3.3031 | 3.255
I1 | A | 15 | 3.026 | 3.156 | 3.3036 | 3.3888 | 3.4478 | 3.4739 | 3.4763 | 3.5029 | 3.4119 | 3.2942 | 3.2494
I1 | A | 16 | 2.993 | 3.1306 | 3.2803 | 3.3774 | 3.4424 | 3.4617 | 3.4684 | 3.4993 | 3.4043 | 3.2938 | 3.2499
I1 | A | 17 | 2.9652 | 3.1064 | 3.2594 | 3.3641 | 3.4347 | 3.4586 | 3.465 | 3.4971 | 3.4036 | 3.2861 | 3.2452
I1 | A | 18 | 2.94 | 3.0833 | 3.2412 | 3.3514 | 3.4276 | 3.45 | 3.4621 | 3.4978 | 3.4023 | 3.2856 | 3.2427
I1 | A | 19 | 2.9195 | 3.0646 | 3.228 | 3.3451 | 3.4243 | 3.4535 | 3.4633 | 3.5035 | 3.4039 | 3.2905 | 3.2398
I1 | A | 20 | 2.8991 | 3.047 | 3.219 | 3.3352 | 3.4254 | 3.455 | 3.4681 | 3.507 | 3.4101 | 3.291 | 3.2504
I1 | A | 21 | 2.8803 | 3.0345 | 3.2082 | 3.335 | 3.4313 | 3.4654 | 3.478 | 3.519 | 3.4192 | 3.3007 | 3.2595
I1 | A | 22 | 2.8653 | 3.0207 | 3.205 | 3.3329 | 3.4454 | 3.4747 | 3.4877 | 3.5307 | 3.4275 | 3.3077 | 3.2629
I1 | A | 23 | 2.8486 | 3.0108 | 3.2023 | 3.3381 | 3.4615 | 3.4881 | 3.5054 | 3.5504 | 3.4425 | 3.3151 | 3.2734
I1 | A | 24 | 2.8323 | 3.0075 | 3.2052 | 3.3558 | 3.4788 | 3.5102 | 3.5246 | 3.5663 | 3.4614 | 3.3292 | 3.2858
I1 | A | 25 | 2.8218 | 2.9972 | 3.2043 | 3.3668 | 3.5017 | 3.5333 | 3.5461 | 3.591 | 3.4792 | 3.3366 | 3.2984
I1 | A | 26 | 2.8064 | 2.998 | 3.2169 | 3.3946 | 3.5265 | 3.5582 | 3.5738 | 3.6184 | 3.4996 | 3.3521 | 3.3158
I1 | A | 27 | 2.797 | 2.9952 | 3.2266 | 3.4075 | 3.5559 | 3.5857 | 3.5986 | 3.6451 | 3.516 | 3.3712 | 3.3267
I1 | A | 28 | 2.7791 | 2.9932 | 3.2443 | 3.443 | 3.5957 | 3.6256 | 3.6371 | 3.6811 | 3.5405 | 3.3893 | 3.3416
I1 | A | 29 | 2.7704 | 3.0014 | 3.26 | 3.4789 | 3.641 | 3.6632 | 3.6729 | 3.7156 | 3.5681 | 3.4065 | 3.3608
I1 | A | 30 | 2.7652 | 3.0021 | 3.2791 | 3.5114 | 3.683 | 3.7084 | 3.7132 | 3.7528 | 3.596 | 3.4273 | 3.3812
I1 | A | 31 | 2.7648 | 3.0151 | 3.3103 | 3.5823 | 3.7416 | 3.7581 | 3.7623 | 3.8012 | 3.6311 | 3.4511 | 3.3964
I1 | A | 32 | 2.7508 | 3.0213 | 3.3436 | 3.6334 | 3.8002 | 3.808 | 3.8109 | 3.844 | 3.6647 | 3.4735 | 3.4273
I1 | B | 1 | 4.4434 | 4.3565 | 4.2611 | 4.2224 | 4.1172 | 4.0296 | 3.9682 | 3.8697 | 3.6945 | 3.4934 | 3.4089
I1 | B | 2 | 4.4111 | 4.3323 | 4.2326 | 4.1957 | 4.0858 | 3.9955 | 3.9369 | 3.8383 | 3.6616 | 3.4653 | 3.3732
I1 | B | 3 | 4.3881 | 4.3015 | 4.2053 | 4.1643 | 4.0621 | 3.9711 | 3.9092 | 3.813 | 3.6306 | 3.4339 | 3.3491
I1 | B | 4 | 4.3528 | 4.2741 | 4.1776 | 4.1399 | 4.0321 | 3.9433 | 3.882 | 3.7853 | 3.6082 | 3.4114 | 3.3229
I1 | B | 5 | 4.33 | 4.2494 | 4.1535 | 4.1172 | 4.01 | 3.9212 | 3.8586 | 3.7642 | 3.5867 | 3.3907 | 3.2991
I1 | B | 6 | 4.3028 | 4.2263 | 4.1282 | 4.0896 | 3.9827 | 3.8999 | 3.8357 | 3.7432 | 3.5654 | 3.3675 | 3.2809
I1 | B | 7 | 4.2713 | 4.1938 | 4.1026 | 4.0698 | 3.9599 | 3.8747 | 3.8131 | 3.7206 | 3.5474 | 3.351 | 3.2658
I1 | B | 8 | 4.2405 | 4.1714 | 4.083 | 4.0406 | 3.9398 | 3.8589 | 3.7946 | 3.7062 | 3.5295 | 3.334 | 3.252
I1 | B | 9 | 4.2147 | 4.1477 | 4.0601 | 4.0272 | 3.9207 | 3.8398 | 3.7793 | 3.6895 | 3.5141 | 3.3182 | 3.2438
I1 | B | 10 | 4.1714 | 4.1146 | 4.0324 | 3.9975 | 3.8945 | 3.8186 | 3.7565 | 3.6738 | 3.4975 | 3.3083 | 3.2286
I1 | B | 11 | 4.1513 | 4.0989 | 4.0189 | 3.989 | 3.8849 | 3.8093 | 3.7467 | 3.6605 | 3.485 | 3.2949 | 3.2142
I1 | B | 12 | 4.1406 | 4.0895 | 4.0095 | 3.9792 | 3.8704 | 3.7988 | 3.7351 | 3.6505 | 3.4735 | 3.2849 | 3.2154
I1 | B | 13 | 4.1184 | 4.0756 | 3.9986 | 3.969 | 3.8649 | 3.7896 | 3.7273 | 3.646 | 3.468 | 3.2767 | 3.2014
I1 | B | 14 | 4.0946 | 4.0575 | 3.9859 | 3.9568 | 3.8552 | 3.7799 | 3.7174 | 3.6347 | 3.4599 | 3.2681 | 3.1927
I1 | B | 15 | 4.0804 | 4.0504 | 3.9768 | 3.9474 | 3.848 | 3.7767 | 3.7122 | 3.6305 | 3.4553 | 3.2607 | 3.1888
I1 | B | 16 | 4.0673 | 4.0394 | 3.9721 | 3.9498 | 3.843 | 3.7756 | 3.7074 | 3.6278 | 3.4494 | 3.2598 | 3.1867
I1 | B | 17 | 4.0802 | 4.0533 | 3.9806 | 3.9606 | 3.8519 | 3.7782 | 3.7133 | 3.6312 | 3.4494 | 3.2563 | 3.1829
I1 | B | 18 | 4.0826 | 4.0591 | 3.9861 | 3.9611 | 3.8522 | 3.7835 | 3.7141 | 3.6323 | 3.4481 | 3.2527 | 3.1809
I1 | B | 19 | 4.1091 | 4.0792 | 3.9984 | 3.9762 | 3.864 | 3.7891 | 3.722 | 3.6405 | 3.4513 | 3.2547 | 3.1796
I1 | B | 20 | 4.1224 | 4.0946 | 4.0131 | 3.9878 | 3.8721 | 3.7978 | 3.7309 | 3.6457 | 3.4561 | 3.2571 | 3.1841
I1 | B | 21 | 4.1685 | 4.1282 | 4.0355 | 4.0062 | 3.8874 | 3.8136 | 3.7427 | 3.6555 | 3.4657 | 3.2625 | 3.1914
I1 | B | 22 | 4.2093 | 4.1614 | 4.0567 | 4.0269 | 3.912 | 3.8317 | 3.7579 | 3.6712 | 3.4729 | 3.2675 | 3.1968
I1 | B | 23 | 4.2878 | 4.2113 | 4.0942 | 4.0565 | 3.9421 | 3.8562 | 3.7803 | 3.688 | 3.4845 | 3.2732 | 3.2013
I1 | B | 24 | 4.3562 | 4.2613 | 4.1313 | 4.099 | 3.9669 | 3.8782 | 3.8002 | 3.7051 | 3.4986 | 3.283 | 3.2068
I1 | B | 25 | 4.4556 | 4.3198 | 4.1744 | 4.1345 | 4.0015 | 3.9067 | 3.8252 | 3.7275 | 3.5161 | 3.2882 | 3.2152
I1 | B | 26 | 4.547 | 4.3832 | 4.2189 | 4.1776 | 4.0354 | 3.9365 | 3.8524 | 3.7497 | 3.5315 | 3.3 | 3.2253
I1 | B | 27 | 4.6231 | 4.4403 | 4.2549 | 4.2054 | 4.0628 | 3.9618 | 3.874 | 3.7701 | 3.5419 | 3.3084 | 3.2293
I1 | B | 28 | 4.7505 | 4.5201 | 4.326 | 4.2719 | 4.1097 | 4.005 | 3.9124 | 3.8015 | 3.5633 | 3.3232 | 3.2405
I1 | B | 29 | 4.8373 | 4.5807 | 4.3721 | 4.31 | 4.1491 | 4.0396 | 3.944 | 3.8296 | 3.586 | 3.336 | 3.2495
I1 | B | 30 | 4.9191 | 4.6492 | 4.4305 | 4.3611 | 4.1882 | 4.0784 | 3.9783 | 3.8565 | 3.6038 | 3.3459 | 3.2645
I1 | B | 31 | 5.0022 | 4.7156 | 4.4853 | 4.4156 | 4.2346 | 4.1199 | 4.0164 | 3.8897 | 3.6309 | 3.3604 | 3.2667
I1 | B | 32 | 5.066 | 4.7696 | 4.5404 | 4.4633 | 4.2785 | 4.1612 | 4.052 | 3.9211 | 3.6533 | 3.3765 | 3.2874
I2 | A | 1 | 2.9503 | 3.1703 | 3.4336 | 3.6819 | 3.9673 | 4.0662 | 4.2319 | 4.4598 | 4.5849 | 4.6428 | 4.6264
I2 | A | 2 | 2.9375 | 3.1234 | 3.3595 | 3.5752 | 3.8552 | 3.9661 | 4.1421 | 4.4026 | 4.5483 | 4.6152 | 4.6027
I2 | A | 3 | 2.9408 | 3.106 | 3.3228 | 3.5125 | 3.7707 | 3.878 | 4.0546 | 4.3245 | 4.4951 | 4.5764 | 4.5683
I2 | A | 4 | 2.9384 | 3.0811 | 3.2721 | 3.4409 | 3.6762 | 3.7909 | 3.9644 | 4.2554 | 4.4436 | 4.5441 | 4.5386
I2 | A | 5 | 2.9355 | 3.0618 | 3.2401 | 3.39 | 3.609 | 3.723 | 3.8946 | 4.2078 | 4.4123 | 4.5284 | 4.5255
I2 | A | 6 | 2.9455 | 3.057 | 3.2157 | 3.352 | 3.5485 | 3.6581 | 3.8258 | 4.1441 | 4.3676 | 4.5005 | 4.5036
I2 | A | 7 | 2.9466 | 3.0448 | 3.1951 | 3.3151 | 3.499 | 3.6125 | 3.7713 | 4.1057 | 4.343 | 4.4924 | 4.4967
I2 | A | 8 | 2.9497 | 3.0375 | 3.1719 | 3.2867 | 3.4443 | 3.5616 | 3.7188 | 4.0691 | 4.3208 | 4.4841 | 4.4958
I2 | A | 9 | 2.969 | 3.0416 | 3.1662 | 3.2673 | 3.4074 | 3.5156 | 3.6608 | 4.007 | 4.2739 | 4.4552 | 4.4678
I2 | A | 10 | 2.9728 | 3.0396 | 3.1503 | 3.2441 | 3.3752 | 3.4838 | 3.627 | 3.9775 | 4.2673 | 4.4635 | 4.479
I2 | A | 11 | 2.9797 | 3.0366 | 3.1407 | 3.2209 | 3.3452 | 3.4486 | 3.584 | 3.9425 | 4.2524 | 4.4644 | 4.4842
I2 | A | 12 | 2.9902 | 3.0394 | 3.1398 | 3.2123 | 3.3207 | 3.4246 | 3.5564 | 3.9225 | 4.2433 | 4.4694 | 4.4886
I2 | A | 13 | 2.9971 | 3.0394 | 3.1334 | 3.2003 | 3.3019 | 3.3978 | 3.527 | 3.9003 | 4.2465 | 4.4757 | 4.5024
I2 | A | 14 | 3.0104 | 3.0463 | 3.1295 | 3.1874 | 3.2842 | 3.3797 | 3.5021 | 3.882 | 4.2362 | 4.4897 | 4.5151
I2 | A | 15 | 3.0171 | 3.0439 | 3.1205 | 3.1727 | 3.2587 | 3.3545 | 3.4695 | 3.8631 | 4.2492 | 4.5163 | 4.539
I2 | A | 16 | 3.0249 | 3.0453 | 3.1163 | 3.1664 | 3.2415 | 3.331 | 3.4465 | 3.8515 | 4.2666 | 4.5423 | 4.5666
I2 | A | 17 | 3.0302 | 3.0446 | 3.1052 | 3.1464 | 3.2144 | 3.3104 | 3.422 | 3.8483 | 4.3051 | 4.5901 | 4.613
I2 | A | 18 | 3.0442 | 3.0505 | 3.1059 | 3.1432 | 3.2021 | 3.2924 | 3.4009 | 3.8442 | 4.3246 | 4.6262 | 4.6435
I2 | A | 19 | 3.0467 | 3.0458 | 3.0956 | 3.1217 | 3.1756 | 3.272 | 3.3778 | 3.8597 | 4.4031 | 4.7035 | 4.7137
I2 | A | 20 | 3.0616 | 3.0555 | 3.1 | 3.1273 | 3.1743 | 3.263 | 3.3693 | 3.8634 | 4.432 | 4.7352 | 4.7449
I2 | A | 21 | 3.0704 | 3.059 | 3.0933 | 3.1149 | 3.1519 | 3.2436 | 3.3471 | 3.876 | 4.5258 | 4.8145 | 4.8095
I2 | A | 22 | 3.0692 | 3.0476 | 3.0758 | 3.0863 | 3.1141 | 3.1997 | 3.3092 | 3.9537 | 4.7035 | 4.9403 | 4.9117
I2 | A | 23 | 3.0781 | 3.0496 | 3.0745 | 3.0776 | 3.0961 | 3.1851 | 3.2866 | 3.9911 | 4.8167 | 5.0088 | 4.9748
I2 | A | 24 | 3.0885 | 3.0509 | 3.0686 | 3.0655 | 3.0694 | 3.1558 | 3.264 | 4.0944 | 4.9782 | 5.1092 | 5.062
I2 | A | 25 | 3.0955 | 3.0536 | 3.0632 | 3.0524 | 3.0483 | 3.1299 | 3.2274 | 4.2585 | 5.1404 | 5.2118 | 5.1403
I2 | A | 26 | 3.1005 | 3.0492 | 3.0485 | 3.028 | 2.9921 | 3.0871 | 3.1668 | 4.562 | 5.3472 | 5.3221 | 5.2337
I2 | A | 27 | 3.1016 | 3.0396 | 3.0279 | 2.9982 | 2.9297 | 2.9971 | 3.0365 | 5.1379 | 5.5439 | 5.4428 | 5.3583
I2 | A | 28 | 3.1126 | 3.0422 | 3.0189 | 2.9706 | 2.8695 | 2.9443 | 2.8872 | 5.537 | 5.6887 | 5.537 | 5.4317
I2 | A | 29 | 3.1192 | 3.0376 | 3.0076 | 2.9444 | 2.7621 | 2.8449 | 2.3662 | 5.8143 | 5.803 | 5.6165 | 5.5136
I2 | A | 30 | 3.1302 | 3.0341 | 2.9874 | 2.9075 | 2.641 | 2.664 | 1.3547 | 6.0074 | 5.9059 | 5.6957 | 5.5839
I2 | A | 31 | 3.1426 | 3.0405 | 2.9811 | 2.8677 | 2.4852 | 2.4115 | 0.83964 | 6.0942 | 5.9562 | 5.7498 | 5.6432
I2 | A | 32 | 3.1537 | 3.0368 | 2.9602 | 2.8154 | 2.2421 | 2.0121 | 0.59045 | 6.1799 | 6.0177 | 5.8092 | 5.7071
I2 | B | 1 | 3.0873 | 3.1926 | 3.3376 | 3.4827 | 3.7011 | 3.7958 | 3.9476 | 4.1796 | 4.3348 | 4.4445 | 4.4676
I2 | B | 2 | 3.058 | 3.1496 | 3.2858 | 3.4218 | 3.6266 | 3.7267 | 3.8786 | 4.1275 | 4.2933 | 4.4129 | 4.4341
I2 | B | 3 | 3.0438 | 3.1273 | 3.2535 | 3.3773 | 3.5749 | 3.6708 | 3.8187 | 4.0702 | 4.2437 | 4.3697 | 4.3938
I2 | B | 4 | 3.0224 | 3.0952 | 3.2109 | 3.3267 | 3.5101 | 3.6114 | 3.7571 | 4.0169 | 4.1973 | 4.3308 | 4.3611
I2 | B | 5 | 3.0027 | 3.069 | 3.1798 | 3.2836 | 3.4594 | 3.5606 | 3.7081 | 3.9779 | 4.1718 | 4.3086 | 4.3415
I2 | B | 6 | 2.9918 | 3.0481 | 3.1496 | 3.2508 | 3.4116 | 3.5171 | 3.661 | 3.9343 | 4.1355 | 4.2775 | 4.313
I2 | B | 7 | 2.9735 | 3.0259 | 3.123 | 3.2148 | 3.3751 | 3.4789 | 3.6206 | 3.9093 | 4.1148 | 4.2697 | 4.2991
I2 | B | 8 | 2.9601 | 3.005 | 3.0962 | 3.1827 | 3.3351 | 3.4403 | 3.583 | 3.8839 | 4.0977 | 4.2564 | 4.2921
I2 | B | 9 | 2.9544 | 2.9913 | 3.0783 | 3.1588 | 3.2989 | 3.4058 | 3.5431 | 3.844 | 4.0629 | 4.2285 | 4.2639
I2 | B | 10 | 2.9407 | 2.9723 | 3.0544 | 3.1298 | 3.2674 | 3.3756 | 3.5147 | 3.8268 | 4.0636 | 4.2317 | 4.2683
I2 | B | 11 | 2.9295 | 2.9562 | 3.0336 | 3.1009 | 3.2322 | 3.3457 | 3.4838 | 3.8121 | 4.0581 | 4.2326 | 4.2677
I2 | B | 12 | 2.9201 | 2.9422 | 3.017 | 3.0846 | 3.2112 | 3.32 | 3.4604 | 3.7948 | 4.0525 | 4.2327 | 4.2696
I2 | B | 13 | 2.9094 | 2.9279 | 2.9972 | 3.0597 | 3.1788 | 3.2918 | 3.4348 | 3.7858 | 4.0636 | 4.2433 | 4.2809
I2 | B | 14 | 2.9021 | 2.9163 | 2.9828 | 3.0402 | 3.1568 | 3.2747 | 3.4154 | 3.7772 | 4.0615 | 4.2498 | 4.2875
I2 | B | 15 | 2.892 | 2.8991 | 2.9606 | 3.0114 | 3.1224 | 3.2428 | 3.3865 | 3.7756 | 4.0769 | 4.273 | 4.3041
I2 | B | 16 | 2.8828 | 2.885 | 2.9416 | 2.9874 | 3.0952 | 3.2124 | 3.3636 | 3.7705 | 4.0995 | 4.2935 | 4.3263
I2 | B | 17 | 2.869 | 2.8651 | 2.9168 | 2.9572 | 3.0558 | 3.1807 | 3.3362 | 3.7784 | 4.1316 | 4.3329 | 4.3625
I2 | B | 18 | 2.8613 | 2.8537 | 2.9002 | 2.9354 | 3.028 | 3.1574 | 3.3136 | 3.7874 | 4.1676 | 4.3614 | 4.3874
I2 | B | 19 | 2.8479 | 2.8327 | 2.8751 | 2.9022 | 2.9907 | 3.1215 | 3.2836 | 3.8135 | 4.2257 | 4.4174 | 4.439
I2 | B | 20 | 2.8433 | 2.8242 | 2.8641 | 2.8888 | 2.9665 | 3.1073 | 3.2733 | 3.8282 | 4.2677 | 4.4562 | 4.4707
I2 | B | 21 | 2.8317 | 2.8051 | 2.838 | 2.8537 | 2.928 | 3.0698 | 3.2406 | 3.864 | 4.3421 | 4.5162 | 4.5247
I2 | B | 22 | 2.8098 | 2.7754 | 2.7993 | 2.8045 | 2.8566 | 2.9986 | 3.1829 | 3.9392 | 4.4791 | 4.6288 | 4.6152
I2 | B | 23 | 2.7999 | 2.7604 | 2.778 | 2.7737 | 2.8127 | 2.9588 | 3.145 | 4.0147 | 4.5852 | 4.6924 | 4.6774
I2 | B | 24 | 2.7865 | 2.7377 | 2.7487 | 2.7321 | 2.7472 | 2.8989 | 3.0852 | 4.1452 | 4.7239 | 4.7807 | 4.7519
I2 | B | 25 | 2.7725 | 2.7172 | 2.718 | 2.6913 | 2.6793 | 2.8251 | 2.9999 | 4.3394 | 4.8787 | 4.8783 | 4.8303
I2 | B | 26 | 2.754 | 2.6876 | 2.6774 | 2.6307 | 2.5723 | 2.7209 | 2.8392 | 4.6888 | 5.0654 | 4.9971 | 4.9288
I2 | B | 27 | 2.7327 | 2.6526 | 2.6269 | 2.5521 | 2.4269 | 2.542 | 2.4558 | 5.2 | 5.2748 | 5.1302 | 5.0528
I2 | B | 28 | 2.7157 | 2.6218 | 2.581 | 2.4873 | 2.2854 | 2.3638 | 1.9059 | 5.5338 | 5.4148 | 5.2385 | 5.1458
I2 | B | 29 | 2.6975 | 2.5896 | 2.534 | 2.4084 | 2.1053 | 2.1028 | 1.2401 | 5.8048 | 5.5558 | 5.3415 | 5.2364
I2 | B | 30 | 2.6757 | 2.5542 | 2.4772 | 2.3186 | 1.9067 | 1.7898 | 0.85539 | 5.961 | 5.6744 | 5.4359 | 5.3295
I2 | B | 31 | 2.658 | 2.522 | 2.4286 | 2.2288 | 1.7293 | 1.5221 | 0.68095 | 6.0431 | 5.7527 | 5.5041 | 5.3935
I2 | B | 32 | 2.6343 | 2.4797 | 2.3623 | 2.1179 | 1.515 | 1.2455 | 0.55741 | 6.1032 | 5.8246 | 5.58 | 5.4749
M1 | A | 1 | 4.3682 | 4.2628 | 4.1926 | 4.1498 | 4.0827 | 4.0265 | 4.0129 | 3.9491 | 3.8711 | 3.7871 | 3.753
M1 | A | 2 | 4.3088 | 4.2004 | 4.1281 | 4.0866 | 4.0158 | 3.9593 | 3.9447 | 3.8807 | 3.7994 | 3.7111 | 3.6745
M1 | A | 3 | 4.253 | 4.1382 | 4.0646 | 4.0242 | 3.9483 | 3.894 | 3.8783 | 3.8146 | 3.7288 | 3.6383 | 3.6009
M1 | A | 4 | 4.1911 | 4.0766 | 4.0048 | 3.9568 | 3.8841 | 3.826 | 3.8122 | 3.7467 | 3.6612 | 3.565 | 3.524
M1 | A | 5 | 4.1279 | 4.0091 | 3.9359 | 3.8926 | 3.8158 | 3.7581 | 3.7446 | 3.6809 | 3.5903 | 3.4934 | 3.4537
M1 | A | 6 | 4.0646 | 3.9438 | 3.867 | 3.8287 | 3.7475 | 3.6928 | 3.6776 | 3.6149 | 3.522 | 3.4225 | 3.3822
M1 | A | 7 | 3.9965 | 3.8759 | 3.802 | 3.7607 | 3.6797 | 3.625 | 3.6106 | 3.5499 | 3.4546 | 3.3557 | 3.313
M1 | A | 8 | 3.9285 | 3.8061 | 3.7315 | 3.6951 | 3.6102 | 3.5595 | 3.5445 | 3.4845 | 3.3894 | 3.2884 | 3.2451
M1 | A | 9 | 3.8576 | 3.7367 | 3.6617 | 3.6251 | 3.5423 | 3.493 | 3.4785 | 3.4206 | 3.3259 | 3.2237 | 3.1811
M1 | A | 10 | 3.7825 | 3.6598 | 3.5918 | 3.5577 | 3.4772 | 3.4271 | 3.4133 | 3.3558 | 3.2623 | 3.1583 | 3.1168
M1 | A | 11 | 3.7037 | 3.5827 | 3.5186 | 3.4877 | 3.4078 | 3.3596 | 3.3469 | 3.294 | 3.1981 | 3.0976 | 3.0567
M1 | A | 12 | 3.6184 | 3.5047 | 3.4408 | 3.4199 | 3.3394 | 3.2935 | 3.2828 | 3.2327 | 3.1392 | 3.0366 | 2.9938
M1 | A | 13 | 3.5208 | 3.4208 | 3.3626 | 3.3464 | 3.2718 | 3.225 | 3.2165 | 3.1689 | 3.0779 | 2.975 | 2.9354
M1 | A | 14 | 3.4153 | 3.3295 | 3.2804 | 3.2698 | 3.1976 | 3.1547 | 3.1488 | 3.1035 | 3.0164 | 2.9174 | 2.8779
M1 | A | 15 | 3.2989 | 3.2329 | 3.1903 | 3.1843 | 3.1227 | 3.0845 | 3.0794 | 3.0385 | 2.9572 | 2.8599 | 2.821
M1 | A | 16 | 3.161 | 3.124 | 3.095 | 3.1017 | 3.0447 | 3.0125 | 3.0087 | 2.9726 | 2.8986 | 2.8026 | 2.7662
M1 | B | 1 | 4.4977 | 4.3683 | 4.2759 | 4.2132 | 4.1245 | 4.0606 | 4.0254 | 3.9317 | 3.8103 | 3.6701 | 3.6137
M1 | B | 2 | 4.4495 | 4.3161 | 4.2219 | 4.1599 | 4.0646 | 4.0001 | 3.9656 | 3.8702 | 3.746 | 3.6034 | 3.5435
M1 | B | 3 | 4.4043 | 4.2635 | 4.1694 | 4.1049 | 4.01 | 3.9428 | 3.9077 | 3.8152 | 3.6856 | 3.5374 | 3.4753
M1 | B | 4 | 4.357 | 4.2141 | 4.1165 | 4.0499 | 3.9511 | 3.8854 | 3.8515 | 3.7573 | 3.6263 | 3.4745 | 3.4104
M1 | B | 5 | 4.3089 | 4.1593 | 4.0594 | 3.9933 | 3.8938 | 3.8282 | 3.7932 | 3.6996 | 3.5676 | 3.4127 | 3.3481
M1 | B | 6 | 4.2613 | 4.1027 | 4.003 | 3.9392 | 3.8358 | 3.769 | 3.7362 | 3.6424 | 3.509 | 3.3536 | 3.2877
M1 | B | 7 | 4.213 | 4.0494 | 3.9487 | 3.8842 | 3.7783 | 3.7132 | 3.6792 | 3.5869 | 3.4525 | 3.2956 | 3.2292
M1 | B | 8 | 4.1649 | 3.9955 | 3.8931 | 3.8295 | 3.7197 | 3.6565 | 3.6227 | 3.5328 | 3.3981 | 3.2388 | 3.1716
M1 | B | 9 | 4.1153 | 3.9416 | 3.8384 | 3.774 | 3.6641 | 3.603 | 3.5681 | 3.4783 | 3.3433 | 3.1838 | 3.1168
M1 | B | 10 | 4.0684 | 3.8861 | 3.7809 | 3.7183 | 3.6085 | 3.5478 | 3.5141 | 3.4257 | 3.2892 | 3.1316 | 3.0656
M1 | B | 11 | 4.0155 | 3.8308 | 3.7226 | 3.6605 | 3.5527 | 3.492 | 3.4594 | 3.3725 | 3.2382 | 3.0817 | 3.0147
M1 | B | 12 | 3.9682 | 3.7725 | 3.6658 | 3.6095 | 3.4947 | 3.4345 | 3.4045 | 3.322 | 3.189 | 3.0317 | 2.9643
M1 | B | 13 | 3.913 | 3.7118 | 3.6045 | 3.5495 | 3.4396 | 3.3804 | 3.3506 | 3.2692 | 3.1394 | 2.982 | 2.9168
M1 | B | 14 | 3.8509 | 3.6474 | 3.5388 | 3.4908 | 3.3788 | 3.3204 | 3.2946 | 3.2166 | 3.0889 | 2.9361 | 2.8702
M1 | B | 15 | 3.7806 | 3.576 | 3.4724 | 3.4233 | 3.3191 | 3.2643 | 3.2378 | 3.1635 | 3.0406 | 2.8886 | 2.8248
M1 | B | 16 | 3.7021 | 3.4943 | 3.3941 | 3.3536 | 3.2532 | 3.1984 | 3.178 | 3.1074 | 2.9901 | 2.843 | 2.7807
M2 | A | 1 | 4.4824 | 4.3845 | 4.3216 | 4.28 | 4.2155 | 4.1667 | 4.1548 | 4.1017 | 4.0348 | 3.9683 | 3.9372
M2 | A | 2 | 4.372 | 4.2706 | 4.2051 | 4.1627 | 4.0934 | 4.0423 | 4.0305 | 3.9761 | 3.9042 | 3.8309 | 3.7958
M2 | A | 3 | 4.2657 | 4.1605 | 4.0908 | 4.0524 | 3.9773 | 3.9252 | 3.9134 | 3.8586 | 3.7819 | 3.7032 | 3.6657
M2 | A | 4 | 4.1577 | 4.0479 | 3.9798 | 3.9422 | 3.862 | 3.8113 | 3.7991 | 3.7436 | 3.6651 | 3.5834 | 3.5456
M2 | A | 5 | 4.0412 | 3.932 | 3.8648 | 3.832 | 3.7481 | 3.6978 | 3.6877 | 3.6325 | 3.5529 | 3.4696 | 3.4322
M2 | A | 6 | 3.912 | 3.8042 | 3.7424 | 3.7125 | 3.6315 | 3.5819 | 3.5729 | 3.5186 | 3.4416 | 3.3569 | 3.3195
M2 | A | 7 | 3.7738 | 3.6762 | 3.6198 | 3.5957 | 3.5154 | 3.4685 | 3.4616 | 3.4129 | 3.3345 | 3.2519 | 3.2163
M2 | A | 8 | 3.6226 | 3.5379 | 3.4887 | 3.4737 | 3.3977 | 3.3559 | 3.3499 | 3.3046 | 3.2315 | 3.1509 | 3.1165
M2 | A | 9 | 3.4608 | 3.3969 | 3.3582 | 3.3504 | 3.2831 | 3.2437 | 3.2409 | 3.2002 | 3.1299 | 3.0521 | 3.0202
M2 | A | 10 | 3.2883 | 3.2493 | 3.2231 | 3.2265 | 3.1692 | 3.1355 | 3.1339 | 3.0993 | 3.035 | 2.9607 | 2.9311
M2 | A | 11 | 3.1126 | 3.1046 | 3.0902 | 3.1042 | 3.0571 | 3.0279 | 3.0297 | 3.0007 | 2.941 | 2.8724 | 2.8437
M2 | A | 12 | 2.9304 | 2.9547 | 2.955 | 2.9805 | 2.9447 | 2.9242 | 2.9261 | 2.9031 | 2.8507 | 2.7853 | 2.7605
M2 | A | 13 | 2.7483 | 2.8052 | 2.8204 | 2.8557 | 2.8359 | 2.8192 | 2.8244 | 2.8072 | 2.7627 | 2.7011 | 2.6781
M2 | A | 14 | 2.5695 | 2.6571 | 2.685 | 2.7297 | 2.725 | 2.7115 | 2.7218 | 2.7087 | 2.6747 | 2.6198 | 2.5972
M2 | A | 15 | 2.4093 | 2.5134 | 2.551 | 2.6022 | 2.613 | 2.6056 | 2.6185 | 2.6127 | 2.5876 | 2.5389 | 2.5192
M2 | A | 16 | 2.2608 | 2.3712 | 2.4137 | 2.4693 | 2.4932 | 2.4898 | 2.5067 | 2.5083 | 2.4961 | 2.4551 | 2.4368
M2 | B | 1 | 4.5476 | 4.4329 | 4.3534 | 4.301 | 4.2232 | 4.1671 | 4.1433 | 4.0719 | 3.9785 | 3.8789 | 3.8364
M2 | B | 2 | 4.4591 | 4.3347 | 4.2518 | 4.1988 | 4.1124 | 4.0551 | 4.0322 | 3.9583 | 3.8595 | 3.752 | 3.7028
M2 | B | 3 | 4.3709 | 4.238 | 4.15 | 4.097 | 4.0062 | 3.9472 | 3.9243 | 3.8514 | 3.748 | 3.6352 | 3.5842
M2 | B | 4 | 4.2772 | 4.1374 | 4.0488 | 3.9953 | 3.9004 | 3.8417 | 3.8188 | 3.7444 | 3.642 | 3.5263 | 3.4739
M2 | B | 5 | 4.181 | 4.0354 | 3.9429 | 3.8937 | 3.7941 | 3.7371 | 3.7154 | 3.6434 | 3.5396 | 3.4219 | 3.3702
M2 | B | 6 | 4.0746 | 3.9247 | 3.8343 | 3.7847 | 3.6879 | 3.6319 | 3.6107 | 3.5405 | 3.4378 | 3.322 | 3.2707
M2 | B | 7 | 3.9642 | 3.8127 | 3.7252 | 3.6809 | 3.5833 | 3.5305 | 3.5094 | 3.4435 | 3.3414 | 3.2262 | 3.1772
M2 | B | 8 | 3.8411 | 3.693 | 3.6105 | 3.5739 | 3.4761 | 3.4259 | 3.4076 | 3.3453 | 3.2474 | 3.1355 | 3.0862
M2 | B | 9 | 3.7093 | 3.5698 | 3.4936 | 3.4612 | 3.3728 | 3.3256 | 3.3081 | 3.2514 | 3.1569 | 3.0482 | 3.0022
M2 | B | 10 | 3.5691 | 3.4441 | 3.3767 | 3.3523 | 3.2705 | 3.2247 | 3.212 | 3.1584 | 3.0693 | 2.9631 | 2.9196
M2 | B | 11 | 3.4185 | 3.3174 | 3.258 | 3.2433 | 3.1686 | 3.128 | 3.1176 | 3.0688 | 2.9849 | 2.8845 | 2.8426
M2 | B | 12 | 3.2495 | 3.1786 | 3.1353 | 3.1291 | 3.0666 | 3.0311 | 3.023 | 2.9789 | 2.9035 | 2.807 | 2.767
M2 | B | 13 | 3.0741 | 3.0404 | 3.0102 | 3.0179 | 2.9663 | 2.9371 | 2.9305 | 2.8918 | 2.8232 | 2.7328 | 2.6947
M2 | B | 14 | 2.8787 | 2.8949 | 2.8817 | 2.901 | 2.8627 | 2.8377 | 2.8358 | 2.8041 | 2.7416 | 2.6559 | 2.6212
M2 | B | 15 | 2.6694 | 2.7348 | 2.7428 | 2.776 | 2.7537 | 2.7329 | 2.7368 | 2.7117 | 2.6624 | 2.5835 | 2.5501
M2 | B | 16 | 2.4337 | 2.5512 | 2.583 | 2.6269 | 2.6281 | 2.6167 | 2.6246 | 2.6098 | 2.5766 | 2.5079 | 2.4771
M3 | A | 1 | 4.4831 | 4.3939 | 4.3322 | 4.2873 | 4.2226 | 4.1685 | 4.1568 | 4.1005 | 4.0346 | 3.9652 | 3.9199
M3 | A | 2 | 4.3301 | 4.2321 | 4.1644 | 4.1178 | 4.0409 | 3.9806 | 3.9666 | 3.9038 | 3.8295 | 3.7471 | 3.6882
M3 | A | 3 | 4.2148 | 4.1122 | 4.041 | 3.9942 | 3.9145 | 3.8525 | 3.838 | 3.7735 | 3.6959 | 3.6087 | 3.5512
M3 | A | 4 | 4.1055 | 4.005 | 3.9316 | 3.8896 | 3.8082 | 3.7466 | 3.7339 | 3.6704 | 3.5924 | 3.5034 | 3.4499
M3 | A | 5 | 3.9943 | 3.8938 | 3.8262 | 3.787 | 3.7069 | 3.6483 | 3.6362 | 3.5743 | 3.4969 | 3.4081 | 3.359
M3 | A | 6 | 3.8683 | 3.7775 | 3.7132 | 3.684 | 3.6043 | 3.5489 | 3.5388 | 3.4797 | 3.4043 | 3.3158 | 3.2712
M3 | A | 7 | 3.7283 | 3.6502 | 3.5936 | 3.573 | 3.4987 | 3.4454 | 3.438 | 3.3833 | 3.3107 | 3.2239 | 3.1832
M3 | A | 8 | 3.5724 | 3.516 | 3.4705 | 3.4615 | 3.3945 | 3.3467 | 3.3421 | 3.292 | 3.2234 | 3.1401 | 3.1036
M3 | A | 9 | 3.3998 | 3.3714 | 3.3429 | 3.3459 | 3.2873 | 3.2458 | 3.2445 | 3.2005 | 3.1372 | 3.0568 | 3.0247
M3 | A | 10 | 3.2115 | 3.2204 | 3.2074 | 3.2255 | 3.1808 | 3.1433 | 3.1472 | 3.1101 | 3.0518 | 2.976 | 2.9482
M3 | A | 11 | 3.0141 | 3.0654 | 3.0739 | 3.1094 | 3.0785 | 3.0489 | 3.0576 | 3.0278 | 2.9765 | 2.9037 | 2.8805
M3 | A | 12 | 2.8017 | 2.9011 | 2.9318 | 2.9876 | 2.9757 | 2.9524 | 2.967 | 2.9446 | 2.8995 | 2.8321 | 2.813
M3 | A | 13 | 2.5831 | 2.7312 | 2.786 | 2.862 | 2.8703 | 2.8552 | 2.8758 | 2.8622 | 2.8265 | 2.7638 | 2.7495
M3 | A | 14 | 2.3606 | 2.5488 | 2.628 | 2.7249 | 2.7585 | 2.7516 | 2.7799 | 2.7763 | 2.751 | 2.6943 | 2.6841
M3 | A | 15 | 2.1384 | 2.3414 | 2.4378 | 2.5499 | 2.6142 | 2.6176 | 2.6536 | 2.6624 | 2.6521 | 2.6045 | 2.5981
M3 | A | 16 | 1.9351 | 2.1155 | 2.2074 | 2.3149 | 2.4101 | 2.4222 | 2.4648 | 2.488 | 2.504 | 2.4709 | 2.465
M3 | B | 1 | 4.5732 | 4.4606 | 4.3835 | 4.3298 | 4.2521 | 4.1931 | 4.1728 | 4.1024 | 4.0151 | 3.9142 | 3.8531
M3 | B | 2 | 4.4354 | 4.3074 | 4.2219 | 4.1637 | 4.0753 | 4.0095 | 3.9887 | 3.9142 | 3.8202 | 3.7099 | 3.6401
M3 | B | 3 | 4.3233 | 4.1875 | 4.0974 | 4.0387 | 3.9464 | 3.8799 | 3.8591 | 3.7851 | 3.6917 | 3.5813 | 3.5133
M3 | B | 4 | 4.2209 | 4.0808 | 3.9873 | 3.9328 | 3.8389 | 3.7727 | 3.7536 | 3.6804 | 3.5895 | 3.4805 | 3.4179
M3 | B | 5 | 4.1212 | 3.9781 | 3.8845 | 3.8333 | 3.7391 | 3.6751 | 3.6568 | 3.5868 | 3.496 | 3.3888 | 3.3317
M3 | B | 6 | 4.0155 | 3.8729 | 3.7816 | 3.7335 | 3.6408 | 3.5799 | 3.5622 | 3.4936 | 3.4062 | 3.3005 | 3.2483
M3 | B | 7 | 3.9023 | 3.7622 | 3.6737 | 3.6318 | 3.5408 | 3.4819 | 3.4672 | 3.4009 | 3.3161 | 3.2125 | 3.1643
M3 | B | 8 | 3.7817 | 3.6487 | 3.5663 | 3.5327 | 3.4458 | 3.3908 | 3.3767 | 3.315 | 3.2325 | 3.1316 | 3.0876
M3 | B | 9 | 3.6496 | 3.53 | 3.4559 | 3.4324 | 3.3505 | 3.2995 | 3.2876 | 3.2292 | 3.1498 | 3.0517 | 3.0118
M3 | B | 10 | 3.5066 | 3.4084 | 3.3441 | 3.33 | 3.2556 | 3.2085 | 3.1991 | 3.1444 | 3.0704 | 2.9739 | 2.9377
M3 | B | 11 | 3.3611 | 3.2896 | 3.2388 | 3.2353 | 3.1688 | 3.1266 | 3.1202 | 3.0691 | 2.9974 | 2.9049 | 2.8719
M3 | B | 12 | 3.1958 | 3.1629 | 3.1279 | 3.1374 | 3.0823 | 3.0438 | 3.0402 | 2.9933 | 2.9263 | 2.8355 | 2.8055
M3 | B | 13 | 3.0156 | 3.0341 | 3.0156 | 3.0413 | 2.9966 | 2.9637 | 2.9633 | 2.9217 | 2.8572 | 2.7699 | 2.7432
M3 | B | 14 | 2.804 | 2.889 | 2.8936 | 2.9371 | 2.9064 | 2.8791 | 2.8833 | 2.8464 | 2.7875 | 2.7028 | 2.6793
M3 | B | 15 | 2.523 | 2.6917 | 2.7295 | 2.7925 | 2.7871 | 2.7665 | 2.776 | 2.7481 | 2.699 | 2.619 | 2.598
M3 | B | 16 | 2.1647 | 2.4003 | 2.4751 | 2.5608 | 2.5956 | 2.5841 | 2.6039 | 2.5886 | 2.5637 | 2.4967 | 2.4759
M4 | A | 1 | 4.3585 | 4.3131 | 4.2688 | 4.2446 | 4.22 | 4.181 | 4.1686 | 4.1324 | 4.1064 | 4.0745 | 4.0623
M4 | A | 2 | 4.2725 | 4.2348 | 4.1883 | 4.1682 | 4.1402 | 4.1019 | 4.0853 | 4.0479 | 4.0149 | 3.9742 | 3.9617
M4 | A | 3 | 4.1703 | 4.1426 | 4.1025 | 4.0811 | 4.0571 | 4.0165 | 3.9997 | 3.9624 | 3.9214 | 3.8722 | 3.8606
M4 | A | 4 | 4.0593 | 4.0451 | 4.0109 | 3.9941 | 3.9685 | 3.929 | 3.9115 | 3.873 | 3.8279 | 3.776 | 3.7531
M4 | A | 5 | 3.9455 | 3.9427 | 3.9171 | 3.9042 | 3.8782 | 3.84 | 3.8237 | 3.7861 | 3.7362 | 3.676 | 3.6557
M4 | A | 6 | 3.8222 | 3.84 | 3.8227 | 3.8147 | 3.7906 | 3.759 | 3.7389 | 3.7025 | 3.6485 | 3.583 | 3.559
M4 | A | 7 | 3.6959 | 3.7358 | 3.7273 | 3.7238 | 3.7033 | 3.6744 | 3.656 | 3.6233 | 3.565 | 3.4938 | 3.4691
M4 | A | 8 | 3.568 | 3.6202 | 3.6313 | 3.6326 | 3.6207 | 3.5911 | 3.5739 | 3.545 | 3.4832 | 3.4071 | 3.3818
M4 | A | 9 | 3.4308 | 3.5092 | 3.5256 | 3.5411 | 3.5297 | 3.5122 | 3.4926 | 3.4663 | 3.4022 | 3.3214 | 3.2953
M4 | A | 10 | 3.2824 | 3.3837 | 3.4214 | 3.441 | 3.4397 | 3.4218 | 3.4026 | 3.379 | 3.3145 | 3.2277 | 3.2014
M4 | A | 11 | 3.1093 | 3.2359 | 3.2884 | 3.3158 | 3.3242 | 3.3132 | 3.2964 | 3.2794 | 3.2097 | 3.1262 | 3.095
M4 | A | 12 | 2.9006 | 3.0389 | 3.1143 | 3.1564 | 3.1775 | 3.1777 | 3.1654 | 3.1551 | 3.0879 | 3.0016 | 2.9691
M4 | A | 13 | 2.6283 | 2.7739 | 2.8836 | 2.9272 | 2.9737 | 2.9932 | 2.9828 | 2.984 | 2.9296 | 2.845 | 2.8144
M4 | A | 14 | 2.2779 | 2.4172 | 2.5452 | 2.5957 | 2.6807 | 2.7265 | 2.7307 | 2.7541 | 2.7151 | 2.6497 | 2.6203
M4 | A | 15 | 1.9482 | 2.0324 | 2.1429 | 2.1806 | 2.2941 | 2.361 | 2.3925 | 2.4414 | 2.4544 | 2.4117 | 2.3891
M4 | A | 16 | 1.6629 | 1.7046 | 1.7699 | 1.8055 | 1.922 | 1.9877 | 2.036 | 2.0954 | 2.1603 | 2.1519 | 2.1444
M4 | B | 1 | 4.4139 | 4.3434 | 4.2827 | 4.2537 | 4.2206 | 4.1817 | 4.1634 | 4.1243 | 4.0916 | 4.0491 | 4.029
M4 | B | 2 | 4.3395 | 4.2735 | 4.2089 | 4.1793 | 4.1415 | 4.1 | 4.0799 | 4.038 | 3.9961 | 3.9493 | 3.9318
M4 | B | 3 | 4.2576 | 4.1906 | 4.127 | 4.0957 | 4.061 | 4.0164 | 3.9963 | 3.9513 | 3.901 | 3.8452 | 3.8286
M4 | B | 4 | 4.1712 | 4.1065 | 4.0456 | 4.0138 | 3.9745 | 3.9319 | 3.9087 | 3.8624 | 3.8112 | 3.7493 | 3.723
M4 | B | 5 | 4.0844 | 4.0211 | 3.9621 | 3.9289 | 3.8895 | 3.8488 | 3.8232 | 3.7781 | 3.7178 | 3.6514 | 3.6239
M4 | B | 6 | 3.9917 | 3.938 | 3.8796 | 3.8532 | 3.8066 | 3.7652 | 3.7412 | 3.6963 | 3.6332 | 3.5569 | 3.5289
M4 | B | 7 | 3.9005 | 3.8515 | 3.7992 | 3.7713 | 3.7284 | 3.6927 | 3.6631 | 3.617 | 3.5495 | 3.4691 | 3.4384
M4 | B | 8 | 3.8134 | 3.7704 | 3.7207 | 3.6971 | 3.6542 | 3.6176 | 3.587 | 3.5415 | 3.4685 | 3.3854 | 3.3527
M4 | B | 9 | 3.7258 | 3.6907 | 3.6452 | 3.6223 | 3.576 | 3.5424 | 3.5116 | 3.4678 | 3.3909 | 3.3001 | 3.2716
M4 | B | 10 | 3.6291 | 3.6006 | 3.5599 | 3.5349 | 3.5005 | 3.464 | 3.4325 | 3.3876 | 3.3048 | 3.2114 | 3.1791
M4 | B | 11 | 3.5125 | 3.503 | 3.4693 | 3.4503 | 3.408 | 3.3753 | 3.3421 | 3.3002 | 3.2134 | 3.1103 | 3.0762
M4 | B | 12 | 3.3559 | 3.3597 | 3.3419 | 3.3281 | 3.2887 | 3.2621 | 3.2287 | 3.1898 | 3.0992 | 2.994 | 2.9538
M4 | B | 13 | 3.0639 | 3.1453 | 3.1474 | 3.1512 | 3.1285 | 3.1111 | 3.075 | 3.043 | 2.9537 | 2.8478 | 2.8094
M4 | B | 14 | 2.5044 | 2.7555 | 2.8303 | 2.8639 | 2.8824 | 2.8843 | 2.8591 | 2.8473 | 2.7676 | 2.6666 | 2.6316
M4 | B | 15 | 1.7677 | 2.1421 | 2.3293 | 2.4018 | 2.5154 | 2.5535 | 2.5602 | 2.5701 | 2.5378 | 2.4579 | 2.4287
M4 | B | 16 | 1.3308 | 1.6001 | 1.8045 | 1.8991 | 2.084 | 2.148 | 2.194 | 2.244 | 2.2655 | 2.2232 | 2.2046
M5 | A | 1 | 4.0274 | 4.0995 | 4.127 | 4.1598 | 4.1801 | 4.1578 | 4.1654 | 4.1647 | 4.155 | 4.166 | 4.1743
M5 | A | 2 | 3.8225 | 3.8976 | 3.9358 | 3.9679 | 3.9864 | 3.9683 | 3.9746 | 3.973 | 3.9487 | 3.9358 | 3.9327
M5 | A | 3 | 3.6544 | 3.7349 | 3.776 | 3.8082 | 3.8353 | 3.8188 | 3.8224 | 3.8204 | 3.7948 | 3.7643 | 3.753
M5 | A | 4 | 3.5124 | 3.5878 | 3.6424 | 3.6781 | 3.7062 | 3.6945 | 3.7024 | 3.706 | 3.6733 | 3.637 | 3.6213
M5 | A | 5 | 3.3901 | 3.4716 | 3.5283 | 3.569 | 3.6036 | 3.5973 | 3.6077 | 3.6154 | 3.5865 | 3.5459 | 3.5249
M5 | A | 6 | 3.2815 | 3.3691 | 3.4313 | 3.477 | 3.5168 | 3.5195 | 3.5338 | 3.5467 | 3.5206 | 3.4777 | 3.4602
M5 | A | 7 | 3.1831 | 3.2742 | 3.349 | 3.3988 | 3.4509 | 3.4569 | 3.474 | 3.4952 | 3.4744 | 3.4266 | 3.4106
M5 | A | 8 | 3.084 | 3.1812 | 3.2655 | 3.3267 | 3.3889 | 3.4034 | 3.426 | 3.4547 | 3.436 | 3.3934 | 3.3753
M5 | A | 9 | 2.9916 | 3.094 | 3.1907 | 3.2606 | 3.3399 | 3.3616 | 3.3907 | 3.4303 | 3.4133 | 3.3722 | 3.3519
M5 | A | 10 | 2.8904 | 3.0035 | 3.1197 | 3.2006 | 3.3009 | 3.3298 | 3.3695 | 3.4215 | 3.4119 | 3.3687 | 3.3453
M5 | A | 11 | 2.7866 | 2.9119 | 3.0444 | 3.152 | 3.2786 | 3.3228 | 3.3717 | 3.4377 | 3.4322 | 3.3869 | 3.3626
M5 | A | 12 | 2.672 | 2.809 | 2.9748 | 3.1081 | 3.2845 | 3.3492 | 3.409 | 3.4975 | 3.4845 | 3.432 | 3.4038
M5 | A | 13 | 2.5324 | 2.6755 | 2.8911 | 3.0877 | 3.3462 | 3.4323 | 3.5145 | 3.6139 | 3.5956 | 3.5156 | 3.4836
M5 | A | 14 | 2.3516 | 2.4721 | 2.7662 | 3.1101 | 3.5648 | 3.6828 | 3.7729 | 3.862 | 3.7949 | 3.6743 | 3.6313
M5 | A | 15 | 2.0564 | 2.038 | 2.3329 | 3.8294 | 4.46 | 4.3592 | 4.3719 | 4.3306 | 4.1509 | 3.9538 | 3.8871
M5 | A | 16 | 1.5522 | 0.97924 | 0.14762 | 5.8707 | 5.445 | 5.2235 | 5.1042 | 4.9029 | 4.6587 | 4.3968 | 4.2947
M5 | B | 1 | 4.3921 | 4.359 | 4.3135 | 4.2944 | 4.2608 | 4.2145 | 4.194 | 4.1519 | 4.0996 | 4.0613 | 4.0523
M5 | B | 2 | 4.2226 | 4.2005 | 4.158 | 4.1408 | 4.1015 | 4.0573 | 4.0309 | 3.9842 | 3.9162 | 3.8513 | 3.8276
M5 | B | 3 | 4.0627 | 4.0566 | 4.0232 | 4.007 | 3.9722 | 3.9261 | 3.9014 | 3.8526 | 3.7748 | 3.6958 | 3.6669
M5 | B | 4 | 3.9152 | 3.9259 | 3.904 | 3.8948 | 3.8592 | 3.8221 | 3.7967 | 3.75 | 3.6709 | 3.5834 | 3.5478
M5 | B | 5 | 3.7813 | 3.8116 | 3.8036 | 3.8008 | 3.7745 | 3.7375 | 3.715 | 3.6721 | 3.5937 | 3.4992 | 3.4662
M5 | B | 6 | 3.6519 | 3.7077 | 3.7161 | 3.7205 | 3.7029 | 3.6725 | 3.6514 | 3.612 | 3.5339 | 3.4405 | 3.4039
M5 | B | 7 | 3.5239 | 3.6135 | 3.6431 | 3.6588 | 3.6477 | 3.6219 | 3.6035 | 3.5695 | 3.4909 | 3.3992 | 3.3585
M5 | B | 8 | 3.3949 | 3.5241 | 3.5763 | 3.6046 | 3.6038 | 3.5841 | 3.5678 | 3.5393 | 3.4633 | 3.3662 | 3.3266
M5 | B | 9 | 3.2578 | 3.4393 | 3.521 | 3.5631 | 3.5798 | 3.5588 | 3.5491 | 3.5242 | 3.4481 | 3.3476 | 3.3085
M5 | B | 10 | 3.1049 | 3.3583 | 3.4816 | 3.5429 | 3.5711 | 3.5589 | 3.5495 | 3.53 | 3.45 | 3.3419 | 3.302
M5 | B | 11 | 2.9041 | 3.2828 | 3.4678 | 3.559 | 3.603 | 3.5873 | 3.5791 | 3.5604 | 3.4714 | 3.3572 | 3.3115
M5 | B | 12 | 2.6036 | 3.2177 | 3.5133 | 3.6441 | 3.6862 | 3.6677 | 3.6558 | 3.629 | 3.5265 | 3.3946 | 3.3446
M5 | B | 13 | 1.994 | 3.1758 | 3.7418 | 3.8807 | 3.8856 | 3.8418 | 3.8098 | 3.7578 | 3.6227 | 3.467 | 3.4049
M5 | B | 14 | 1.0185 | 6.0416 | 4.6273 | 4.4894 | 4.2955 | 4.1686 | 4.0894 | 3.9831 | 3.7927 | 3.5866 | 3.5098
M5 | B | 15 | 0.43854 | 6.1036 | 5.5092 | 5.1983 | 4.8516 | 4.6577 | 4.5203 | 4.3297 | 4.0666 | 3.7933 | 3.6983
M5 | B | 16 | 0.1692 | 6.0624 | 5.7383 | 5.5311 | 5.2667 | 5.0883 | 4.9596 | 4.7373 | 4.4485 | 4.1238 | 4.0029
M6 | A | 1 | 3.0624 | 3.1513 | 3.3099 | 3.4611 | 3.6964 | 3.7843 | 3.9409 | 4.167 | 4.2858 | 4.3818 | 4.3908
M6 | A | 2 | 3.0256 | 3.0975 | 3.2185 | 3.3258 | 3.5063 | 3.5811 | 3.7052 | 3.9203 | 4.0448 | 4.1456 | 4.156
M6 | A | 3 | 3.0024 | 3.0616 | 3.1594 | 3.2479 | 3.3925 | 3.4624 | 3.5658 | 3.7578 | 3.8774 | 3.9772 | 3.997
M6 | A | 4 | 2.9847 | 3.0368 | 3.1214 | 3.1952 | 3.3182 | 3.3828 | 3.4691 | 3.6449 | 3.7441 | 3.8498 | 3.8677
M6 | A | 5 | 2.9731 | 3.0192 | 3.0942 | 3.1575 | 3.2657 | 3.3222 | 3.3963 | 3.5512 | 3.6471 | 3.7405 | 3.7661
M6 | A | 6 | 2.9603 | 3.0004 | 3.0697 | 3.1254 | 3.2235 | 3.2808 | 3.3494 | 3.4959 | 3.5862 | 3.6754 | 3.6999
M6 | A | 7 | 2.9438 | 2.9792 | 3.0442 | 3.0949 | 3.1878 | 3.2465 | 3.3094 | 3.4567 | 3.5427 | 3.6289 | 3.6581
M6 | A | 8 | 2.9244 | 2.9537 | 3.0144 | 3.0591 | 3.1525 | 3.2075 | 3.2711 | 3.4156 | 3.5084 | 3.5955 | 3.6237
M6 | A | 9 | 2.9 | 2.9243 | 2.9822 | 3.0225 | 3.1095 | 3.171 | 3.2324 | 3.3831 | 3.4773 | 3.57 | 3.6017
M6 | A | 10 | 2.872 | 2.8917 | 2.9432 | 2.9834 | 3.0711 | 3.1299 | 3.1952 | 3.3534 | 3.4542 | 3.5581 | 3.5894
M6 | A | 11 | 2.8348 | 2.8482 | 2.893 | 2.9269 | 3.0144 | 3.0766 | 3.1489 | 3.3263 | 3.4421 | 3.5608 | 3.5978
M6 | A | 12 | 2.7939 | 2.7958 | 2.8343 | 2.8612 | 2.9451 | 3.0202 | 3.0908 | 3.2941 | 3.4369 | 3.5702 | 3.6136
M6 | A | 13 | 2.7381 | 2.7286 | 2.755 | 2.7689 | 2.8426 | 2.9324 | 3.0006 | 3.245 | 3.4236 | 3.5964 | 3.6562
M6 | A | 14 | 2.6706 | 2.633 | 2.6433 | 2.6329 | 2.6789 | 2.775 | 2.8379 | 3.1531 | 3.427 | 3.6914 | 3.7693
M6 | A | 15 | 2.5866 | 2.5194 | 2.4993 | 2.4455 | 2.4277 | 2.5189 | 2.5287 | 2.9354 | 3.4784 | 3.9333 | 4.0431
M6 | A | 16 | 2.4784 | 2.3665 | 2.2886 | 2.1661 | 2.0063 | 2.0256 | 1.8443 | 1.338 | 6.2526 | 5.1432 | 4.9174
M6 | B | 1 | 3.7055 | 3.9002 | 4.0252 | 4.1216 | 4.1737 | 4.1436 | 4.1612 | 4.1634 | 4.0981 | 4.0338 | 4.0062
M6 | B | 2 | 3.555 | 3.7355 | 3.857 | 3.9487 | 4.0092 | 3.9837 | 4.0009 | 4.0055 | 3.9358 | 3.8592 | 3.8218
M6 | B | 3 | 3.4548 | 3.6212 | 3.738 | 3.8271 | 3.8843 | 3.8716 | 3.8874 | 3.8964 | 3.8264 | 3.7414 | 3.7054
M6 | B | 4 | 3.3787 | 3.5255 | 3.6464 | 3.7293 | 3.7906 | 3.7829 | 3.7965 | 3.8099 | 3.7395 | 3.6544 | 3.6198
M6 | B | 5 | 3.3145 | 3.4525 | 3.5644 | 3.6484 | 3.7079 | 3.709 | 3.7205 | 3.7377 | 3.6722 | 3.5906 | 3.5531
M6 | B | 6 | 3.2611 | 3.3955 | 3.5138 | 3.591 | 3.6552 | 3.6586 | 3.6724 | 3.6948 | 3.6271 | 3.5452 | 3.5121
M6 | B | 7 | 3.2129 | 3.3509 | 3.4669 | 3.5433 | 3.6127 | 3.6223 | 3.6388 | 3.6606 | 3.5986 | 3.517 | 3.4814
M6 | B | 8 | 3.1666 | 3.301 | 3.4249 | 3.5027 | 3.581 | 3.5924 | 3.6108 | 3.6365 | 3.5742 | 3.4908 | 3.458
M6 | B | 9 | 3.1173 | 3.2568 | 3.385 | 3.4671 | 3.5512 | 3.5701 | 3.5904 | 3.6201 | 3.5576 | 3.4752 | 3.4382
M6 | B | 10 | 3.0672 | 3.2109 | 3.3463 | 3.4386 | 3.5334 | 3.5552 | 3.5791 | 3.6128 | 3.5506 | 3.462 | 3.4285
M6 | B | 11 | 3.0031 | 3.1553 | 3.307 | 3.41 | 3.5238 | 3.5514 | 3.5782 | 3.6214 | 3.5564 | 3.4588 | 3.4226
M6 | B | 12 | 2.9332 | 3.087 | 3.2649 | 3.3913 | 3.5187 | 3.5584 | 3.587 | 3.64 | 3.5669 | 3.461 | 3.4192
M6 | B | 13 | 2.8408 | 3.0004 | 3.208 | 3.3567 | 3.5312 | 3.5728 | 3.6154 | 3.677 | 3.5947 | 3.4686 | 3.4201
M6 | B | 14 | 2.7071 | 2.8597 | 3.1085 | 3.3048 | 3.5566 | 3.6231 | 3.6763 | 3.7556 | 3.6435 | 3.4884 | 3.436
M6 | B | 15 | 2.531 | 2.6332 | 2.9205 | 3.2173 | 3.6651 | 3.7477 | 3.8282 | 3.9318 | 3.7505 | 3.5421 | 3.4578
M6 | B | 16 | 2.2678 | 2.2023 | 2.3415 | 2.4523 | 4.515 | 4.3206 | 4.3798 | 4.3762 | 4.0393 | 3.6679 | 3.5532
M7 | A | 1 | 2.9169 | 3.0932 | 3.3246 | 3.5338 | 3.8279 | 3.936 | 4.1311 | 4.4009 | 4.5548 | 4.626 | 4.6108
M7 | A | 2 | 2.9177 | 3.0555 | 3.2385 | 3.3965 | 3.6487 | 3.7584 | 3.9427 | 4.2538 | 4.4522 | 4.5559 | 4.5511
M7 | A | 3 | 2.9214 | 3.0306 | 3.1821 | 3.3121 | 3.5147 | 3.6225 | 3.7958 | 4.13 | 4.3608 | 4.5008 | 4.5057
M7 | A | 4 | 2.9388 | 3.0227 | 3.1526 | 3.2547 | 3.4181 | 3.521 | 3.6834 | 4.0321 | 4.2964 | 4.4622 | 4.4754
M7 | A | 5 | 2.9543 | 3.0204 | 3.1294 | 3.2178 | 3.3495 | 3.4534 | 3.5944 | 3.9468 | 4.2431 | 4.4425 | 4.4586
M7 | A | 6 | 2.9767 | 3.0247 | 3.1215 | 3.188 | 3.2983 | 3.4029 | 3.5319 | 3.8936 | 4.2062 | 4.4406 | 4.4617
M7 | A | 7 | 2.9944 | 3.0289 | 3.1112 | 3.1678 | 3.2617 | 3.3577 | 3.4791 | 3.8541 | 4.2034 | 4.458 | 4.4863
M7 | A | 8 | 3.0144 | 3.0331 | 3.1053 | 3.1487 | 3.2243 | 3.3181 | 3.4319 | 3.8241 | 4.2158 | 4.4989 | 4.5293
M7 | A | 9 | 3.031 | 3.0376 | 3.0974 | 3.1337 | 3.1949 | 3.2909 | 3.3957 | 3.8204 | 4.265 | 4.5715 | 4.5929
M7 | A | 10 | 3.0442 | 3.04 | 3.0878 | 3.1154 | 3.1597 | 3.2576 | 3.3604 | 3.834 | 4.3705 | 4.6827 | 4.6932
M7 | A | 11 | 3.0588 | 3.0408 | 3.0778 | 3.0972 | 3.1362 | 3.2225 | 3.3261 | 3.8833 | 4.5432 | 4.8205 | 4.814
M7 | A | 12 | 3.0715 | 3.0434 | 3.0697 | 3.0706 | 3.0918 | 3.1809 | 3.2861 | 4.0041 | 4.7928 | 4.9994 | 4.9614
M7 | A | 13 | 3.0851 | 3.0452 | 3.053 | 3.0442 | 3.0418 | 3.1398 | 3.2367 | 4.2875 | 5.1286 | 5.1913 | 5.128
M7 | A | 14 | 3.0989 | 3.039 | 3.0355 | 3.0092 | 2.962 | 3.0658 | 3.1351 | 4.963 | 5.4792 | 5.3985 | 5.3023
M7 | A | 15 | 3.1091 | 3.032 | 3.01 | 2.9574 | 2.8183 | 2.9116 | 2.6765 | 5.6948 | 5.7516 | 5.5758 | 5.4714
M7 | A | 16 | 3.1253 | 3.0239 | 2.9798 | 2.8678 | 2.5091 | 2.4964 | 0.69358 | 6.0108 | 5.9188 | 5.7194 | 5.6141
M7 | B | 1 | 3.0442 | 3.1323 | 3.2584 | 3.3947 | 3.6048 | 3.7069 | 3.8631 | 4.1192 | 4.2946 | 4.4134 | 4.4385
M7 | B | 2 | 3.0071 | 3.077 | 3.186 | 3.2987 | 3.4879 | 3.5844 | 3.7343 | 4.0063 | 4.1962 | 4.333 | 4.3618
M7 | B | 3 | 2.9766 | 3.033 | 3.1274 | 3.2254 | 3.3891 | 3.4935 | 3.6341 | 3.9155 | 4.1248 | 4.2711 | 4.3063
M7 | B | 4 | 2.9528 | 2.9958 | 3.0833 | 3.164 | 3.311 | 3.4141 | 3.5556 | 3.8517 | 4.0665 | 4.2313 | 4.2688
M7 | B | 5 | 2.932 | 2.965 | 3.0415 | 3.1124 | 3.2494 | 3.3548 | 3.4908 | 3.7961 | 4.0357 | 4.2106 | 4.2444
M7 | B | 6 | 2.9153 | 2.9384 | 3.009 | 3.0755 | 3.1974 | 3.3077 | 3.4386 | 3.766 | 4.0196 | 4.2033 | 4.2425
M7 | B | 7 | 2.9008 | 2.9144 | 2.978 | 3.0334 | 3.1504 | 3.257 | 3.3963 | 3.7514 | 4.0281 | 4.22 | 4.2547
M7 | B | 8 | 2.8821 | 2.8877 | 2.9404 | 2.9881 | 3.0968 | 3.2146 | 3.3533 | 3.7415 | 4.0501 | 4.2519 | 4.2914
M7 | B | 9 | 2.8649 | 2.8631 | 2.907 | 2.944 | 3.0465 | 3.1677 | 3.3112 | 3.7513 | 4.1043 | 4.3145 | 4.3414
M7 | B | 10 | 2.8455 | 2.8307 | 2.8682 | 2.8946 | 2.9864 | 3.1127 | 3.2709 | 3.7954 | 4.197 | 4.3989 | 4.4257
M7 | B | 11 | 2.8244 | 2.7965 | 2.8278 | 2.8411 | 2.9147 | 3.0534 | 3.2195 | 3.8558 | 4.3432 | 4.5173 | 4.5296
M7 | B | 12 | 2.7997 | 2.7601 | 2.7752 | 2.773 | 2.8162 | 2.97 | 3.1499 | 4.0052 | 4.5618 | 4.6732 | 4.6546
M7 | B | 13 | 2.7734 | 2.7176 | 2.7189 | 2.6921 | 2.6854 | 2.8459 | 3.0233 | 4.3294 | 4.8647 | 4.8678 | 4.8209
M7 | B | 14 | 2.7403 | 2.6651 | 2.6427 | 2.5874 | 2.495 | 2.6278 | 2.6704 | 5.0115 | 5.1986 | 5.0862 | 5.0064
M7 | B | 15 | 2.7029 | 2.6036 | 2.5546 | 2.4406 | 2.1822 | 2.2221 | 1.3839 | 5.6615 | 5.5014 | 5.2923 | 5.1977
M7 | B | 16 | 2.6607 | 2.5325 | 2.4392 | 2.2544 | 1.7646 | 1.5564 | 0.62787 | 5.9722 | 5.7125 | 5.4811 | 5.3747

## RSB Stray Light
Band | Band Median |  |  |  | Band Average
| dn | L | R | Uncertainty | dn | L | R | Uncertainty
I1 | 0.30951 | 0.072312 | 0.32869 | 1.8791 | 0.30961 | 0.072015 | 0.32734 | 1.8703
I2 | 0.32145 | 0.037713 | 0.15085 | 0.47495 | 0.33373 | 0.038516 | 0.15406 | 0.49083
I3 | 0.24069 | 0.0044996 | 0.061638 | 0.33696 | 0.24041 | 0.0045588 | 0.06245 | 0.75056
M1 | 2.6582 | 0.14636 | 0.32596 | 2.2874 | 2.5953 | 0.14406 | 0.32084 | 2.3321
M2 | 2.6825 | 0.11709 | 0.29272 | 1.6336 | 2.6884 | 0.11767 | 0.29417 | 1.668
M3 | 2.8809 | 0.11192 | 0.34975 | 0.99758 | 2.881 | 0.11174 | 0.34918 | 0.99202
M4 | 2.9285 | 0.082802 | 0.3943 | 0.81818 | 2.9078 | 0.082319 | 0.39199 | 0.81001
M5 | 2.5873 | 0.054369 | 0.54369 | 0.80181 | 2.607 | 0.054552 | 0.54552 | 0.8149
M6 | 2.9759 | 0.038963 | 0.40586 | 0.42887 | 2.9868 | 0.039094 | 0.40723 | 0.42725
M7 | 3.3266 | 0.033521 | 0.52377 | 0.26802 | 3.3611 | 0.033903 | 0.52974 | 0.26765
M8 | 0.43954 | 0.013224 | 0.24488 | 0.23007 | 0.44497 | 0.013188 | 0.24422 | 0.23571
M9 | 0.35918 | 0.0089333 | 0.14889 | 0.2521 | 0.36133 | 0.0090208 | 0.15035 | 0.25267
M10 | 0.21602 | 0.0046334 | 0.063471 | 0.095896 | 0.20905 | 0.00448 | 0.06137 | 0.092353
M11 | 0.084585 | 0.00074511 | 0.62092 | 2.5077 | 0.090609 | 0.000792 | 0.66 | 2.5106
Band | Detector | Detector Dependent
|  | dn | L | R | Uncertainty |  | Notes:
1. This table provides VIIRS F2 RSB stray light rejection.
I1 | 1 | 0.1694 | 0.038729 | 0.17604 | 1.2894 |  | 2. The results were derived from FP-12 using the studio lamp and ABC blackbody.
I1 | 2 | 0.29519 | 0.067502 | 0.30683 | 1.869 |  | 3. dn and L are the measured stray light response and radiance. R is the ratio of the modeled Earth and measured irradiances.
I1 | 3 | 0.32752 | 0.075018 | 0.34099 | 1.985 |  | 4. The testing was conducted at Raytheon El Segundo, winter 2013/2014.
I1 | 4 | 0.33989 | 0.078131 | 0.35514 | 2.2214 |  | 5. The test was set to diagnostic mode.
I1 | 5 | 0.37 | 0.085006 | 0.38639 | 2.147
I1 | 6 | 0.39815 | 0.091356 | 0.41525 | 1.9153
I1 | 7 | 0.27759 | 0.064221 | 0.29191 | 1.762
I1 | 8 | 0.27787 | 0.063951 | 0.29069 | 2.0837
I1 | 9 | 0.30871 | 0.070841 | 0.322 | 1.876
I1 | 10 | 0.30943 | 0.071093 | 0.32315 | 1.8822
I1 | 11 | 0.35306 | 0.082599 | 0.37545 | 1.9532
I1 | 12 | 0.41964 | 0.098196 | 0.44634 | 1.9272
I1 | 13 | 0.22153 | 0.051523 | 0.2342 | 1.4501
I1 | 14 | 0.33331 | 0.077604 | 0.35274 | 1.893
I1 | 15 | 0.38701 | 0.090754 | 0.41252 | 1.9755
I1 | 16 | 0.29172 | 0.068723 | 0.31238 | 1.8899
I1 | 17 | 0.2868 | 0.067026 | 0.30466 | 1.7968
I1 | 18 | 0.27483 | 0.063829 | 0.29013 | 1.7284
I1 | 19 | 0.33648 | 0.077761 | 0.35346 | 1.9896
I1 | 20 | 0.20692 | 0.04768 | 0.21673 | 1.6191
I1 | 21 | 0.33941 | 0.078726 | 0.35785 | 1.8318
I1 | 22 | 0.51579 | 0.11943 | 0.54287 | 2.1104
I1 | 23 | 0.31473 | 0.073648 | 0.33476 | 1.9755
I1 | 24 | 0.23547 | 0.055836 | 0.2538 | 1.7551
I1 | 25 | 0.34405 | 0.081833 | 0.37197 | 1.8309
I1 | 26 | 0.24613 | 0.05824 | 0.26473 | 1.8831
I1 | 27 | 0.30464 | 0.072024 | 0.32738 | 1.822
I1 | 28 | 0.37252 | 0.086751 | 0.39432 | 2.237
I1 | 29 | 0.32698 | 0.076308 | 0.34686 | 1.7773
I1 | 30 | 0.20612 | 0.048326 | 0.21966 | 1.7153
I1 | 31 | 0.20692 | 0.049216 | 0.22371 | 1.828
I1 | 32 | 0.30959 | 0.0726 | 0.33 | 1.828
I2 | 1 | 0.43766 | 0.049477 | 0.19791 | 0.56074
I2 | 2 | 0.30623 | 0.034896 | 0.13958 | 0.46641
I2 | 3 | 0.39005 | 0.044632 | 0.17853 | 0.4521
I2 | 4 | 0.27887 | 0.031638 | 0.12655 | 0.378
I2 | 5 | 0.23151 | 0.026265 | 0.10506 | 0.3954
I2 | 6 | 0.29988 | 0.033991 | 0.13597 | 0.4592
I2 | 7 | 0.41652 | 0.047337 | 0.18935 | 0.55573
I2 | 8 | 0.36462 | 0.040637 | 0.16255 | 0.50256
I2 | 9 | 0.34488 | 0.038437 | 0.15375 | 0.48193
I2 | 10 | 0.35139 | 0.039127 | 0.15651 | 0.46797
I2 | 11 | 0.39015 | 0.044223 | 0.17689 | 0.45475
I2 | 12 | 0.34463 | 0.03903 | 0.15612 | 0.48878
I2 | 13 | 0.43991 | 0.050216 | 0.20086 | 0.553
I2 | 14 | 0.26629 | 0.030637 | 0.12255 | 0.41658
I2 | 15 | 0.29349 | 0.033971 | 0.13588 | 0.43374
I2 | 16 | 0.29063 | 0.033721 | 0.13488 | 0.46635
I2 | 17 | 0.4976 | 0.058045 | 0.23218 | 0.63746
I2 | 18 | 0.33314 | 0.038394 | 0.15358 | 0.51683
I2 | 19 | 0.28996 | 0.033621 | 0.13448 | 0.43107
I2 | 20 | 0.29243 | 0.033761 | 0.13504 | 0.51787
I2 | 21 | 0.29967 | 0.034949 | 0.1398 | 0.48381
I2 | 22 | 0.32353 | 0.037384 | 0.14954 | 0.51157
I2 | 23 | 0.32016 | 0.037699 | 0.1508 | 0.45825
I2 | 24 | 0.30792 | 0.036257 | 0.14503 | 0.53733
I2 | 25 | 0.28218 | 0.033678 | 0.13471 | 0.41409
I2 | 26 | 0.23602 | 0.028098 | 0.11239 | 0.45697
I2 | 27 | 0.32274 | 0.037728 | 0.15091 | 0.4663
I2 | 28 | 0.23851 | 0.027798 | 0.11119 | 0.45925
I2 | 29 | 0.35101 | 0.041191 | 0.16476 | 0.60412
I2 | 30 | 0.36604 | 0.042698 | 0.17079 | 0.52675
I2 | 31 | 0.45574 | 0.054985 | 0.21994 | 0.66693
I2 | 32 | 0.31596 | 0.037978 | 0.15191 | 0.48478
I3 | 1 | 0.22831 | 0.0043721 | 0.059891 | 0.33493
I3 | 2 | 0.22483 | 0.0042043 | 0.057593 | 0.33016
I3 | 3 | 0.25678 | 0.0047761 | 0.065426 | 0.34578
I3 | 4 | 0.10666 | 0.0037491 | 0.051358 | 13.5532
I3 | 5 | 0.2558 | 0.0047323 | 0.064826 | 0.35162
I3 | 6 | 0.29317 | 0.005409 | 0.074096 | 0.33879
I3 | 7 | 0.26146 | 0.0048109 | 0.065903 | 0.34997
I3 | 8 | 0.26375 | 0.0048529 | 0.066478 | 0.32916
I3 | 9 | 0.2404 | 0.0044474 | 0.060923 | 0.31217
I3 | 10 | 0.23123 | 0.004324 | 0.059233 | 0.33387
I3 | 11 | 0.26157 | 0.0048717 | 0.066736 | 0.32563
I3 | 12 | 0.23937 | 0.0044163 | 0.060497 | 0.33503
I3 | 13 | 0.25897 | 0.0047391 | 0.06492 | 0.31813
I3 | 14 | 0.22875 | 0.0042662 | 0.05844 | 0.33745
I3 | 15 | 0.2507 | 0.0046317 | 0.063447 | 0.32424
I3 | 16 | 0.22247 | 0.0041379 | 0.056683 | 0.32946
I3 | 17 | 0.23496 | 0.0044054 | 0.060348 | 0.33316
I3 | 18 | 0.24098 | 0.0044942 | 0.061565 | 0.34533
I3 | 19 | 0.25237 | 0.0047193 | 0.064648 | 0.34016
I3 | 20 | 0.22281 | 0.0041888 | 0.057381 | 0.33562
I3 | 21 | 0.25115 | 0.0047216 | 0.06468 | 0.33425
I3 | 22 | 0.23586 | 0.0044694 | 0.061225 | 0.36009
I3 | 23 | 0.26048 | 0.0049166 | 0.067351 | 0.34238
I3 | 24 | 0.25683 | 0.004854 | 0.066493 | 0.34727
I3 | 25 | 0.23835 | 0.0045049 | 0.061711 | 0.32571
I3 | 26 | 0.23684 | 0.0044763 | 0.061319 | 0.3481
I3 | 27 | 0.22734 | 0.0042966 | 0.058858 | 0.33648
I3 | 28 | 0.21484 | 0.0040713 | 0.055771 | 0.3274
I3 | 29 | 0.28709 | 0.0054546 | 0.074721 | 0.34053
I3 | 30 | 0.21512 | 0.0041088 | 0.056285 | 0.35613
I3 | 31 | 0.24227 | 0.0046274 | 0.063389 | 0.33815
I3 | 32 | 0.25166 | 0.0048318 | 0.066189 | 0.35772
M1 | 1 | 2.6786 | 0.15107 | 0.33646 | 2.2887
M1 | 2 | 2.481 | 0.14005 | 0.31192 | 2.4191
M1 | 3 | 2.7218 | 0.15256 | 0.33977 | 2.4012
M1 | 4 | 2.8677 | 0.1593 | 0.35479 | 2.5255
M1 | 5 | 2.2071 | 0.12172 | 0.2711 | 2.2444
M1 | 6 | 2.775 | 0.15276 | 0.34023 | 2.4511
M1 | 7 | 2.5402 | 0.14022 | 0.3123 | 2.3957
M1 | 8 | 2.233 | 0.12282 | 0.27354 | 2.1901
M1 | 9 | 2.6378 | 0.14468 | 0.32223 | 2.2627
M1 | 10 | 2.8382 | 0.1561 | 0.34766 | 2.278
M1 | 11 | 2.6915 | 0.14803 | 0.32969 | 2.2511
M1 | 12 | 3.1556 | 0.1745 | 0.38865 | 2.3948
M1 | 13 | 2.1023 | 0.11626 | 0.25893 | 2.2463
M1 | 14 | 2.3386 | 0.12956 | 0.28856 | 2.2861
M1 | 15 | 2.8837 | 0.16149 | 0.35966 | 2.474
M1 | 16 | 2.3723 | 0.13379 | 0.29798 | 2.2054
M2 | 1 | 2.5861 | 0.1143 | 0.28576 | 1.5987
M2 | 2 | 3.0191 | 0.13405 | 0.33512 | 1.803
M2 | 3 | 2.896 | 0.12801 | 0.32001 | 1.5994
M2 | 4 | 3.1587 | 0.1393 | 0.34824 | 1.7924
M2 | 5 | 2.8174 | 0.12397 | 0.30991 | 1.6037
M2 | 6 | 2.1624 | 0.094713 | 0.23678 | 1.705
M2 | 7 | 2.7187 | 0.11854 | 0.29634 | 1.6123
M2 | 8 | 2.3722 | 0.10295 | 0.25738 | 1.616
M2 | 9 | 2.9123 | 0.12625 | 0.31562 | 1.8614
M2 | 10 | 2.9671 | 0.12877 | 0.32193 | 1.6512
M2 | 11 | 2.6462 | 0.11564 | 0.2891 | 1.7118
M2 | 12 | 2.5734 | 0.11259 | 0.28147 | 1.6589
M2 | 13 | 2.2281 | 0.097143 | 0.24286 | 1.5786
M2 | 14 | 2.6216 | 0.1143 | 0.28575 | 1.5558
M2 | 15 | 2.3263 | 0.10131 | 0.25327 | 1.6055
M2 | 16 | 3.0086 | 0.13087 | 0.32719 | 1.7341
M3 | 1 | 2.9285 | 0.1126 | 0.35188 | 1.0031
M3 | 2 | 2.9111 | 0.11252 | 0.35161 | 0.98804
M3 | 3 | 2.7522 | 0.1061 | 0.33156 | 0.94674
M3 | 4 | 3.2669 | 0.12561 | 0.39253 | 1.0642
M3 | 5 | 2.8507 | 0.10947 | 0.34208 | 0.93563
M3 | 6 | 2.7475 | 0.10564 | 0.33013 | 0.91328
M3 | 7 | 2.849 | 0.10997 | 0.34366 | 0.99527
M3 | 8 | 3.0944 | 0.11944 | 0.37327 | 0.9998
M3 | 9 | 2.9653 | 0.11475 | 0.35861 | 1.0438
M3 | 10 | 2.9222 | 0.11324 | 0.35386 | 0.99536
M3 | 11 | 2.9766 | 0.11534 | 0.36045 | 1.0125
M3 | 12 | 2.6284 | 0.10224 | 0.31951 | 0.88937
M3 | 13 | 3.1756 | 0.12433 | 0.38852 | 1.0926
M3 | 14 | 2.7358 | 0.10738 | 0.33556 | 0.93688
M3 | 15 | 2.484 | 0.097869 | 0.30584 | 1.0472
M3 | 16 | 2.8077 | 0.11132 | 0.34789 | 1.0084
M4 | 1 | 2.6573 | 0.074538 | 0.35494 | 0.76341
M4 | 2 | 2.8644 | 0.080634 | 0.38397 | 0.78223
M4 | 3 | 3.0437 | 0.085528 | 0.40727 | 0.82552
M4 | 4 | 3.0847 | 0.08668 | 0.41276 | 0.81529
M4 | 5 | 2.7926 | 0.078333 | 0.37301 | 0.77807
M4 | 6 | 3.0526 | 0.085778 | 0.40847 | 0.82264
M4 | 7 | 2.7864 | 0.078298 | 0.37285 | 0.77832
M4 | 8 | 2.9415 | 0.083097 | 0.3957 | 0.83627
M4 | 9 | 3.054 | 0.086122 | 0.4101 | 0.85653
M4 | 10 | 3.0041 | 0.085318 | 0.40627 | 0.83352
M4 | 11 | 2.9155 | 0.082508 | 0.39289 | 0.8398
M4 | 12 | 3.0705 | 0.087355 | 0.41598 | 0.83452
M4 | 13 | 3.0315 | 0.086702 | 0.41286 | 0.82108
M4 | 14 | 2.8222 | 0.081138 | 0.38637 | 0.79579
M4 | 15 | 2.6861 | 0.077227 | 0.36775 | 0.79245
M4 | 16 | 2.7172 | 0.077849 | 0.37071 | 0.78477
M5 | 1 | 2.8549 | 0.059239 | 0.59239 | 0.85598
M5 | 2 | 2.5157 | 0.052704 | 0.52704 | 0.79663
M5 | 3 | 2.5496 | 0.053159 | 0.53159 | 0.79373
M5 | 4 | 2.6823 | 0.055926 | 0.55926 | 0.82448
M5 | 5 | 2.5139 | 0.052164 | 0.52164 | 0.80046
M5 | 6 | 2.8178 | 0.058751 | 0.58751 | 0.83917
M5 | 7 | 2.612 | 0.05446 | 0.5446 | 0.7978
M5 | 8 | 2.474 | 0.051335 | 0.51335 | 0.79186
M5 | 9 | 2.6763 | 0.055802 | 0.55802 | 0.79822
M5 | 10 | 2.6516 | 0.055286 | 0.55286 | 0.82711
M5 | 11 | 2.5406 | 0.053479 | 0.53479 | 0.80316
M5 | 12 | 2.5188 | 0.052894 | 0.52894 | 0.79191
M5 | 13 | 2.5785 | 0.054278 | 0.54278 | 0.82519
M5 | 14 | 2.7448 | 0.057777 | 0.57777 | 0.87018
M5 | 15 | 2.3846 | 0.050673 | 0.50673 | 0.79482
M5 | 16 | 2.5961 | 0.054908 | 0.54908 | 0.82777
M6 | 1 | 3.047 | 0.039306 | 0.40944 | 0.40023
M6 | 2 | 2.9095 | 0.037823 | 0.39399 | 0.4089
M6 | 3 | 2.8753 | 0.037379 | 0.38937 | 0.40447
M6 | 4 | 3.0613 | 0.03949 | 0.41136 | 0.42861
M6 | 5 | 3.1406 | 0.040671 | 0.42366 | 0.41185
M6 | 6 | 3.2437 | 0.04233 | 0.44094 | 0.4369
M6 | 7 | 2.8719 | 0.037335 | 0.3889 | 0.41265
M6 | 8 | 2.8706 | 0.037318 | 0.38873 | 0.40471
M6 | 9 | 3.0965 | 0.040564 | 0.42254 | 0.46301
M6 | 10 | 2.9368 | 0.038619 | 0.40229 | 0.43521
M6 | 11 | 3.0581 | 0.040213 | 0.41889 | 0.45578
M6 | 12 | 3.0554 | 0.040179 | 0.41853 | 0.45363
M6 | 13 | 2.8901 | 0.038439 | 0.4004 | 0.43012
M6 | 14 | 3.0149 | 0.039948 | 0.41613 | 0.44087
M6 | 15 | 2.8756 | 0.038102 | 0.3969 | 0.42914
M6 | 16 | 2.841 | 0.037785 | 0.39359 | 0.41998
M7 | 1 | 3.3109 | 0.033274 | 0.51992 | 0.2729
M7 | 2 | 3.3872 | 0.03421 | 0.53454 | 0.26798
M7 | 3 | 3.3573 | 0.033573 | 0.52457 | 0.25051
M7 | 4 | 3.4196 | 0.034196 | 0.53431 | 0.27231
M7 | 5 | 3.2786 | 0.033114 | 0.51741 | 0.26546
M7 | 6 | 3.5259 | 0.035259 | 0.55092 | 0.29223
M7 | 7 | 3.4781 | 0.035129 | 0.54888 | 0.26681
M7 | 8 | 3.3139 | 0.03347 | 0.52297 | 0.25003
M7 | 9 | 3.3032 | 0.033362 | 0.52129 | 0.26484
M7 | 10 | 3.5369 | 0.035723 | 0.55817 | 0.26806
M7 | 11 | 3.2886 | 0.033215 | 0.51898 | 0.27146
M7 | 12 | 3.2806 | 0.033134 | 0.51772 | 0.26436
M7 | 13 | 3.3393 | 0.033894 | 0.52959 | 0.26791
M7 | 14 | 3.5049 | 0.0354 | 0.55312 | 0.26806
M7 | 15 | 3.2658 | 0.033311 | 0.52049 | 0.27008
M7 | 16 | 3.1873 | 0.032191 | 0.50299 | 0.26941
M8 | 1 | 0.4661 | 0.011956 | 0.2214 | 0.2016
M8 | 2 | 0.47958 | 0.012373 | 0.22913 | 0.23328
M8 | 3 | 0.37992 | 0.010068 | 0.18644 | 0.21621
M8 | 4 | 0.57699 | 0.016156 | 0.29918 | 0.2752
M8 | 5 | 0.45254 | 0.012739 | 0.23591 | 0.22794
M8 | 6 | 0.43689 | 0.012648 | 0.23422 | 0.20373
M8 | 7 | 0.3531 | 0.010505 | 0.19453 | 0.22232
M8 | 8 | 0.45732 | 0.013788 | 0.25534 | 0.2737
M8 | 9 | 0.44219 | 0.01373 | 0.25426 | 0.25832
M8 | 10 | 0.45673 | 0.013907 | 0.25754 | 0.25523
M8 | 11 | 0.42667 | 0.013355 | 0.24731 | 0.1841
M8 | 12 | 0.41762 | 0.013092 | 0.24245 | 0.2647
M8 | 13 | 0.43154 | 0.013744 | 0.25453 | 0.2322
M8 | 14 | 0.5228 | 0.016599 | 0.30739 | 0.30574
M8 | 15 | 0.39648 | 0.012747 | 0.23605 | 0.20597
M8 | 16 | 0.42301 | 0.0136 | 0.25185 | 0.21113
M9 | 1 | 0.39221 | 0.009119 | 0.15198 | 0.24441
M9 | 2 | 0.40165 | 0.0093383 | 0.15564 | 0.23806
M9 | 3 | 0.35671 | 0.0083648 | 0.13941 | 0.23128
M9 | 4 | 0.3669 | 0.008824 | 0.14707 | 0.23976
M9 | 5 | 0.37104 | 0.0088864 | 0.14811 | 0.2307
M9 | 6 | 0.34677 | 0.0084093 | 0.14016 | 0.22935
M9 | 7 | 0.37861 | 0.0093327 | 0.15554 | 0.25138
M9 | 8 | 0.42171 | 0.010501 | 0.17501 | 0.26702
M9 | 9 | 0.34473 | 0.0086526 | 0.14421 | 0.26598
M9 | 10 | 0.35043 | 0.0089185 | 0.14864 | 0.26788
M9 | 11 | 0.36166 | 0.0092222 | 0.1537 | 0.25117
M9 | 12 | 0.31179 | 0.0080908 | 0.13485 | 0.25444
M9 | 13 | 0.36744 | 0.0096086 | 0.16014 | 0.28549
M9 | 14 | 0.32881 | 0.0087134 | 0.14522 | 0.25281
M9 | 15 | 0.33576 | 0.0089482 | 0.14914 | 0.25729
M9 | 16 | 0.34508 | 0.0094035 | 0.15673 | 0.27568
M10 | 1 | 0.18383 | 0.0038604 | 0.052882 | 0.10655
M10 | 2 | 0.18543 | 0.0038477 | 0.052709 | 0.095889
M10 | 3 | 0.17796 | 0.0037104 | 0.050828 | 0.088114
M10 | 4 | 0.28357 | 0.0059266 | 0.081186 | 0.078946
M10 | 5 | 0.12387 | 0.0026013 | 0.035634 | 0.095903
M10 | 6 | 0.25639 | 0.0054739 | 0.074984 | 0.084465
M10 | 7 | 0.21764 | 0.0046249 | 0.063354 | 0.10514
M10 | 8 | 0.2144 | 0.0046419 | 0.063587 | 0.10776
M10 | 9 | 0.21401 | 0.0046013 | 0.063031 | 0.10899
M10 | 10 | 0.22638 | 0.0048897 | 0.066983 | 0.084541
M10 | 11 | 0.27922 | 0.0060171 | 0.082427 | 0.095091
M10 | 12 | 0.23806 | 0.0052136 | 0.071419 | 0.1079
M10 | 13 | 0.24753 | 0.0053467 | 0.073242 | 0.11106
M10 | 14 | 0.16215 | 0.0035592 | 0.048756 | 0.096583
M10 | 15 | 0.23376 | 0.005131 | 0.070287 | 0.068662
M10 | 16 | 0.10064 | 0.0022342 | 0.030605 | 0.042044
M11 | 1 | 0.083588 | 0.00074393 | 0.61994 | 2.7589
M11 | 2 | 0.11774 | 0.0010008 | 0.834 | 2.7014
M11 | 3 | 0.062542 | 0.00055662 | 0.46385 | 2.0244
M11 | 4 | 0.048833 | 0.00043462 | 0.36218 | 2.6025
M11 | 5 | 0.10674 | 0.00091794 | 0.76495 | 2.791
M11 | 6 | 0.10831 | 0.00093151 | 0.77626 | 2.7716
M11 | 7 | 0.15865 | 0.0013802 | 1.1502 | 2.8169
M11 | 8 | 0.08285 | 0.00074565 | 0.62138 | 2.3383
M11 | 9 | 0.073634 | 0.00065534 | 0.54612 | 2.26
M11 | 10 | 0.083275 | 0.00073282 | 0.61069 | 2.8522
M11 | 11 | 0.060297 | 0.00051856 | 0.43213 | 2.3712
M11 | 12 | 0.058178 | 0.00050324 | 0.41936 | 2.2338
M11 | 13 | 0.093928 | 0.00080778 | 0.67315 | 2.2278
M11 | 14 | 0.085582 | 0.00074456 | 0.62047 | 2.413
M11 | 15 | 0.094796 | 0.0008342 | 0.69517 | 2.2783
M11 | 16 | 0.13081 | 0.0011642 | 0.97015 | 2.7287