## Summary
NASA Summary of J1 VIIRS Pre-Themal Vacuum and Thermal Vacuum Spectral Calibration
Contents: data list and spectral band performance metrics (center wavelength, centroid wavelength, bandwifth, 1% limits, ND transmittence, and IOOB)
All quantities are listed by electronics side and subsample. Detector dependent and detector averaged quantities are listed.
Version 1: | Initial results from pre-thermal vacuum and thermal vacuum testing.
2014-10-15

## VIIRS Data
Band | E-Side | Test | UAID | Collects | Det Row | Wave Start | Wave End | Delta Wave | Grating | Order Filter | ND Filter OD | Slit Widths |  |  | SpMA Source | SpMA Bulb | RSO Run | RSO Row | Date | Test Phase | Smile Correction | Comments
[12 empty] | Entrance | Middle | Exit
I1 | A | FP15 | 4302602 | 40 | All | 563 | 719 | 4 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | B |  | B | 2014-07-10 | Pre-TV
I1 | B | FP15 | 4302870 | 40 | All | 563 | 719 | 4 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | D |  | B | 2014-08-14 | Pre-TV |  | Retest of FP15 with bulb D after lock-in amplifier fix
I1 | A | FP16 | 4302606 | 40 | All | 347 | 561.5 | 5.5 | 1 | 5 | None | 2000 | 1712 | 510 | Tungsten | B |  | A | 2014-07-10 | Pre-TV
I1 | A | FP16 | 4302607 | 28 | All | 563 | 725 | 6 | 1 | 5 | None | 2000 | 1712 | 510 | Tungsten | B |  | A | 2014-07-10 | Pre-TV
I1 | A | FP16 | 4302608 | 25 | All | 581 | 725 | 6 | 1 | 2 | None | 2000 | 1712 | 510 | Tungsten | B |  | C | 2014-07-10 | Pre-TV
I1 | A | FP16 | 4302609 | 69 | All | 730.5 | 1104.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 510 | Tungsten | B |  | C | 2014-07-10 | Pre-TV
I2 | B | FP15 | 4302535 | 45 | All | 799 | 931 | 3 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
I2 | B | FP16 | 4303539 | 23 | All | 799 | 931 | 6 | 1 | 2 | 0.1 | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
I2 | B | FP16 | 4302540 | 52 | All | 352 | 632.5 | 5.5 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | B |  | A | 2014-07-08 | Pre-TV
I2 | B | FP16 | 4301541 | 36 | All | 605 | 797.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
I2 | B | FP16 | 4300542 | 23 | All | 799 | 931 | 6 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
I2 | B | FP16 | 4299543 | 31 | All | 936.5 | 1101.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
I3 | A | FP15 | 4303383 | 211 | Odd | 1506 | 1715 | 0.99524 | 2 | 7 | None | 2000 | 976 | 1150 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad slit alignment
I3 | A | FP15 | 4303395 | 211 | Odd | 1506 | 1715 | 0.99524 | 2 | 7 | None | 2000 | 80 | 96 | Tungsten | E |  | F | 2014-09-10 | TV - Nom
I3 | A | FP16 | 4303399 | 19 | Odd | 1506 | 1722 | 12 | 2 | 7 | 1.8 | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303400 | 43 | Odd | 990 | 1494 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303401 | 19 | Odd | 1506 | 1722 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303402 | 44 | Odd | 1734 | 2250 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303403 | 19 | Odd | 2200 | 2416 | 12 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | E |  | G | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303404 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | E |  | H | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303405 | 50 | Odd | 2810 | 5553.9 | 55.998 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-11 | TV - Nom
I3 | A | FP16 | 4303406 | 33 | Odd | 5274 | 7066.1 | 56.0031 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-11 | TV - Nom
I3 | A | FP15 | 4303446 | 211 | Even | 1506 | 1716 | 1 | 2 | 7 | None | 2000 | 76 | 90 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303450 | 19 | Even | 1506 | 1722 | 12 | 2 | 7 | 1.8 | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303451 | 43 | Even | 990 | 1494 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303452 | 19 | Even | 1506 | 1722 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303453 | 44 | Even | 1734 | 2250 | 12 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303454 | 19 | Even | 2200 | 2416 | 12 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | E |  | G | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303455 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | E |  | H | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303456 | 50 | Even | 2810 | 5553.9 | 55.998 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-13 | TV - Nom
I3 | A | FP16 | 4303457 | 33 | Even | 5274 | 7066.1 | 56.0031 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-13 | TV - Nom
I4 | B | FP15 | 4303667 | 211 | Odd | 3326.1 | 4166 | 3.9995 | 3 | 12 | None | 2000 | 150 | 175 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303671 | 16 | Odd | 3326.1 | 4166 | 55.9933 | 3 | 12 | 1.3 | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303672 | 90 | Odd | 1000 | 2246 | 14 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | F |  | F | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303673 | 16 | Odd | 2200 | 2410 | 14 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | F |  | G | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303674 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | F |  | H | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303675 | 8 | Odd | 2878 | 3270.1 | 56.0143 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303676 | 16 | Odd | 3326.1 | 4166 | 55.9933 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303677 | 24 | Odd | 4222.1 | 5510 | 55.9957 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
I4 | B | FP16 | 4303678 | 35 | Odd | 5229.9 | 7134 | 56.0029 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-23 | TV - Nom
I4 | B | FP15 | 4303696 | 211 | Even | 3326.1 | 4166 | 3.9995 | 3 | 12 | None | 2000 | 150 | 175 | Ceramic | A |  | I | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303700 | 16 | Even | 3326.1 | 4166 | 55.9933 | 3 | 12 | 1.3 | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303701 | 90 | Even | 1000 | 2246 | 14 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | F |  | F | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303702 | 16 | Even | 2200 | 2410 | 14 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | F |  | G | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303703 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | F |  | H | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303704 | 8 | Even | 2878 | 3270.1 | 56.0143 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303705 | 16 | Even | 3326.1 | 4166 | 55.9933 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303706 | 24 | Even | 4222.1 | 5510 | 55.9957 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-25 | TV - Nom
I4 | B | FP16 | 4303707 | 35 | Even | 5229.9 | 7134 | 56.0029 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-25 | TV - Nom
I5 | A | FP15 | 4303409 | 44 | Odd | 9852.1 | 13291.9 | 79.9953 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303414 | 29 | Odd | 1804 | 2251 | 15.9643 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303415 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | E |  | G | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303416 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | E |  | H | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303417 | 47 | Odd | 2900 | 5476 | 56 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303418 | 29 | Odd | 5400 | 7080 | 60 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303419 | 25 | Odd | 7030.2 | 9909.9 | 119.9875 | 4 | 11 | None | 2000 | 1712 | 590 | Ceramic | A |  | K | 2014-09-11 | TV - Nom
I5 | A | FP16 | 4303420 | 68 | Odd | 9756.8 | 16121.9 | 95.0015 | 4 | 16 | None | 2000 | 1712 | 590 | Ceramic | A |  | L | 2014-09-11 | TV - Nom
I5 | A | FP15 | 4303459 | 44 | Even | 9852.1 | 13291.9 | 79.9953 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-13 | TV - Nom
I5 | A | FP16 | 4303462 | 29 | Even | 1804 | 2251 | 15.9643 | 2 | 7 | None | 2000 | 1712 | 590 | Tungsten | E |  | F | 2014-09-13 | TV - Nom
I5 | A | FP16 | 4303463 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 590 | Tungsten | E |  | G | 2014-09-13 | TV - Nom
I5 | A | FP16 | 4303464 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 590 | Tungsten | E |  | H | 2014-09-13 | TV - Nom
I5 | A | FP16 | 4303465 | 47 | Even | 2900 | 5476 | 56 | 3 | 12 | None | 2000 | 1712 | 590 | Ceramic | A |  | I | 2014-09-14 | TV - Nom
I5 | A | FP16 | 4303466 | 29 | Even | 5400 | 7080 | 60 | 3 | 10 | None | 2000 | 1712 | 590 | Ceramic | A |  | J | 2014-09-14 | TV - Nom
I5 | A | FP16 | 4303467 | 25 | Even | 7030.2 | 9909.9 | 119.9875 | 4 | 11 | None | 2000 | 1712 | 590 | Ceramic | A |  | K | 2014-09-14 | TV - Nom
I5 | A | FP16 | 4303468 | 68 | Even | 9756.8 | 16121.9 | 95.0015 | 4 | 16 | None | 2000 | 1712 | 590 | Ceramic | A |  | L | 2014-09-14 | TV - Nom
M1 | A | FP15 | 4302491 | 36 | All | 375 | 445 | 2 | 1 | 5 | None | 2000 | 586 | 770 | Tungsten | A |  | B | 2014-07-01 | Pre-TV
M1 | A | FP16 | 4302497 | 19 | All | 375 | 447 | 4 | 1 | 5 | 0.5 | 2000 | 1712 | 830 | Tungsten | A |  | A | 2014-07-01 | Pre-TV
M1 | A | FP16 | 4302498 | 12 | All | 343 | 370.5 | 2.5 | 1 | 5 | None | 2000 | 1712 | 830 | Tungsten | A |  | A | 2014-07-01 | Pre-TV
M1 | A | FP16 | 4302499 | 19 | All | 375 | 447 | 4 | 1 | 5 | None | 2000 | 1712 | 830 | Tungsten | A |  | A | 2014-07-01 | Pre-TV
M1 | A | FP16 | 4302500 | 34 | All | 452.5 | 634 | 5.5 | 1 | 5 | None | 2000 | 1712 | 830 | Tungsten | A |  | A | 2014-07-01 | Pre-TV
M1 | A | FP16 | 4202501 | 91 | All | 606.5 | 1101.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 830 | Tungsten | A |  | C | 2014-07-01 | Pre-TV
M2 | B | FP15 | 4302545 | 47 | All | 416 | 475.8 | 1.3 | 1 | 5 | None | 2000 | 402 | 471 | Tungsten | B |  | B | 2014-07-08 | Pre-TV
M2 | B | FP16 | 4302553 | 18 | All | 416 | 477.2 | 3.6 | 1 | 5 | 0.7 | 2000 | 1712 | 834 | Tungsten | B |  | A | 2014-07-08 | Pre-TV
M2 | B | FP16 | 4302554 | 14 | All | 344 | 415.5 | 5.5 | 1 | 5 | None | 2000 | 1712 | 834 | Tungsten | B |  | A | 2014-07-08 | Pre-TV
M2 | B | FP16 | 4302555 | 18 | All | 416 | 477.2 | 3.6 | 1 | 5 | None | 2000 | 1712 | 834 | Tungsten | B |  | A | 2014-07-08 | Pre-TV
M2 | B | FP16 | 4302556 | 29 | All | 482 | 636 | 5.5 | 1 | 5 | None | 2000 | 1712 | 834 | Tungsten | B |  | A | 2014-07-08 | Pre-TV
M2 | B | FP16 | 4302557 | 91 | All | 608.5 | 1103.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 834 | Tungsten | B |  | C | 2014-07-08 | Pre-TV
M3 | B | FP15 | 4302559 | 58 | All | 454 | 522.4 | 1.2 | 1 | 5 | None | 2000 | 260 | 304 | Tungsten | B |  | B | 2014-07-08 | Pre-TV
M3 | B | FP16 | 4302566 | 18 | All | 454 | 522 | 4 | 1 | 5 | 1.1 | 2000 | 1712 | 836 | Tungsten | B |  | A | 2014-07-09 | Pre-TV
M3 | B | FP16 | 4302567 | 19 | All | 350 | 449 | 5.5 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | B |  | A | 2014-07-09 | Pre-TV
M3 | B | FP16 | 4302568 | 18 | All | 454 | 522 | 4 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | B |  | A | 2014-07-09 | Pre-TV
M3 | B | FP16 | 4302569 | 20 | All | 527.5 | 632 | 5.5 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | B |  | A | 2014-07-09 | Pre-TV
M3 | B | FP16 | 4302570 | 91 | All | 604.5 | 1099.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 836 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M4 | A | FP15 | 4302503 | 77 | All | 522 | 590.4 | 0.9 | 1 | 5 | None | 2000 | 220 | 256 | Tungsten | A |  | B | 2014-07-02 | Pre-TV
M4 | B | FP15 | 4302866 | 77 | All | 522 | 590.4 | 0.9 | 1 | 5 | None | 2000 | 290 | 338 | Tungsten | D |  | B | 2014-08-14 | Pre-TV |  | Retest of FP15 with bulb D after lock-in amplifier fix
M4 | A | FP16 | 4302507 | 18 | All | 522 | 590 | 4 | 1 | 5 | 1.2 | 2000 | 1712 | 836 | Tungsten | A |  | A | 2014-07-02 | Pre-TV
M4 | A | FP16 | 4302508 | 31 | All | 354 | 519 | 5.5 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | A |  | A | 2014-07-02 | Pre-TV
M4 | A | FP16 | 4302509 | 18 | All | 522 | 590 | 4 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | A |  | A | 2014-07-02 | Pre-TV
M4 | A | FP16 | 4302510 | 8 | All | 595.5 | 634 | 5.5 | 1 | 5 | None | 2000 | 1712 | 836 | Tungsten | A |  | A | 2014-07-02 | Pre-TV
M4 | A | FP16 | 4302511 | 91 | All | 606.5 | 1101.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 836 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M5 | B | FP15 | 4302572 | 79 | All | 637 | 707.2 | 0.9 | 1 | 2 | None | 2000 | 202 | 237 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M5 | B | FP15 | 4304460 | 79 | All | 637 | 707.2 | 0.9 | 1 | 2 | None | 2000 | 214 | 248 | Tungsten | F |  | C | 2014-10-30 | TV - Hot |  | Retest of FP15 with bulb F after lock-in amplifier fix
M5 | B | FP16 | 4302580 | 19 | All | 637 | 709 | 4 | 1 | 2 | 1.3 | 2000 | 1712 | 852 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M5 | B | FP16 | 4302581 | 53 | All | 349 | 635 | 5.5 | 1 | 5 | None | 2000 | 1712 | 852 | Tungsten | B |  | A | 2014-07-09 | Pre-TV
M5 | B | FP16 | 4302582 | 6 | All | 607.5 | 635 | 5.5 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M5 | B | FP16 | 4302583 | 19 | All | 637 | 709 | 4 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M5 | B | FP16 | 4302584 | 71 | All | 714.5 | 1099.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | B |  | C | 2014-07-09 | Pre-TV
M6 | A | FP15 | 4302513 | 66 | All | 720 | 772 | 0.8 | 1 | 2 | None | 2000 | 166 | 195 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M6 | A | FP16 | 4302517 | 19 | All | 720 | 774 | 3 | 1 | 2 | 1.4 | 2000 | 1712 | 852 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M6 | A | FP16 | 4302518 | 53 | All | 349 | 635 | 5.5 | 1 | 5 | None | 2000 | 1712 | 852 | Tungsten | A |  | A | 2014-07-02 | Pre-TV
M6 | A | FP16 | 4302519 | 21 | All | 607.5 | 717.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M6 | A | FP16 | 4302520 | 19 | All | 720 | 774 | 3 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M6 | A | FP16 | 4302521 | 60 | All | 779.5 | 1104 | 5.5 | 1 | 2 | None | 2000 | 1712 | 852 | Tungsten | A |  | C | 2014-07-02 | Pre-TV
M7 | A | FP15 | 4302619 | 104 | All | 799 | 931.6 | 1.2874 | 1 | 2 | None | 2000 | 372 | 436 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
M7 | A | FP16 | 4302627 | 23 | All | 799 | 931 | 6 | 1 | 2 | 0.8 | 2000 | 1712 | 734 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
M7 | A | FP16 | 4302628 | 52 | All | 351 | 631.5 | 5.5 | 1 | 5 | None | 2000 | 1712 | 734 | Tungsten | D |  | A | 2014-07-12 | Pre-TV
M7 | A | FP16 | 4302629 | 36 | All | 604 | 796.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 734 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
M7 | A | FP16 | 4302630 | 23 | All | 799 | 931 | 6 | 1 | 2 | None | 2000 | 1712 | 734 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
M7 | A | FP16 | 4302631 | 31 | All | 936.5 | 1101.5 | 5.5 | 1 | 2 | None | 2000 | 1712 | 734 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
M8 | B | FP15 | 4303579 | 104 | Odd | 1204 | 1276.1 | 0.7 | 2 | 7 | None | 2000 | 86 | 100 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303583 | 19 | Odd | 1204 | 1276 | 4 | 2 | 7 | 1.9 | 2000 | 1500 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom |  | Slit configuration does not match OOB sweeps
M8 | B | FP16 | 4303584 | 13 | Odd | 996 | 1188 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303585 | 19 | Odd | 1204 | 1276 | 4 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303586 | 61 | Odd | 1284 | 2244 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303587 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303588 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303589 | 42 | Odd | 2788 | 5453 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303590 | 33 | Odd | 5128 | 7208 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-19 | TV - Nom
M8 | B | FP15 | 4303606 | 104 | Even | 1204 | 1276.1 | 0.7 | 2 | 7 | None | 2000 | 90 | 106 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303610 | 19 | Even | 1204 | 1276 | 4 | 2 | 7 | 1.9 | 2000 | 1650 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom |  | Slit configuration does not match OOB sweeps
M8 | B | FP16 | 4303611 | 13 | Even | 996 | 1188 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4304344 | 13 | Even | 996 | 1188 | 16 | 2 | 2 | None | 2000 | 1712 | 1180 | Tungsten | F |  | N | 2014-10-22 | TV - Hot
M8 | B | FP16 | 4303612 | 19 | Even | 1204 | 1276 | 4 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4304345 | 19 | Even | 1204 | 1276 | 4 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-10-22 | TV - Hot
M8 | B | FP16 | 4303613 | 61 | Even | 1284 | 2244 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303614 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303615 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303616 | 42 | Even | 2788 | 5453 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-19 | TV - Nom
M8 | B | FP16 | 4303617 | 33 | Even | 5128 | 7208 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-19 | TV - Nom
M9 | A | FP15 | 4303370 | 140 | Odd | 1350 | 1405 | 0.39568 | 2 | 7 | none | 2000 | 358 | 420 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP15 | 4303482 | 140 | Odd | 1350 | 1405.6 | 0.4 | 2 | 7 | none | 2000 | 54 | 65 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303373 | 14 | Odd | 1350 | 1408.5 | 4.5 | 2 | 7 | 0.8 | 2000 | 1712 | 1180 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303485 | 14 | Odd | 1350 | 1408.5 | 4.5 | 2 | 7 | 2.3 | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303374 | 24 | Odd | 990 | 1335 | 15 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303486 | 24 | Odd | 990 | 1335 | 15 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4304347 | 24 | Odd | 990 | 1335 | 15 | 2 | 2 | none | 2000 | 1712 | 1180 | Tungsten | F |  | N | 2014-10-22 | TV - Hot |  | Repeat of 4303486 with order filter N
M9 | A | FP16 | 4304348 | 24 | Odd | 990 | 1335 | 15 | 2 | 2 | none | 2000 | 1712 | 1180 | Tungsten | F |  | N | 2014-10-22 | TV - Hot |  | Repeat of 4304347
M9 | A | FP16 | 4303375 | 14 | Odd | 1350 | 1408.5 | 4.5 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303487 | 14 | Odd | 1350 | 1408.5 | 4.5 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4304349 | 14 | Odd | 1350 | 1408.5 | 4.5 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-10-22 | TV - Hot |  | Repeat of 4303487
M9 | A | FP16 | 4303376 | 60 | Odd | 1420 | 2246 | 14 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | D |  | F | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303488 | 60 | Odd | 1420 | 2246 | 14 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303377 | 16 | Odd | 2200 | 2410 | 14 | 2 | 8 | none | 2000 | 1712 | 1180 | Tungsten | D |  | G | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303489 | 16 | Odd | 2200 | 2410 | 14 | 2 | 8 | none | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303378 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | none | 2000 | 1712 | 1180 | Tungsten | D |  | H | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303490 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | none | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303379 | 42 | Odd | 2779 | 5444 | 65 | 3 | 12 | none | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303491 | 42 | Odd | 2779 | 5444 | 65 | 3 | 12 | none | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-14 | TV - Nom
M9 | A | FP16 | 4303380 | 33 | Odd | 5118.9 | 7199 | 65.0031 | 3 | 10 | none | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-08 | TV - Nom |  | Bad alingment
M9 | A | FP16 | 4303492 | 33 | Odd | 5118.9 | 7199 | 65.0031 | 3 | 10 | none | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-14 | TV - Nom
M9 | A | FP15 | 4303422 | 140 | Even | 1350 | 1405.6 | 0.4 | 2 | 7 | none | 2000 | 64 | 65 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303425 | 14 | Even | 1350 | 1408.5 | 4.5 | 2 | 7 | 2.4 | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303426 | 24 | Even | 990 | 1335 | 15 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303427 | 14 | Even | 1350 | 1408.5 | 4.5 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303428 | 60 | Even | 1420 | 2246 | 14 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303429 | 16 | Even | 2200 | 2410 | 14 | 2 | 8 | none | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303430 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | none | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303431 | 42 | Even | 2779 | 5444 | 65 | 3 | 12 | none | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-12 | TV - Nom
M9 | A | FP16 | 4303432 | 33 | Even | 5118.9 | 7199 | 65.0031 | 3 | 10 | none | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-12 | TV - Nom
M10 | B | FP15 | 4303592 | 211 | Odd | 1506 | 1715 | 0.99524 | 2 | 7 | none | 2000 | 118 | 140 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M10 | B | FP16 | 4303597 | 19 | Odd | 1506 | 1722 | 12 | 2 | 7 | 1.6 | 2000 | 1550 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom |  | Slit configuration does not match OOB sweeps
M10 | B | FP16 | 4303598 | 42 | Odd | 1002 | 1494 | 12 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M10 | B | FP16 | 4303599 | 19 | Odd | 1506 | 1722 | 12 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M10 | B | FP16 | 4303600 | 43 | Odd | 1734 | 2238 | 12 | 2 | 7 | none | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-19 | TV - Nom
M10 | B | FP16 | 4303601 | 19 | Odd | 2200 | 2416 | 12 | 2 | 8 | none | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-19 | TV - Nom
M10 | B | FP16 | 4303602 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | none | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303603 | 42 | Odd | 2806 | 5471 | 65 | 3 | 12 | none | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303604 | 31 | Odd | 5146 | 7096.1 | 65.0033 | 3 | 10 | none | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-20 | TV - Nom
M10 | B | FP15 | 4303619 | 211 | Even | 1506 | 1716 | 1 | 2 | 7 | None | 2000 | 116 | 136 | Tungsten | E |  | F | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303623 | 19 | Even | 1506 | 1722 | 12 | 2 | 7 | 1.7 | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303624 | 42 | Even | 1002 | 1494 | 12 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-20 | TV - Nom |  | Repeat of 4303624 with order filter N
M10 | B | FP16 | 4304351 | 42 | Even | 1002 | 1494 | 12 | 2 | 2 | None | 2000 | 1712 | 1180 | Tungsten | F |  | N | 2014-10-22 | TV-Hot
M10 | B | FP16 | 4303625 | 19 | Even | 1506 | 1722 | 12 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4304552 | 19 | Even | 1506 | 1722 | 12 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-10-22 | TV-Hot |  | Repeat of 4303625
M10 | B | FP16 | 4303626 | 43 | Even | 1734 | 2238 | 12 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303627 | 19 | Even | 2200 | 2416 | 12 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303628 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303629 | 42 | Even | 2806 | 5471 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303630 | 31 | Even | 5146 | 7096.1 | 65.0033 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-20 | TV - Nom
M10 | B | FP16 | 4303631 | 31 | Even | 5146 | 7096.1 | 65.0033 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-20 | TV - Nom |  | Repeat of 4303630
M11 | A | FP15 | 4303434 | 136 | Odd | 2164 | 2339.5 | 1.3 | 2 | 8 | None | 2000 | 170 | 199 | Tungsten | E |  | G | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303438 | 13 | Odd | 2164 | 2344 | 15 | 2 | 8 | 1.2 | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303913 | 13 | Odd | 2164 | 2344 | 15 | 2 | 8 | 1.3 | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4303439 | 78 | Odd | 994 | 2149 | 15 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303914 | 78 | Odd | 994 | 2149 | 15 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4304354 | 78 | Odd | 994 | 2149 | 15 | 2 | 2 | None | 2000 | 1712 | 1180 | Tungsten | F |  | N | 2014-10-22 | TV - Hot |  | Repeat of 4303914 with order filter N
M11 | A | FP16 | 4303440 | 13 | Odd | 2164 | 2344 | 15 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303915 | 13 | Odd | 2164 | 2344 | 15 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4304355 | 13 | Odd | 2164 | 2344 | 15 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-10-22 | TV - Hot |  | Repeat of 4303915
M11 | A | FP16 | 4303441 | 5 | Odd | 2360 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303916 | 5 | Odd | 2360 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4303442 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303917 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4303443 | 42 | Odd | 2781 | 5446.1 | 65.0024 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-12 | TV - Nom
M11 | A | FP16 | 4303918 | 42 | Odd | 2781 | 5446.1 | 65.0024 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP16 | 4303444 | 33 | Odd | 5121 | 7201.1 | 65.0031 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-13 | TV - Nom
M11 | A | FP16 | 4303919 | 33 | Odd | 5121 | 7201.1 | 65.0031 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-10-03 | TV - Nom |  | Retest of FP-16 with different ND filter
M11 | A | FP15 | 4303470 | 136 | Even | 2164 | 2339.5 | 1.3 | 2 | 8 | None | 2000 | 180 | 210 | Tungsten | E |  | G | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303474 | 13 | Even | 2164 | 2344 | 15 | 2 | 8 | 1.2 | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303475 | 78 | Even | 994 | 2149 | 15 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | E |  | F | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303476 | 13 | Even | 2164 | 2344 | 15 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303477 | 5 | Even | 2360 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | E |  | G | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303478 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | E |  | H | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303479 | 42 | Even | 2781 | 5446.1 | 65.0024 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-14 | TV - Nom
M11 | A | FP16 | 4303480 | 33 | Even | 5121 | 7201.1 | 65.0031 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-14 | TV - Nom
M12 | B | FP15 | 4303653 | 94 | Odd | 3401 | 4014.9 | 6.6011 | 3 | 12 | None | 2000 | 146 | 171 | Ceramic | A |  | I | 2014-09-22 | TV - Nom
M12 | B | FP16 | 4303657 | 13 | Odd | 3401 | 4048.9 | 53.9917 | 3 | 12 | 1.6 | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303658 | 13 | Odd | 3401 | 4048.9 | 53.9917 | 3 | 12 | 1.6 | 2000 | 1980 | 1180 | Ceramic | A |  | I | 2014-09-23 | TV - Nom |  | Slit configuration does not match OOB sweeps
M12 | B | FP16 | 4303659 | 79 | Odd | 1000 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303660 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303661 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303662 | 9 | Odd | 2914.9 | 3349.9 | 54.375 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303663 | 13 | Odd | 3401 | 4048.9 | 53.9917 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303664 | 27 | Odd | 4103.1 | 5506.9 | 53.9923 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-23 | TV - Nom
M12 | B | FP16 | 4303665 | 31 | Odd | 5237.1 | 7187 | 64.9967 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-23 | TV - Nom
M12 | B | FP15 | 4303680 | 94 | Even | 3401 | 4014.9 | 6.6011 | 3 | 12 | None | 2000 | 138 | 162 | Ceramic | A |  | I | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303683 | 13 | Even | 3401 | 4048.9 | 53.9917 | 3 | 12 | 1.6 | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303684 | 79 | Even | 1000 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303685 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303686 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303687 | 9 | Even | 2914.9 | 3349.9 | 54.375 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303688 | 13 | Even | 3401 | 4048.9 | 53.9917 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303689 | 27 | Even | 4103.1 | 5506.9 | 53.9923 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-24 | TV - Nom
M12 | B | FP16 | 4303690 | 31 | Even | 5237.1 | 7187 | 64.9967 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-24 | TV - Nom
M13 | A | FP15 | 4303866 | 73 | Odd | 3782 | 4329.2 | 7.6 | 3 | 12 | None | 2000 | 150 | 175 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303872 | 19 | Odd | 3782 | 4340 | 31 | 3 | 12 | 1.3 | 2000 | 1000 | 1180 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303873 | 79 | Odd | 1000 | 2248 | 16 | 2 | 7 | None | 2000 | 1000 | 1180 | Tungsten | F |  | F | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303874 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1000 | 1180 | Tungsten | F |  | G | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303875 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1000 | 1180 | Tungsten | F |  | H | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303876 | 15 | Odd | 2851.9 | 3719.9 | 62 | 3 | 12 | None | 2000 | 1000 | 1180 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303877 | 19 | Odd | 3782 | 4340 | 31 | 3 | 12 | None | 2000 | 1000 | 1180 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303878 | 18 | Odd | 4402 | 5455.9 | 61.9941 | 3 | 12 | None | 2000 | 1000 | 1180 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M13 | A | FP16 | 4303879 | 31 | Odd | 5146 | 7096.1 | 65.0033 | 3 | 10 | None | 2000 | 1000 | 1180 | Ceramic | A |  | J | 2014-10-01 | TV - Nom
M13 | A | FP15 | 4303882 | 73 | Even | 3782 | 4329.2 | 7.6 | 3 | 12 | None | 2000 | 140 | 162 | Ceramic | A |  | I | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303888 | 19 | Even | 3782 | 4340 | 31 | 3 | 12 | 1.3 | 2000 | 1100 | 1180 | Ceramic | A |  | I | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303889 | 79 | Even | 1000 | 2248 | 16 | 2 | 7 | None | 2000 | 1100 | 1180 | Tungsten | F |  | F | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303890 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1100 | 1180 | Tungsten | F |  | G | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303891 | 34 | Even | 2370 | 2882 | 15.5152 | 2 | 4 | None | 2000 | 1100 | 1180 | Tungsten | F |  | H | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303892 | 2 | Even | 2914 | 2930 | 16 | 2 | 4 | None | 2000 | 1100 | 1180 | Tungsten | F |  | H | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303893 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1100 | 1180 | Tungsten | F |  | H | 2014-10-02 | TV - Nom |  | Repeat of 4303891 and 4303892
M13 | A | FP16 | 4303894 | 15 | Even | 2851.9 | 3719.9 | 62 | 3 | 12 | None | 2000 | 1100 | 1180 | Ceramic | A |  | I | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303895 | 19 | Even | 3782 | 4340 | 31 | 3 | 12 | None | 2000 | 1100 | 1180 | Ceramic | A |  | I | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303896 | 18 | Even | 4402 | 5455.9 | 61.9941 | 3 | 12 | None | 2000 | 1100 | 1180 | Ceramic | A |  | I | 2014-10-02 | TV - Nom
M13 | A | FP16 | 4303897 | 31 | Even | 5146 | 7096.1 | 65.0033 | 3 | 10 | None | 2000 | 1100 | 1180 | Ceramic | A |  | J | 2014-10-02 | TV - Nom
M14 | B | FP15 | 4303709 | 54 | Odd | 8035 | 9095.2 | 20.0038 | 4 | 11 | None | 2000 | 414 | 483 | Ceramic | A |  | K | 2014-09-25 | TV - Nom
M14 | B | FP16 | 4303713 | 19 | Odd | 8035 | 9115.1 | 60.0056 | 4 | 11 | 0.9 | 2000 | 1900 | 1180 | Ceramic | A |  | K | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303714 | 29 | Odd | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1900 | 1180 | Tungsten | F |  | F | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303715 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1900 | 1180 | Tungsten | F |  | G | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303716 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1900 | 1180 | Tungsten | F |  | H | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303717 | 40 | Odd | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1900 | 1180 | Ceramic | A |  | I | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303718 | 27 | Odd | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1900 | 1180 | Ceramic | A |  | J | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303719 | 9 | Odd | 6955.1 | 7914.9 | 119.975 | 4 | 11 | None | 2000 | 1900 | 1180 | Ceramic | A |  | K | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303720 | 19 | Odd | 8035 | 9115.1 | 60.0056 | 4 | 11 | None | 2000 | 1900 | 1180 | Ceramic | A |  | K | 2014-09-26 | TV - Nom
M14 | B | FP16 | 4303721 | 58 | Odd | 9175.1 | 16015 | 119.9982 | 4 | 16 | None | 2000 | 1900 | 1180 | Ceramic | A |  | L | 2014-09-26 | TV - Nom
M14 | B | FP15 | 4303754 | 54 | Even | 8035 | 9095.2 | 20.0038 | 4 | 11 | None | 2000 | 454 | 531 | Ceramic | A |  | K | 2014-09-27 | TV - Nom
M14 | B | FP16 | 4303759 | 19 | Even | 8035 | 9115.1 | 60.0056 | 4 | 11 | 0.9 | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-27 | TV - Nom
M14 | B | FP16 | 4303760 | 19 | Even | 8035 | 9115.1 | 60.0056 | 4 | 11 | 0.6 | 2000 | 1200 | 1180 | Ceramic | A |  | K | 2014-09-27 | TV - Nom |  | Slit configuration does not match OOB sweeps
M14 | B | FP16 | 4303761 | 29 | Even | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-27 | TV - Nom
M14 | B | FP16 | 4303762 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-27 | TV - Nom
M14 | B | FP16 | 4303763 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-27 | TV - Nom
M14 | B | FP16 | 4303764 | 40 | Even | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-28 | TV - Nom
M14 | B | FP16 | 4303765 | 27 | Even | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-28 | TV - Nom
M14 | B | FP16 | 4303766 | 9 | Even | 6955.1 | 7914.9 | 119.975 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-28 | TV - Nom
M14 | B | FP16 | 4303767 | 19 | Even | 8035 | 9115.1 | 60.0056 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-28 | TV - Nom
M14 | B | FP16 | 4303768 | 58 | Even | 9175.1 | 16015 | 119.9982 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-28 | TV - Nom
M15 | A | FP15 | 4303802 | 46 | Odd | 9666.9 | 11917.2 | 50.0067 | 4 | 16 | None | 2000 | 1506 | 1761 | Ceramic | A |  | L | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303805 | 29 | Odd | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303806 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303807 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303808 | 40 | Odd | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303809 | 27 | Odd | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303810 | 22 | Odd | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303811 | 18 | Odd | 9666.9 | 11876.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-29 | TV - Nom
M15 | A | FP16 | 4303812 | 33 | Odd | 12006.9 | 16166.8 | 129.9969 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-29 | TV - Nom
M15 | A | FP15 | 4303850 | 46 | Even | 9666.9 | 11917.2 | 50.0067 | 4 | 16 | None | 2000 | 1452 | 1698 | Ceramic | A |  | L | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303857 | 29 | Even | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303858 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303859 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303860 | 40 | Even | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303861 | 27 | Even | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303862 | 22 | Even | 7008.1 | 9738 | 129.9952 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303863 | 18 | Even | 9666.9 | 11876.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-10-01 | TV - Nom
M15 | A | FP16 | 4303864 | 33 | Even | 12006.9 | 16166.8 | 129.9969 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-10-01 | TV - Nom
M16A | B | FP15 | 4303739 | 33 | Odd | 11026.8 | 13106.9 | 65.0031 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303743 | 29 | Odd | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303744 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303745 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303746 | 40 | Odd | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303747 | 27 | Odd | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303748 | 22 | Odd | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303749 | 22 | Odd | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-27 | TV - Nom |  | Repeat of 4303748
M16A | B | FP16 | 4303750 | 11 | Odd | 9596.9 | 10896.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303751 | 17 | Odd | 11026.8 | 13106.9 | 130.0062 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-27 | TV - Nom
M16A | B | FP16 | 4303752 | 23 | Odd | 13237.1 | 16097 | 129.9955 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-27 | TV - Nom
M16A | B | FP15 | 4303770 | 33 | Even | 11026.8 | 13106.9 | 65.0031 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303774 | 29 | Even | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303775 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303776 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303777 | 40 | Even | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303778 | 27 | Even | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303779 | 22 | Even | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303780 | 11 | Even | 9596.9 | 10896.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303781 | 17 | Even | 11026.8 | 13106.9 | 130.0062 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-28 | TV - Nom
M16A | B | FP16 | 4303782 | 23 | Even | 13237.1 | 16097 | 129.9955 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-28 | TV - Nom
M16B | A | FP15 | 4303814 | 33 | Odd | 11026.8 | 13106.9 | 65.0031 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-29 | TV - Nom
M16B | A | FP16 | 4303818 | 29 | Odd | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-29 | TV - Nom
M16B | A | FP16 | 4303819 | 15 | Odd | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-29 | TV - Nom
M16B | A | FP16 | 4303820 | 36 | Odd | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-29 | TV - Nom
M16B | A | FP16 | 4303821 | 40 | Odd | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303822 | 27 | Odd | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303823 | 22 | Odd | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303824 | 11 | Odd | 9596.9 | 10896.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303825 | 17 | Odd | 11026.8 | 13106.9 | 130.0062 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303826 | 23 | Odd | 13237.1 | 16097 | 129.9955 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP15 | 4303828 | 33 | Even | 11026.8 | 13106.9 | 65.0031 | 4 | 16 | None | 2000 | 1712 | 2000 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303832 | 29 | Even | 1800 | 2248 | 16 | 2 | 7 | None | 2000 | 1712 | 1180 | Tungsten | F |  | F | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303833 | 15 | Even | 2200 | 2424 | 16 | 2 | 8 | None | 2000 | 1712 | 1180 | Tungsten | F |  | G | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303834 | 36 | Even | 2370 | 2930 | 16 | 2 | 4 | None | 2000 | 1712 | 1180 | Tungsten | F |  | H | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303835 | 40 | Even | 2900 | 5435 | 65 | 3 | 12 | None | 2000 | 1712 | 1180 | Ceramic | A |  | I | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303836 | 27 | Even | 5383 | 7073 | 65 | 3 | 10 | None | 2000 | 1712 | 1180 | Ceramic | A |  | J | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303837 | 22 | Even | 7008 | 9738 | 130 | 4 | 11 | None | 2000 | 1712 | 1180 | Ceramic | A |  | K | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303838 | 11 | Even | 9596.9 | 10896.9 | 130 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303839 | 17 | Even | 11026.8 | 13106.9 | 130.0062 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
M16B | A | FP16 | 4303840 | 23 | Even | 13237.1 | 16097 | 129.9955 | 4 | 16 | None | 2000 | 1712 | 1180 | Ceramic | A |  | L | 2014-09-30 | TV - Nom
DNBMGS | A | FP15 | 4302633 | 35 | All | 462 | 972 | 15 | 1 | 5 | None | 2000 | 100 | 118 | Tungsten | D |  | A | 2014-07-12 | Pre-TV
DNBMGS | A | FP15 | 4302634 | 35 | All | 462 | 972 | 15 | 1 | 5 | None | 2000 | 82 | 97 | Tungsten | D |  | A | 2014-07-12 | Pre-TV
DNBMGS | A | FP16 | 4302635 | 23 | All | 650 | 980 | 15 | 1 | 2 | None | 2000 | 82 | 97 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
DNBMGS | A | FP16 | 4302639 | 12 | All | 462 | 627 | 15 | 1 | 5 | 1.7 | 2000 | 1712 | 2000 | Tungsten | D |  | A | 2014-07-12 | Pre-TV
DNBMGS | A | FP16 | 4302640 | 29 | All | 552 | 972 | 15 | 1 | 2 | 1.7 | 2000 | 1712 | 2000 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
DNBMGS | A | FP16 | 4302641 | 20 | All | 342 | 627 | 15 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | D |  | A | 2014-07-12 | Pre-TV
DNBMGS | A | FP16 | 4302642 | 38 | All | 552 | 1107 | 15 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | D |  | C | 2014-07-12 | Pre-TV
DNBLGS | B | FP15 | 4302595 | 35 | All | 462 | 972 | 15 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | B |  | A | 2014-07-10 | Pre-TV
DNBLGS | B | FP16 | 4302596 | 23 | All | 650 | 980 | 15 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-10 | Pre-TV
DNBLGS | B | FP16 | 4302599 | 20 | All | 342 | 627 | 15 | 1 | 5 | None | 2000 | 1712 | 2000 | Tungsten | B |  | A | 2014-07-10 | Pre-TV
DNBLGS | B | FP16 | 4302600 | 38 | All | 552 | 1107 | 15 | 1 | 2 | None | 2000 | 1712 | 2000 | Tungsten | B |  | C | 2014-07-10 | Pre-TV

## RSO
SpMA Source | Bulb | A / C | RSO Run | RSO Rows | # of Repeats | Bulb Hours | Date | Comments
Tungsten | A | C |  | A, B, C, D, E, F, G, H | 3 |  | 2014-06-18 | Merlin lock-in amplifier problem that led to wavelength shifts in the RSO
Tungsten | A | A | 1 | A, B, C, D, E, F, G, H | 3 |  | 2014-06-21
Tungsten | A | A | 2 | A, B, C, D | 3 |  | 2014-07-02
Tungsten | B | A | 1 | A, B, C, D | 3 |  | 2014-07-07
Tungsten | B | A | 2 | A, C | 2 |  | 2014-07-09
Tungsten | B | A | 3 | A, B, C, D | 2 |  | 2014-07-10
Tungsten | D | A | 1 | A, B, C, D | 3 |  | 2014-07-11
Tungsten | D | A | 2 | A, B, C, D | 3 |  | 2014-07-13
Tungsten | D | A | 3 | A, B, C, D, E, F, G, H | 3 |  | 2014-08-02
Tungsten | E | A | 1 | D, F, G, H | 3 |  | 2014-09-09
Tungsten | E | A | 2 | D, F, G, H | 2 |  | 2014-09-12
Tungsten | E | A | 3 | D, F, G, H | 2 |  | 2014-09-19
Tungsten | E | A | 4 | D, F, G, H | 2 |  | 2014-09-20
Tungsten | F | A | 1 | D, F, G, H | 3 |  | 2014-09-21
Tungsten | F | A | 2 | F, G, H | 3 |  | 2014-10-13
Tungsten | F | A | 3 | F, G, N | 3 |  | 2014-10-21 | Added new row N from 1000-1500 nm using grating 2 and order filter 2
Ceramic | A | C |  | I, J, K, L, M | 3 |  | 2014-06-18 | Merlin lock-in amplifier problem that led to wavelength shifts in the RSO
Ceramic | A | A | 1 | I, J, K, L, M | 3 |  | 2014-06-21
Ceramic | A | A | 2 | I, J, K, L, M | 3 |  | 2014-08-09
Ceramic | A | A | 3 | I, J, K, L, M | 3 |  | 2014-10-03
Ceramic | A | A | 4 | I, J, K, L, M | 3 |  | 2014-10-28

## Band Median
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source.
3. The VisNIR bands and the DNB were measured duirng pre-TV testing under ambient conditions. | [7 empty] | Non-compliance with specification (all detectors)
3. The SMWIR and LWIR bands were measured duirng TV testing at the nominal iinstrument temperature plateau. | [7 empty] | Non-compliance with specification (some detectors)
5. Odd and even detector arrays were measured separately for the SWIR, MWIR, and LWIR bands.
6. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
7. All values listed are the band median.
Band | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
| Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
| Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
I1 | 642.5108 | 643.8713 | 640 | 6 | 78.7184 | 80 | 6 | 593.4889 | 692.9282 | 565 | 715 | 0.10963 | 0.5
I2 | 867.4666 | 867.5605 | 865 | 8 | 36.3794 | 39 | 5 | 841.6046 | 893.6353 | 802 | 928 | 0.13791 | 0.7
I3 | 1603.5394 | 1604.1971 | 1610 | 14 | 60.5752 | 60 | 9 | 1544.8572 | 1668.1555 | 1509 | 1709 | 0.4517 | 0.7
I4 | 3747.1569 | 3748.053 | 3740 | 40 | 387.1992 | 380 | 30 | 3473.8135 | 4013.9694 | 3340 | 4140 | 0.18272 | 0.5
I5 | 11484.5381 | 11481.9697 | 11450 | 125 | 1867.7285 | 1900 | 100 | 10167.4429 | 13086.5566 | 9900 | 12900 | 0.28489 | 0.4
M1 | 411.1671 | 411.2074 | 412 | 2 | 17.7062 | 20 | 2 | 395.2781 | 425.4519 | 376 | 444 | 0.25526 | 1
M2 | 444.8706 | 444.7615 | 445 | 3 | 17.0384 | 18 | 2 | 429.3997 | 457.9775 | 417 | 473 | 0.33062 | 1
M3 | 488.7797 | 489.0874 | 488 | 4 | 18.8728 | 20 | 3 | 473.2912 | 504.4037 | 455 | 521 | 0.3004 | 0.7
M4 | 556.2954 | 556.4501 | 555 | 4 | 18.3606 | 20 | 3 | 540.1233 | 573.583 | 523 | 589 | 0.25179 | 0.7
M5 | 667.3452 | 667.4207 | 672 | 5 | 19.2337 | 20 | 3 | 649.922 | 684.9951 | 638 | 706 | 0.27469 | 0.7
M6 | 746.0195 | 745.9355 | 746 | 2 | 13.5057 | 15 | 2 | 733.9673 | 757.9894 | 721 | 771 | 0.26673 | 0.8
M7 | 867.6665 | 867.7979 | 865 | 8 | 36.2125 | 39 | 5 | 842.9968 | 892.5548 | 801 | 929 | 0.14946 | 0.7
M8 | 1238.4045 | 1238.545 | 1240 | 5 | 26.1273 | 20 | 4 | 1214.0549 | 1264.9846 | 1205 | 1275 | 0.48832 | 0.8
M9 | 1375.4962 | 1375.4647 | 1378 | 4 | 13.6769 | 15 | 3 | 1362.1338 | 1389.6602 | 1351 | 1405 | 0.47494 | 1
M10 | 1603.871 | 1604.5633 | 1610 | 14 | 60.2169 | 60 | 9 | 1546.1192 | 1667.706 | 1509 | 1709 | 0.44404 | 0.7
M11 | 2258.5601 | 2259.1034 | 2250 | 13 | 52.0094 | 50 | 6 | 2209.7556 | 2314.2686 | 2167 | 2333 | 0.37239 | 1
M12 | 3697.9368 | 3699.3541 | 3700 | 32 | 195.3706 | 180 | 20 | 3517.861 | 3892.993 | 3410 | 3990 | 0.36545 | 1.1
M13 | 4070.2067 | 4070.2161 | 4050 | 34 | 154.5793 | 155 | 20 | 3908.3124 | 4215.2251 | 3790 | 4310 | 0.76215 | 1.3
M14 | 8580.0674 | 8583.5527 | 8550 | 70 | 339.2397 | 300 | 40 | 8336.7593 | 8877.5938 | 8050 | 9050 | 0.24005 | 0.9
M15 | 10731.208 | 10722.5889 | 10763 | 113 | 999.5352 | 1000 | 100 | 9919.8677 | 11636.3945 | 9700 | 11740 | 0.42036 | 0.4
M16A | 11882.8081 | 11877.0615 | 12013 | 88 | 914.9209 | 950 | 50 | 11102.9922 | 12692.6733 | 11060 | 13050 | 0.56629 | 0.4
M16B | 11884.1445 | 11876.168 | 12013 | 88 | 935.7363 | 950 | 50 | 11101.5312 | 12699.2109 | 11060 | 13050 | 0.5961 | 0.4
DNBMGS | 693.1572 | 704.4458 | 700 | 14 | 376.9548 | 400 | 20 | 489.9801 | 904.5822 | 470 | 960 | 0.033253 | 0.1
DNBLGS | 707.2167 | 716.3859 | 700 | 14 | 357.523 | 400 | 20 | 481.7885 | 904.6895 | 470 | 960 | 0.036848 | 0.1

## Band Average
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source.
3. The VisNIR bands and the DNB were measured duirng pre-TV testing under ambient conditions. | [7 empty] | Non-compliance with specification (all detectors)
3. The SMWIR and LWIR bands were measured duirng TV testing at the nominal iinstrument temperature plateau. | [7 empty] | Non-compliance with specification (some detectors)
5. Odd and even detector arrays were measured separately for the SWIR, MWIR, and LWIR bands.
6. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
7. All values listed are the band median.
Band | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
| Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
| Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
I1 | 642.4444 | 643.7801 | 640 | 6 | 78.6632 | 80 | 6 | 593.3232 | 692.8145 | 565 | 715 | 0.11028 | 0.5
I2 | 867.3384 | 867.4437 | 865 | 8 | 36.361 | 39 | 5 | 841.5573 | 893.5139 | 802 | 928 | 0.14415 | 0.7
I3 | 1603.1459 | 1603.7835 | 1610 | 8 | 60.4557 | 60 | 5 | 1544.535 | 1667.4492 | 1509 | 1709 | 0.45542 | 0.7
I4 | 3746.8157 | 3748.1336 | 3740 | 8 | 387.1465 | 380 | 5 | 3473.394 | 4013.3862 | 3340 | 4140 | 0.18413 | 0.5
I5 | 11481.7177 | 11482.1784 | 11450 | 8 | 1872.3682 | 1900 | 5 | 10162.8357 | 13078.3581 | 9900 | 12900 | 0.30474 | 0.4
M1 | 411.1176 | 411.1498 | 412 | 2 | 17.5334 | 20 | 2 | 395.2645 | 425.373 | 376 | 444 | 0.25982 | 1
M2 | 444.8196 | 444.7147 | 445 | 3 | 17.0495 | 18 | 2 | 429.3313 | 457.9292 | 417 | 473 | 0.33436 | 1
M3 | 488.7392 | 489.0475 | 488 | 4 | 18.879 | 20 | 3 | 473.243 | 504.4041 | 455 | 521 | 0.30182 | 0.7
M4 | 556.2233 | 556.3889 | 555 | 4 | 18.3539 | 20 | 3 | 540.0498 | 573.5044 | 523 | 589 | 0.25308 | 0.7
M5 | 667.2989 | 667.3746 | 672 | 5 | 19.1972 | 20 | 3 | 649.8758 | 684.9454 | 638 | 706 | 0.2768 | 0.7
M6 | 745.979 | 745.8973 | 746 | 2 | 13.5075 | 15 | 2 | 733.9343 | 757.9552 | 721 | 771 | 0.26688 | 0.8
M7 | 867.531 | 867.6694 | 865 | 8 | 36.1952 | 39 | 5 | 842.8955 | 892.4758 | 801 | 929 | 0.14336 | 0.7
M8 | 1238.3412 | 1238.4827 | 1240 | 5 | 26.1278 | 20 | 4 | 1214.0041 | 1264.9232 | 1205 | 1275 | 0.49592 | 0.8
M9 | 1375.3673 | 1375.368 | 1378 | 4 | 13.7239 | 15 | 3 | 1362.0414 | 1389.4563 | 1351 | 1405 | 0.48445 | 1
M10 | 1603.7141 | 1604.4465 | 1610 | 14 | 60.0038 | 60 | 9 | 1545.8583 | 1667.503 | 1509 | 1709 | 0.44395 | 0.7
M11 | 2258.19 | 2258.7969 | 2250 | 13 | 51.9918 | 50 | 6 | 2209.4298 | 2314.3241 | 2167 | 2333 | 0.37254 | 1
M12 | 3697.6745 | 3699.2493 | 3700 | 32 | 195.3219 | 180 | 20 | 3517.9997 | 3892.966 | 3410 | 3990 | 0.36626 | 1.1
M13 | 4069.7433 | 4069.699 | 4050 | 34 | 154.6367 | 155 | 20 | 3907.7097 | 4215.3819 | 3790 | 4310 | 0.77559 | 1.3
M14 | 8580.1075 | 8583.3827 | 8550 | 70 | 339.3076 | 300 | 40 | 8336.3815 | 8878.6135 | 8050 | 9050 | 0.24125 | 0.9
M15 | 10730.9038 | 10721.3343 | 10763 | 113 | 1000.3617 | 1000 | 100 | 9915.0123 | 11638.0508 | 9700 | 11740 | 0.42076 | 0.4
M16A | 11882.0111 | 11876.8437 | 12013 | 88 | 914.8541 | 950 | 50 | 11103.9011 | 12691.9383 | 11060 | 13050 | 0.58256 | 0.4
M16B | 11882.4419 | 11875.4802 | 12013 | 88 | 934.801 | 950 | 50 | 11101.0848 | 12698.0162 | 11060 | 13050 | 0.60225 | 0.4
DNBMGS | 694.2006 | 705.7937 | 700 | 14 | 375.6181 | 400 | 20 | 490.228 | 904.5798 | 470 | 960 | 0.033294 | 0.1
DNBLGS | 709.9395 | 716.7815 | 700 | 14 | 352.4864 | 400 | 20 | 481.4898 | 904.6764 | 470 | 960 | 0.036792 | 0.1

## DNB
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source. | [8 empty] | Non-compliance with specification
3. The DNB was measured duirng pre-TV testing under ambient conditions.
4. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
Band | Detector | Subsample | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
|  |  | Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
|  |  | Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
DNBMGS | 1 | 1 | 692.8925 | 704.4531 | 700 | 14 | 375.9581 | 400 | 20 | 489.8931 | 904.4888 | 470 | 960 | 0.033296 | 0.1
DNBMGS | 2 | 1 | 692.4819 | 703.2037 | 700 | 14 | 376.4191 | 400 | 20 | 489.1703 | 904.494 | 470 | 960 | 0.032469 | 0.1
DNBMGS | 3 | 1 | 692.2896 | 703.6826 | 700 | 14 | 376.3656 | 400 | 20 | 488.8826 | 904.5053 | 470 | 960 | 0.032046 | 0.1
DNBMGS | 4 | 1 | 692.4443 | 703.6312 | 700 | 14 | 375.5696 | 400 | 20 | 489.2053 | 904.4567 | 470 | 960 | 0.033749 | 0.1
DNBMGS | 5 | 1 | 691.5529 | 702.8278 | 700 | 14 | 374.2631 | 400 | 20 | 488.9355 | 904.5344 | 470 | 960 | 0.033028 | 0.1
DNBMGS | 6 | 1 | 692.8437 | 702.9418 | 700 | 14 | 377.5682 | 400 | 20 | 488.5811 | 904.556 | 470 | 960 | 0.03202 | 0.1
DNBMGS | 7 | 1 | 693.1293 | 703.6225 | 700 | 14 | 377.9241 | 400 | 20 | 488.6613 | 904.5785 | 470 | 960 | 0.031977 | 0.1
DNBMGS | 8 | 1 | 693.3317 | 703.9153 | 700 | 14 | 378.1889 | 400 | 20 | 489.0224 | 904.5938 | 470 | 960 | 0.031932 | 0.1
DNBMGS | 9 | 1 | 693.1851 | 705.4538 | 700 | 14 | 375.6346 | 400 | 20 | 490.1412 | 904.5258 | 470 | 960 | 0.034649 | 0.1
DNBMGS | 10 | 1 | 693.9468 | 706.5613 | 700 | 14 | 377.7307 | 400 | 20 | 490.0672 | 904.5859 | 470 | 960 | 0.033081 | 0.1
DNBMGS | 11 | 1 | 692.9156 | 704.4385 | 700 | 14 | 374.0957 | 400 | 20 | 491.2605 | 904.6173 | 470 | 960 | 0.034669 | 0.1
DNBMGS | 12 | 1 | 694.7727 | 707.8441 | 700 | 14 | 377.848 | 400 | 20 | 491.672 | 904.6508 | 470 | 960 | 0.033347 | 0.1
DNBMGS | 13 | 1 | 694.4007 | 707.5102 | 700 | 14 | 377.4904 | 400 | 20 | 491.9984 | 904.6531 | 470 | 960 | 0.033266 | 0.1
DNBMGS | 14 | 1 | 695.1274 | 709.1263 | 700 | 14 | 378.2517 | 400 | 20 | 492.0239 | 904.6678 | 470 | 960 | 0.03324 | 0.1
DNBMGS | 15 | 1 | 695.9843 | 711.0172 | 700 | 14 | 378.3223 | 400 | 20 | 492.0516 | 904.6917 | 470 | 960 | 0.034172 | 0.1
DNBMGS | 16 | 1 | 705.9121 | 712.4696 | 700 | 14 | 358.2586 | 400 | 20 | 492.0811 | 904.677 | 470 | 960 | 0.035769 | 0.1
DNBLGS | 1 | 1 | 695.7487 | 714.2003 | 700 | 14 | 377.8284 | 400 | 20 | 480.1755 | 904.6124 | 470 | 960 | 0.035267 | 0.1
DNBLGS | 2 | 1 | 705.6776 | 715.3679 | 700 | 14 | 357.7458 | 400 | 20 | 480.5462 | 904.6119 | 470 | 960 | 0.036803 | 0.1
DNBLGS | 3 | 1 | 704.4489 | 714.9005 | 700 | 14 | 360.5117 | 400 | 20 | 480.7723 | 904.6326 | 470 | 960 | 0.037342 | 0.1
DNBLGS | 4 | 1 | 704.2597 | 714.9466 | 700 | 14 | 363.4465 | 400 | 20 | 480.9977 | 904.6625 | 470 | 960 | 0.035602 | 0.1
DNBLGS | 5 | 1 | 705.91 | 715.7745 | 700 | 14 | 358.5427 | 400 | 20 | 481.317 | 904.6551 | 470 | 960 | 0.036893 | 0.1
DNBLGS | 6 | 1 | 696.7238 | 715.0206 | 700 | 14 | 379.6468 | 400 | 20 | 481.376 | 904.6887 | 470 | 960 | 0.035711 | 0.1
DNBLGS | 7 | 1 | 704.8663 | 715.2385 | 700 | 14 | 362.8638 | 400 | 20 | 481.5616 | 904.6906 | 470 | 960 | 0.035838 | 0.1
DNBLGS | 8 | 1 | 706.0926 | 716.2225 | 700 | 14 | 360.6437 | 400 | 20 | 481.7152 | 904.6814 | 470 | 960 | 0.036272 | 0.1
DNBLGS | 9 | 1 | 709.1245 | 716.5492 | 700 | 14 | 355.0679 | 400 | 20 | 481.9362 | 904.6986 | 470 | 960 | 0.03729 | 0.1
DNBLGS | 10 | 1 | 708.7416 | 717.0477 | 700 | 14 | 357.3002 | 400 | 20 | 481.8838 | 904.7067 | 470 | 960 | 0.036971 | 0.1
DNBLGS | 11 | 1 | 708.3408 | 717.0281 | 700 | 14 | 357.2783 | 400 | 20 | 481.8618 | 904.7071 | 470 | 960 | 0.036351 | 0.1
DNBLGS | 12 | 1 | 710.0593 | 717.9284 | 700 | 14 | 354.6613 | 400 | 20 | 481.8912 | 904.7132 | 470 | 960 | 0.036607 | 0.1
DNBLGS | 13 | 1 | 719.7759 | 717.8384 | 700 | 14 | 333.1762 | 400 | 20 | 481.9335 | 904.6903 | 470 | 960 | 0.037321 | 0.1
DNBLGS | 14 | 1 | 722.683 | 719.1696 | 700 | 14 | 328.6157 | 400 | 20 | 481.9759 | 904.7021 | 470 | 960 | 0.038434 | 0.1
DNBLGS | 15 | 1 | 725.7201 | 720.0388 | 700 | 14 | 320.5889 | 400 | 20 | 481.964 | 904.6712 | 470 | 960 | 0.037655 | 0.1
DNBLGS | 16 | 1 | 730.8596 | 721.2333 | 700 | 14 | 311.8649 | 400 | 20 | 481.9282 | 904.6987 | 470 | 960 | 0.038323 | 0.1

## VisNIR
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source. | [8 empty] | Non-compliance with specification
3. The VisNIR bands were measured duirng pre-TV testing under ambient conditions.
4. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
Band | Detector | Subsample | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
|  |  | Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
|  |  | Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
I1 | 1 | 1 | 641.7863 | 643.0822 | 640 | 6 | 79.0798 | 80 | 6 | 592.1928 | 692.2943 | 565 | 715 | 0.10953 | 0.5
I1 | 2 | 1 | 641.894 | 643.2017 | 640 | 6 | 79.1026 | 80 | 6 | 592.3796 | 692.5121 | 565 | 715 | 0.10889 | 0.5
I1 | 3 | 1 | 641.9971 | 643.3102 | 640 | 6 | 79.1253 | 80 | 6 | 592.4979 | 692.6413 | 565 | 715 | 0.10425 | 0.5
I1 | 4 | 1 | 642.0909 | 643.4108 | 640 | 6 | 79.1396 | 80 | 6 | 592.5596 | 692.908 | 565 | 715 | 0.11191 | 0.5
I1 | 5 | 1 | 642.1772 | 643.4988 | 640 | 6 | 79.1483 | 80 | 6 | 592.6924 | 693.0004 | 565 | 715 | 0.10384 | 0.5
I1 | 6 | 1 | 642.2562 | 643.579 | 640 | 6 | 79.1435 | 80 | 6 | 592.7979 | 693.1118 | 565 | 715 | 0.10698 | 0.5
I1 | 7 | 1 | 642.328 | 643.6519 | 640 | 6 | 79.1393 | 80 | 6 | 592.8842 | 693.1241 | 565 | 715 | 0.10727 | 0.5
I1 | 8 | 1 | 642.3939 | 643.7209 | 640 | 6 | 79.1163 | 80 | 6 | 593.0316 | 693.1795 | 565 | 715 | 0.1036 | 0.5
I1 | 9 | 1 | 642.4484 | 643.7752 | 640 | 6 | 79.0948 | 80 | 6 | 593.1994 | 693.1718 | 565 | 715 | 0.10579 | 0.5
I1 | 10 | 1 | 642.4972 | 643.8299 | 640 | 6 | 79.0632 | 80 | 6 | 593.2785 | 693.1947 | 565 | 715 | 0.10535 | 0.5
I1 | 11 | 1 | 642.5468 | 643.8685 | 640 | 6 | 79.0152 | 80 | 6 | 593.3432 | 693.2327 | 565 | 715 | 0.10673 | 0.5
I1 | 12 | 1 | 642.5903 | 643.9039 | 640 | 6 | 78.9508 | 80 | 6 | 593.4105 | 693.2582 | 565 | 715 | 0.10903 | 0.5
I1 | 13 | 1 | 642.6262 | 643.935 | 640 | 6 | 78.8849 | 80 | 6 | 593.6349 | 693.1507 | 565 | 715 | 0.10544 | 0.5
I1 | 14 | 1 | 642.6503 | 643.9562 | 640 | 6 | 78.8268 | 80 | 6 | 593.6187 | 693.1629 | 565 | 715 | 0.10781 | 0.5
I1 | 15 | 1 | 642.6664 | 643.9646 | 640 | 6 | 78.7775 | 80 | 6 | 593.6591 | 693.1414 | 565 | 715 | 0.10653 | 0.5
I1 | 16 | 1 | 642.6716 | 643.9632 | 640 | 6 | 78.7406 | 80 | 6 | 593.6243 | 693.1774 | 565 | 715 | 0.11114 | 0.5
I1 | 17 | 1 | 642.6688 | 643.9594 | 640 | 6 | 78.6964 | 80 | 6 | 593.7065 | 693.1209 | 565 | 715 | 0.11117 | 0.5
I1 | 18 | 1 | 642.6647 | 643.951 | 640 | 6 | 78.654 | 80 | 6 | 593.6954 | 693.0623 | 565 | 715 | 0.1129 | 0.5
I1 | 19 | 1 | 642.662 | 643.9473 | 640 | 6 | 78.6119 | 80 | 6 | 593.7169 | 693.0197 | 565 | 715 | 0.11338 | 0.5
I1 | 20 | 1 | 642.6488 | 643.9431 | 640 | 6 | 78.564 | 80 | 6 | 593.8192 | 692.877 | 565 | 715 | 0.10104 | 0.5
I1 | 21 | 1 | 642.6246 | 643.9267 | 640 | 6 | 78.5262 | 80 | 6 | 593.6062 | 692.9479 | 565 | 715 | 0.11257 | 0.5
I1 | 22 | 1 | 642.603 | 643.9059 | 640 | 6 | 78.4788 | 80 | 6 | 593.7001 | 692.7974 | 565 | 715 | 0.11255 | 0.5
I1 | 23 | 1 | 642.5904 | 643.903 | 640 | 6 | 78.4244 | 80 | 6 | 593.6325 | 692.8394 | 565 | 715 | 0.11555 | 0.5
I1 | 24 | 1 | 642.5826 | 643.9039 | 640 | 6 | 78.3593 | 80 | 6 | 593.6631 | 692.7208 | 565 | 715 | 0.1102 | 0.5
I1 | 25 | 1 | 642.5493 | 643.8792 | 640 | 6 | 78.333 | 80 | 6 | 593.6062 | 692.6434 | 565 | 715 | 0.10994 | 0.5
I1 | 26 | 1 | 642.5251 | 643.8745 | 640 | 6 | 78.2677 | 80 | 6 | 593.6093 | 692.5692 | 565 | 715 | 0.10966 | 0.5
I1 | 27 | 1 | 642.4957 | 643.8671 | 640 | 6 | 78.2236 | 80 | 6 | 593.4842 | 692.5847 | 565 | 715 | 0.11433 | 0.5
I1 | 28 | 1 | 642.4658 | 643.8696 | 640 | 6 | 78.1781 | 80 | 6 | 593.54 | 692.425 | 565 | 715 | 0.11727 | 0.5
I1 | 29 | 1 | 642.4427 | 643.8749 | 640 | 6 | 78.0611 | 80 | 6 | 593.5483 | 692.2639 | 565 | 715 | 0.1117 | 0.5
I1 | 30 | 1 | 642.4062 | 643.8583 | 640 | 6 | 77.9388 | 80 | 6 | 593.3762 | 692.2214 | 565 | 715 | 0.12036 | 0.5
I1 | 31 | 1 | 642.3618 | 643.8375 | 640 | 6 | 77.8284 | 80 | 6 | 593.4369 | 691.9447 | 565 | 715 | 0.12407 | 0.5
I1 | 32 | 1 | 642.3074 | 643.8107 | 640 | 6 | 77.7297 | 80 | 6 | 593.3582 | 691.7487 | 565 | 715 | 0.1195 | 0.5
I1 | 1 | 2 | 641.7863 | 643.0817 | 640 | 6 | 79.0799 | 80 | 6 | 592.1882 | 692.2948 | 565 | 715 | 0.1132 | 0.5
I1 | 2 | 2 | 641.8942 | 643.2022 | 640 | 6 | 79.1022 | 80 | 6 | 592.3873 | 692.5045 | 565 | 715 | 0.10334 | 0.5
I1 | 3 | 2 | 641.9969 | 643.3104 | 640 | 6 | 79.1257 | 80 | 6 | 592.4927 | 692.6459 | 565 | 715 | 0.10659 | 0.5
I1 | 4 | 2 | 642.091 | 643.4109 | 640 | 6 | 79.1395 | 80 | 6 | 592.5649 | 692.9103 | 565 | 715 | 0.10987 | 0.5
I1 | 5 | 2 | 642.1774 | 643.4991 | 640 | 6 | 79.1481 | 80 | 6 | 592.6949 | 692.9993 | 565 | 715 | 0.10301 | 0.5
I1 | 6 | 2 | 642.2562 | 643.5792 | 640 | 6 | 79.1434 | 80 | 6 | 592.7928 | 693.1147 | 565 | 715 | 0.10901 | 0.5
I1 | 7 | 2 | 642.328 | 643.6516 | 640 | 6 | 79.1393 | 80 | 6 | 592.8852 | 693.1246 | 565 | 715 | 0.10908 | 0.5
I1 | 8 | 2 | 642.394 | 643.7204 | 640 | 6 | 79.1163 | 80 | 6 | 593.03 | 693.18 | 565 | 715 | 0.10432 | 0.5
I1 | 9 | 2 | 642.4485 | 643.7757 | 640 | 6 | 79.0946 | 80 | 6 | 593.2052 | 693.1708 | 565 | 715 | 0.1073 | 0.5
I1 | 10 | 2 | 642.4972 | 643.83 | 640 | 6 | 79.0635 | 80 | 6 | 593.2771 | 693.1973 | 565 | 715 | 0.10593 | 0.5
I1 | 11 | 2 | 642.5469 | 643.8691 | 640 | 6 | 79.0149 | 80 | 6 | 593.3539 | 693.2345 | 565 | 715 | 0.10718 | 0.5
I1 | 12 | 2 | 642.5904 | 643.9033 | 640 | 6 | 78.9506 | 80 | 6 | 593.4108 | 693.2484 | 565 | 715 | 0.11301 | 0.5
I1 | 13 | 2 | 642.6262 | 643.9355 | 640 | 6 | 78.8851 | 80 | 6 | 593.6143 | 693.1694 | 565 | 715 | 0.10808 | 0.5
I1 | 14 | 2 | 642.6503 | 643.956 | 640 | 6 | 78.8268 | 80 | 6 | 593.62 | 693.162 | 565 | 715 | 0.1093 | 0.5
I1 | 15 | 2 | 642.6665 | 643.9649 | 640 | 6 | 78.7773 | 80 | 6 | 593.6562 | 693.1479 | 565 | 715 | 0.10756 | 0.5
I1 | 16 | 2 | 642.6716 | 643.963 | 640 | 6 | 78.7404 | 80 | 6 | 593.614 | 693.1721 | 565 | 715 | 0.11124 | 0.5
I1 | 17 | 2 | 642.6686 | 643.9595 | 640 | 6 | 78.6958 | 80 | 6 | 593.7499 | 693.1468 | 565 | 715 | 0.10808 | 0.5
I1 | 18 | 2 | 642.6652 | 643.9523 | 640 | 6 | 78.6531 | 80 | 6 | 593.7325 | 693.0406 | 565 | 715 | 0.10766 | 0.5
I1 | 19 | 2 | 642.6622 | 643.9489 | 640 | 6 | 78.6117 | 80 | 6 | 593.7469 | 693.0093 | 565 | 715 | 0.1149 | 0.5
I1 | 20 | 2 | 642.6488 | 643.9451 | 640 | 6 | 78.5649 | 80 | 6 | 593.7842 | 692.946 | 565 | 715 | 0.1034 | 0.5
I1 | 21 | 2 | 642.6246 | 643.9258 | 640 | 6 | 78.5254 | 80 | 6 | 593.6439 | 692.8997 | 565 | 715 | 0.1096 | 0.5
I1 | 22 | 2 | 642.6033 | 643.9066 | 640 | 6 | 78.4769 | 80 | 6 | 593.7986 | 692.7365 | 565 | 715 | 0.10516 | 0.5
I1 | 23 | 2 | 642.5903 | 643.9039 | 640 | 6 | 78.424 | 80 | 6 | 593.6303 | 692.8398 | 565 | 715 | 0.10999 | 0.5
I1 | 24 | 2 | 642.5828 | 643.903 | 640 | 6 | 78.3588 | 80 | 6 | 593.6788 | 692.7033 | 565 | 715 | 0.11062 | 0.5
I1 | 25 | 2 | 642.5492 | 643.8779 | 640 | 6 | 78.3337 | 80 | 6 | 593.6125 | 692.6551 | 565 | 715 | 0.1125 | 0.5
I1 | 26 | 2 | 642.5245 | 643.8722 | 640 | 6 | 78.2697 | 80 | 6 | 593.5845 | 692.5643 | 565 | 715 | 0.11443 | 0.5
I1 | 27 | 2 | 642.4954 | 643.8655 | 640 | 6 | 78.2246 | 80 | 6 | 593.4819 | 692.5719 | 565 | 715 | 0.1182 | 0.5
I1 | 28 | 2 | 642.4664 | 643.8704 | 640 | 6 | 78.1756 | 80 | 6 | 593.5962 | 692.3871 | 565 | 715 | 0.11273 | 0.5
I1 | 29 | 2 | 642.4424 | 643.8738 | 640 | 6 | 78.0616 | 80 | 6 | 593.4937 | 692.3055 | 565 | 715 | 0.11727 | 0.5
I1 | 30 | 2 | 642.4061 | 643.8556 | 640 | 6 | 77.939 | 80 | 6 | 593.347 | 692.2036 | 565 | 715 | 0.1118 | 0.5
I1 | 31 | 2 | 642.3625 | 643.8379 | 640 | 6 | 77.8265 | 80 | 6 | 593.4636 | 691.9448 | 565 | 715 | 0.12001 | 0.5
I1 | 32 | 2 | 642.3074 | 643.8102 | 640 | 6 | 77.7298 | 80 | 6 | 593.249 | 691.8494 | 565 | 715 | 0.12342 | 0.5
I2 | 1 | 1 | 866.5508 | 866.6556 | 865 | 8 | 36.2001 | 39 | 5 | 841.1268 | 892.5084 | 802 | 928 | 0.20084 | 0.7
I2 | 2 | 1 | 866.6663 | 866.7789 | 865 | 8 | 36.2121 | 39 | 5 | 841.1843 | 892.7579 | 802 | 928 | 0.18743 | 0.7
I2 | 3 | 1 | 866.7783 | 866.8959 | 865 | 8 | 36.2293 | 39 | 5 | 841.2387 | 892.9188 | 802 | 928 | 0.17833 | 0.7
I2 | 4 | 1 | 866.8813 | 867.0016 | 865 | 8 | 36.2465 | 39 | 5 | 841.2932 | 893.0597 | 802 | 928 | 0.17355 | 0.7
I2 | 5 | 1 | 866.9741 | 867.0963 | 865 | 8 | 36.2672 | 39 | 5 | 841.3446 | 893.1815 | 802 | 928 | 0.16875 | 0.7
I2 | 6 | 1 | 867.0601 | 867.1824 | 865 | 8 | 36.2855 | 39 | 5 | 841.3951 | 893.2937 | 802 | 928 | 0.15889 | 0.7
I2 | 7 | 1 | 867.1375 | 867.2596 | 865 | 8 | 36.3003 | 39 | 5 | 841.4397 | 893.369 | 802 | 928 | 0.15381 | 0.7
I2 | 8 | 1 | 867.2063 | 867.3265 | 865 | 8 | 36.3185 | 39 | 5 | 841.4936 | 893.4529 | 802 | 928 | 0.14057 | 0.7
I2 | 9 | 1 | 867.2668 | 867.3862 | 865 | 8 | 36.3318 | 39 | 5 | 841.5364 | 893.4894 | 802 | 928 | 0.14016 | 0.7
I2 | 10 | 1 | 867.318 | 867.434 | 865 | 8 | 36.3445 | 39 | 5 | 841.5572 | 893.5422 | 802 | 928 | 0.1451 | 0.7
I2 | 11 | 1 | 867.363 | 867.4778 | 865 | 8 | 36.3529 | 39 | 5 | 841.6009 | 893.5796 | 802 | 928 | 0.14526 | 0.7
I2 | 12 | 1 | 867.4053 | 867.5189 | 865 | 8 | 36.3669 | 39 | 5 | 841.6328 | 893.6193 | 802 | 928 | 0.13762 | 0.7
I2 | 13 | 1 | 867.4381 | 867.5494 | 865 | 8 | 36.3748 | 39 | 5 | 841.6601 | 893.6309 | 802 | 928 | 0.14565 | 0.7
I2 | 14 | 1 | 867.4733 | 867.581 | 865 | 8 | 36.3817 | 39 | 5 | 841.6935 | 893.6356 | 802 | 928 | 0.13659 | 0.7
I2 | 15 | 1 | 867.502 | 867.6066 | 865 | 8 | 36.3913 | 39 | 5 | 841.707 | 893.6723 | 802 | 928 | 0.13766 | 0.7
I2 | 16 | 1 | 867.5217 | 867.6239 | 865 | 8 | 36.402 | 39 | 5 | 841.7075 | 893.701 | 802 | 928 | 0.1338 | 0.7
I2 | 17 | 1 | 867.5374 | 867.6384 | 865 | 8 | 36.4078 | 39 | 5 | 841.7261 | 893.7065 | 802 | 928 | 0.13209 | 0.7
I2 | 18 | 1 | 867.5515 | 867.6519 | 865 | 8 | 36.4156 | 39 | 5 | 841.7296 | 893.7241 | 802 | 928 | 0.13149 | 0.7
I2 | 19 | 1 | 867.5604 | 867.6594 | 865 | 8 | 36.4175 | 39 | 5 | 841.7568 | 893.7047 | 802 | 928 | 0.11734 | 0.7
I2 | 20 | 1 | 867.5613 | 867.6589 | 865 | 8 | 36.4255 | 39 | 5 | 841.7105 | 893.7408 | 802 | 928 | 0.1327 | 0.7
I2 | 21 | 1 | 867.5578 | 867.6558 | 865 | 8 | 36.4252 | 39 | 5 | 841.7 | 893.7281 | 802 | 928 | 0.13484 | 0.7
I2 | 22 | 1 | 867.5558 | 867.6523 | 865 | 8 | 36.4289 | 39 | 5 | 841.6915 | 893.7259 | 802 | 928 | 0.13815 | 0.7
I2 | 23 | 1 | 867.5562 | 867.6539 | 865 | 8 | 36.4275 | 39 | 5 | 841.6959 | 893.7184 | 802 | 928 | 0.13592 | 0.7
I2 | 24 | 1 | 867.5527 | 867.6495 | 865 | 8 | 36.4322 | 39 | 5 | 841.6818 | 893.7244 | 802 | 928 | 0.1336 | 0.7
I2 | 25 | 1 | 867.5359 | 867.6337 | 865 | 8 | 36.4223 | 39 | 5 | 841.6667 | 893.6948 | 802 | 928 | 0.1384 | 0.7
I2 | 26 | 1 | 867.5206 | 867.6188 | 865 | 8 | 36.42 | 39 | 5 | 841.6455 | 893.6816 | 802 | 928 | 0.12762 | 0.7
I2 | 27 | 1 | 867.4985 | 867.5972 | 865 | 8 | 36.4168 | 39 | 5 | 841.6115 | 893.6627 | 802 | 928 | 0.12916 | 0.7
I2 | 28 | 1 | 867.4841 | 867.5833 | 865 | 8 | 36.4095 | 39 | 5 | 841.5861 | 893.6515 | 802 | 928 | 0.12396 | 0.7
I2 | 29 | 1 | 867.4712 | 867.5682 | 865 | 8 | 36.3892 | 39 | 5 | 841.5557 | 893.6351 | 802 | 928 | 0.13942 | 0.7
I2 | 30 | 1 | 867.4626 | 867.5529 | 865 | 8 | 36.3712 | 39 | 5 | 841.5312 | 893.5963 | 802 | 928 | 0.1408 | 0.7
I2 | 31 | 1 | 867.4564 | 867.5377 | 865 | 8 | 36.3563 | 39 | 5 | 841.5063 | 893.5701 | 802 | 928 | 0.15204 | 0.7
I2 | 32 | 1 | 867.4305 | 867.5068 | 865 | 8 | 36.3442 | 39 | 5 | 841.4658 | 893.515 | 802 | 928 | 0.14864 | 0.7
I2 | 1 | 2 | 866.5505 | 866.6565 | 865 | 8 | 36.2024 | 39 | 5 | 841.1276 | 892.5366 | 802 | 928 | 0.19705 | 0.7
I2 | 2 | 2 | 866.6658 | 866.7789 | 865 | 8 | 36.2136 | 39 | 5 | 841.1839 | 892.7827 | 802 | 928 | 0.18206 | 0.7
I2 | 3 | 2 | 866.7779 | 866.8966 | 865 | 8 | 36.2316 | 39 | 5 | 841.2383 | 892.9421 | 802 | 928 | 0.17312 | 0.7
I2 | 4 | 2 | 866.8811 | 867.0028 | 865 | 8 | 36.2478 | 39 | 5 | 841.2971 | 893.0669 | 802 | 928 | 0.16391 | 0.7
I2 | 5 | 2 | 866.9738 | 867.0969 | 865 | 8 | 36.269 | 39 | 5 | 841.3478 | 893.1899 | 802 | 928 | 0.16397 | 0.7
I2 | 6 | 2 | 867.0596 | 867.1832 | 865 | 8 | 36.2873 | 39 | 5 | 841.3969 | 893.3052 | 802 | 928 | 0.1648 | 0.7
I2 | 7 | 2 | 867.137 | 867.2594 | 865 | 8 | 36.3023 | 39 | 5 | 841.434 | 893.3829 | 802 | 928 | 0.14953 | 0.7
I2 | 8 | 2 | 867.2059 | 867.3266 | 865 | 8 | 36.3204 | 39 | 5 | 841.4933 | 893.4616 | 802 | 928 | 0.14575 | 0.7
I2 | 9 | 2 | 867.2664 | 867.3862 | 865 | 8 | 36.3338 | 39 | 5 | 841.5363 | 893.4963 | 802 | 928 | 0.14049 | 0.7
I2 | 10 | 2 | 867.3175 | 867.4338 | 865 | 8 | 36.347 | 39 | 5 | 841.5524 | 893.5588 | 802 | 928 | 0.14318 | 0.7
I2 | 11 | 2 | 867.3627 | 867.4782 | 865 | 8 | 36.3553 | 39 | 5 | 841.5918 | 893.6035 | 802 | 928 | 0.1362 | 0.7
I2 | 12 | 2 | 867.4047 | 867.5193 | 865 | 8 | 36.369 | 39 | 5 | 841.629 | 893.6318 | 802 | 928 | 0.13219 | 0.7
I2 | 13 | 2 | 867.4377 | 867.5495 | 865 | 8 | 36.3771 | 39 | 5 | 841.6613 | 893.6385 | 802 | 928 | 0.13684 | 0.7
I2 | 14 | 2 | 867.4728 | 867.582 | 865 | 8 | 36.387 | 39 | 5 | 841.6894 | 893.6572 | 802 | 928 | 0.13005 | 0.7
I2 | 15 | 2 | 867.5015 | 867.6067 | 865 | 8 | 36.393 | 39 | 5 | 841.7029 | 893.6894 | 802 | 928 | 0.1275 | 0.7
I2 | 16 | 2 | 867.5211 | 867.6238 | 865 | 8 | 36.4044 | 39 | 5 | 841.7023 | 893.7242 | 802 | 928 | 0.13255 | 0.7
I2 | 17 | 2 | 867.537 | 867.6389 | 865 | 8 | 36.4101 | 39 | 5 | 841.7296 | 893.7114 | 802 | 928 | 0.12599 | 0.7
I2 | 18 | 2 | 867.5509 | 867.6523 | 865 | 8 | 36.4183 | 39 | 5 | 841.7293 | 893.7395 | 802 | 928 | 0.12938 | 0.7
I2 | 19 | 2 | 867.5599 | 867.6593 | 865 | 8 | 36.4205 | 39 | 5 | 841.7398 | 893.7308 | 802 | 928 | 0.13242 | 0.7
I2 | 20 | 2 | 867.5611 | 867.6594 | 865 | 8 | 36.4274 | 39 | 5 | 841.7136 | 893.7506 | 802 | 928 | 0.1284 | 0.7
I2 | 21 | 2 | 867.5573 | 867.6559 | 865 | 8 | 36.4293 | 39 | 5 | 841.7006 | 893.7368 | 802 | 928 | 0.13329 | 0.7
I2 | 22 | 2 | 867.5551 | 867.6525 | 865 | 8 | 36.4366 | 39 | 5 | 841.6896 | 893.7442 | 802 | 928 | 0.13437 | 0.7
I2 | 23 | 2 | 867.556 | 867.6542 | 865 | 8 | 36.428 | 39 | 5 | 841.6931 | 893.731 | 802 | 928 | 0.13006 | 0.7
I2 | 24 | 2 | 867.5525 | 867.6497 | 865 | 8 | 36.4341 | 39 | 5 | 841.6799 | 893.7451 | 802 | 928 | 0.1341 | 0.7
I2 | 25 | 2 | 867.5352 | 867.6334 | 865 | 8 | 36.4241 | 39 | 5 | 841.6632 | 893.7097 | 802 | 928 | 0.14055 | 0.7
I2 | 26 | 2 | 867.5199 | 867.6191 | 865 | 8 | 36.4221 | 39 | 5 | 841.6399 | 893.6973 | 802 | 928 | 0.13497 | 0.7
I2 | 27 | 2 | 867.4979 | 867.5973 | 865 | 8 | 36.4178 | 39 | 5 | 841.6083 | 893.6783 | 802 | 928 | 0.13308 | 0.7
I2 | 28 | 2 | 867.4835 | 867.5833 | 865 | 8 | 36.4113 | 39 | 5 | 841.5818 | 893.6716 | 802 | 928 | 0.13404 | 0.7
I2 | 29 | 2 | 867.4706 | 867.5684 | 865 | 8 | 36.3907 | 39 | 5 | 841.5484 | 893.6533 | 802 | 928 | 0.13691 | 0.7
I2 | 30 | 2 | 867.4619 | 867.5525 | 865 | 8 | 36.374 | 39 | 5 | 841.5253 | 893.6161 | 802 | 928 | 0.13937 | 0.7
I2 | 31 | 2 | 867.4556 | 867.5378 | 865 | 8 | 36.3584 | 39 | 5 | 841.504 | 893.5797 | 802 | 928 | 0.14492 | 0.7
I2 | 32 | 2 | 867.4296 | 867.5072 | 865 | 8 | 36.3471 | 39 | 5 | 841.465 | 893.5365 | 802 | 928 | 0.15431 | 0.7
M1 | 1 | 1 | 411.1141 | 410.946 | 412 | 2 | 16.9013 | 20 | 2 | 395.0637 | 424.9392 | 376 | 444 | 0.30052 | 1
M1 | 2 | 1 | 411.3509 | 411.0869 | 412 | 2 | 16.7548 | 20 | 2 | 395.1219 | 425.0707 | 376 | 444 | 0.30064 | 1
M1 | 3 | 1 | 411.442 | 411.2094 | 412 | 2 | 16.8279 | 20 | 2 | 395.2012 | 425.3008 | 376 | 444 | 0.27773 | 1
M1 | 4 | 1 | 411.4143 | 411.2934 | 412 | 2 | 17.1179 | 20 | 2 | 395.265 | 425.482 | 376 | 444 | 0.2626 | 1
M1 | 5 | 1 | 411.3868 | 411.3527 | 412 | 2 | 17.3169 | 20 | 2 | 395.3332 | 425.5987 | 376 | 444 | 0.24093 | 1
M1 | 6 | 1 | 411.3508 | 411.3848 | 412 | 2 | 17.4805 | 20 | 2 | 395.3799 | 425.6661 | 376 | 444 | 0.2351 | 1
M1 | 7 | 1 | 411.3093 | 411.3881 | 412 | 2 | 17.5938 | 20 | 2 | 395.4099 | 425.688 | 376 | 444 | 0.2278 | 1
M1 | 8 | 1 | 411.2556 | 411.3604 | 412 | 2 | 17.6748 | 20 | 2 | 395.4142 | 425.6604 | 376 | 444 | 0.22539 | 1
M1 | 9 | 1 | 411.1982 | 411.3199 | 412 | 2 | 17.7375 | 20 | 2 | 395.3951 | 425.6271 | 376 | 444 | 0.22955 | 1
M1 | 10 | 1 | 411.136 | 411.2672 | 412 | 2 | 17.7847 | 20 | 2 | 395.3749 | 425.5769 | 376 | 444 | 0.2319 | 1
M1 | 11 | 1 | 411.0684 | 411.2053 | 412 | 2 | 17.8214 | 20 | 2 | 395.3411 | 425.5133 | 376 | 444 | 0.24047 | 1
M1 | 12 | 1 | 410.9789 | 411.1187 | 412 | 2 | 17.8569 | 20 | 2 | 395.2913 | 425.4218 | 376 | 444 | 0.24999 | 1
M1 | 13 | 1 | 410.8761 | 411.0182 | 412 | 2 | 17.8864 | 20 | 2 | 395.2426 | 425.2903 | 376 | 444 | 0.26052 | 1
M1 | 14 | 1 | 410.7791 | 410.9234 | 412 | 2 | 17.9138 | 20 | 2 | 395.1833 | 425.1833 | 376 | 444 | 0.27666 | 1
M1 | 15 | 1 | 410.6681 | 410.8146 | 412 | 2 | 17.9241 | 20 | 2 | 395.1279 | 424.9939 | 376 | 444 | 0.30384 | 1
M1 | 16 | 1 | 410.5531 | 410.708 | 412 | 2 | 17.9415 | 20 | 2 | 395.0874 | 424.9562 | 376 | 444 | 0.29348 | 1
M2 | 1 | 1 | 444.3832 | 444.276 | 445 | 3 | 17.1659 | 18 | 2 | 428.7017 | 457.5455 | 417 | 473 | 0.35959 | 1
M2 | 2 | 1 | 444.5592 | 444.4576 | 445 | 3 | 17.1071 | 18 | 2 | 428.9982 | 457.6822 | 417 | 473 | 0.35887 | 1
M2 | 3 | 1 | 444.6994 | 444.6039 | 445 | 3 | 17.0724 | 18 | 2 | 429.1373 | 457.8474 | 417 | 473 | 0.34183 | 1
M2 | 4 | 1 | 444.8129 | 444.7194 | 445 | 3 | 17.0512 | 18 | 2 | 429.2496 | 457.9633 | 417 | 473 | 0.33467 | 1
M2 | 5 | 1 | 444.902 | 444.8076 | 445 | 3 | 17.0429 | 18 | 2 | 429.3494 | 458.0444 | 417 | 473 | 0.33199 | 1
M2 | 6 | 1 | 444.963 | 444.8659 | 445 | 3 | 17.0356 | 18 | 2 | 429.4319 | 458.0905 | 417 | 473 | 0.32523 | 1
M2 | 7 | 1 | 444.9957 | 444.8972 | 445 | 3 | 17.0077 | 18 | 2 | 429.4963 | 458.1039 | 417 | 473 | 0.32924 | 1
M2 | 8 | 1 | 445.0087 | 444.9075 | 445 | 3 | 17.0371 | 18 | 2 | 429.5238 | 458.1248 | 417 | 473 | 0.32607 | 1
M2 | 9 | 1 | 445.0066 | 444.9029 | 445 | 3 | 17.0395 | 18 | 2 | 429.537 | 458.1233 | 417 | 473 | 0.32263 | 1
M2 | 10 | 1 | 444.9861 | 444.8819 | 445 | 3 | 17.0215 | 18 | 2 | 429.5526 | 458.0793 | 417 | 473 | 0.31565 | 1
M2 | 11 | 1 | 444.9538 | 444.8475 | 445 | 3 | 17.035 | 18 | 2 | 429.5307 | 458.06 | 417 | 473 | 0.3239 | 1
M2 | 12 | 1 | 444.8961 | 444.7883 | 445 | 3 | 17.0376 | 18 | 2 | 429.4847 | 457.9917 | 417 | 473 | 0.32059 | 1
M2 | 13 | 1 | 444.8451 | 444.7347 | 445 | 3 | 17.0394 | 18 | 2 | 429.4406 | 457.9425 | 417 | 473 | 0.32201 | 1
M2 | 14 | 1 | 444.7759 | 444.6613 | 445 | 3 | 17.0384 | 18 | 2 | 429.3675 | 457.8661 | 417 | 473 | 0.33638 | 1
M2 | 15 | 1 | 444.7058 | 444.5879 | 445 | 3 | 17.0385 | 18 | 2 | 429.2989 | 457.7719 | 417 | 473 | 0.3358 | 1
M2 | 16 | 1 | 444.6205 | 444.496 | 445 | 3 | 17.0218 | 18 | 2 | 429.2001 | 457.6304 | 417 | 473 | 0.36534 | 1
M3 | 1 | 1 | 488.4659 | 488.7717 | 488 | 4 | 18.8486 | 20 | 3 | 472.8764 | 504.2115 | 455 | 521 | 0.29519 | 0.7
M3 | 2 | 1 | 488.6071 | 488.9145 | 488 | 4 | 18.8622 | 20 | 3 | 473.0857 | 504.3017 | 455 | 521 | 0.30217 | 0.7
M3 | 3 | 1 | 488.7201 | 489.0268 | 488 | 4 | 18.8632 | 20 | 3 | 473.2342 | 504.3605 | 455 | 521 | 0.31341 | 0.7
M3 | 4 | 1 | 488.8024 | 489.1115 | 488 | 4 | 18.8694 | 20 | 3 | 473.2891 | 504.4195 | 455 | 521 | 0.3228 | 0.7
M3 | 5 | 1 | 488.8573 | 489.1694 | 488 | 4 | 18.873 | 20 | 3 | 473.3432 | 504.5015 | 455 | 521 | 0.30262 | 0.7
M3 | 6 | 1 | 488.8918 | 489.2082 | 488 | 4 | 18.8725 | 20 | 3 | 473.4037 | 504.526 | 455 | 521 | 0.29294 | 0.7
M3 | 7 | 1 | 488.9016 | 489.2201 | 488 | 4 | 18.8726 | 20 | 3 | 473.4204 | 504.5608 | 455 | 521 | 0.30064 | 0.7
M3 | 8 | 1 | 488.8916 | 489.2109 | 488 | 4 | 18.8712 | 20 | 3 | 473.4174 | 504.5427 | 455 | 521 | 0.29508 | 0.7
M3 | 9 | 1 | 488.8785 | 489.1961 | 488 | 4 | 18.8709 | 20 | 3 | 473.4137 | 504.4995 | 455 | 521 | 0.29575 | 0.7
M3 | 10 | 1 | 488.8495 | 489.1633 | 488 | 4 | 18.8734 | 20 | 3 | 473.3807 | 504.4694 | 455 | 521 | 0.30698 | 0.7
M3 | 11 | 1 | 488.8105 | 489.1208 | 488 | 4 | 18.8781 | 20 | 3 | 473.3468 | 504.4152 | 455 | 521 | 0.30856 | 0.7
M3 | 12 | 1 | 488.7571 | 489.0633 | 488 | 4 | 18.8828 | 20 | 3 | 473.2934 | 504.3922 | 455 | 521 | 0.30015 | 0.7
M3 | 13 | 1 | 488.6978 | 489.0002 | 488 | 4 | 18.8918 | 20 | 3 | 473.251 | 504.3553 | 455 | 521 | 0.28637 | 0.7
M3 | 14 | 1 | 488.6358 | 488.9337 | 488 | 4 | 18.9029 | 20 | 3 | 473.1536 | 504.3375 | 455 | 521 | 0.30845 | 0.7
M3 | 15 | 1 | 488.565 | 488.862 | 488 | 4 | 18.9128 | 20 | 3 | 473.0546 | 504.3038 | 455 | 521 | 0.29985 | 0.7
M3 | 16 | 1 | 488.4944 | 488.7871 | 488 | 4 | 18.9181 | 20 | 3 | 472.9245 | 504.2682 | 455 | 521 | 0.29811 | 0.7
M4 | 1 | 1 | 556.1332 | 556.2256 | 555 | 4 | 18.2876 | 20 | 3 | 540.0069 | 573.4348 | 523 | 589 | 0.25468 | 0.7
M4 | 2 | 1 | 556.2716 | 556.3601 | 555 | 4 | 18.3698 | 20 | 3 | 540.089 | 573.6075 | 523 | 589 | 0.25365 | 0.7
M4 | 3 | 1 | 556.381 | 556.4705 | 555 | 4 | 18.4469 | 20 | 3 | 540.1577 | 573.7143 | 523 | 589 | 0.25355 | 0.7
M4 | 4 | 1 | 556.4572 | 556.5532 | 555 | 4 | 18.5197 | 20 | 3 | 540.2006 | 573.7971 | 523 | 589 | 0.25003 | 0.7
M4 | 5 | 1 | 556.5045 | 556.6146 | 555 | 4 | 18.561 | 20 | 3 | 540.2412 | 573.8288 | 523 | 589 | 0.24115 | 0.7
M4 | 6 | 1 | 556.5209 | 556.6486 | 555 | 4 | 18.5493 | 20 | 3 | 540.2615 | 573.8373 | 523 | 589 | 0.24306 | 0.7
M4 | 7 | 1 | 556.5021 | 556.6482 | 555 | 4 | 18.494 | 20 | 3 | 540.2761 | 573.788 | 523 | 589 | 0.23763 | 0.7
M4 | 8 | 1 | 556.4598 | 556.6212 | 555 | 4 | 18.4339 | 20 | 3 | 540.2442 | 573.7311 | 523 | 589 | 0.2473 | 0.7
M4 | 9 | 1 | 556.3993 | 556.5727 | 555 | 4 | 18.3914 | 20 | 3 | 540.2171 | 573.6611 | 523 | 589 | 0.24952 | 0.7
M4 | 10 | 1 | 556.3193 | 556.5081 | 555 | 4 | 18.3514 | 20 | 3 | 540.1595 | 573.5585 | 523 | 589 | 0.26108 | 0.7
M4 | 11 | 1 | 556.2254 | 556.4298 | 555 | 4 | 18.3123 | 20 | 3 | 540.0759 | 573.4654 | 523 | 589 | 0.27427 | 0.7
M4 | 12 | 1 | 556.1149 | 556.3331 | 555 | 4 | 18.2682 | 20 | 3 | 539.9988 | 573.3376 | 523 | 589 | 0.26958 | 0.7
M4 | 13 | 1 | 556.0039 | 556.2338 | 555 | 4 | 18.2224 | 20 | 3 | 539.8871 | 573.2225 | 523 | 589 | 0.26166 | 0.7
M4 | 14 | 1 | 555.882 | 556.1188 | 555 | 4 | 18.1879 | 20 | 3 | 539.7793 | 573.1326 | 523 | 589 | 0.24536 | 0.7
M4 | 15 | 1 | 555.7605 | 556.0023 | 555 | 4 | 18.1487 | 20 | 3 | 539.6652 | 573.0312 | 523 | 589 | 0.24292 | 0.7
M4 | 16 | 1 | 555.6375 | 555.8814 | 555 | 4 | 18.1175 | 20 | 3 | 539.5363 | 572.9221 | 523 | 589 | 0.26384 | 0.7
M5 | 1 | 1 | 666.962 | 667.0623 | 672 | 5 | 18.9798 | 20 | 3 | 649.6962 | 684.563 | 638 | 706 | 0.29733 | 0.7
M5 | 2 | 1 | 667.1182 | 667.2177 | 672 | 5 | 19.0364 | 20 | 3 | 649.817 | 684.7078 | 638 | 706 | 0.28924 | 0.7
M5 | 3 | 1 | 667.2444 | 667.3361 | 672 | 5 | 19.0976 | 20 | 3 | 649.9207 | 684.8575 | 638 | 706 | 0.27031 | 0.7
M5 | 4 | 1 | 667.3386 | 667.4232 | 672 | 5 | 19.1357 | 20 | 3 | 649.9849 | 684.9753 | 638 | 706 | 0.27229 | 0.7
M5 | 5 | 1 | 667.403 | 667.483 | 672 | 5 | 19.1719 | 20 | 3 | 650.046 | 685.0315 | 638 | 706 | 0.26604 | 0.7
M5 | 6 | 1 | 667.4349 | 667.514 | 672 | 5 | 19.2005 | 20 | 3 | 650.0529 | 685.0815 | 638 | 706 | 0.26567 | 0.7
M5 | 7 | 1 | 667.4558 | 667.5322 | 672 | 5 | 19.223 | 20 | 3 | 650.0609 | 685.1022 | 638 | 706 | 0.27696 | 0.7
M5 | 8 | 1 | 667.4647 | 667.5408 | 672 | 5 | 19.243 | 20 | 3 | 650.0558 | 685.1047 | 638 | 706 | 0.27556 | 0.7
M5 | 9 | 1 | 667.4546 | 667.5261 | 672 | 5 | 19.2529 | 20 | 3 | 650.0184 | 685.1031 | 638 | 706 | 0.26339 | 0.7
M5 | 10 | 1 | 667.436 | 667.5043 | 672 | 5 | 19.2632 | 20 | 3 | 649.9719 | 685.0939 | 638 | 706 | 0.26828 | 0.7
M5 | 11 | 1 | 667.3958 | 667.4629 | 672 | 5 | 19.2684 | 20 | 3 | 649.9233 | 685.0446 | 638 | 706 | 0.26598 | 0.7
M5 | 12 | 1 | 667.3519 | 667.4181 | 672 | 5 | 19.2705 | 20 | 3 | 649.8519 | 685.0148 | 638 | 706 | 0.28053 | 0.7
M5 | 13 | 1 | 667.3004 | 667.3647 | 672 | 5 | 19.2713 | 20 | 3 | 649.7662 | 684.9692 | 638 | 706 | 0.29171 | 0.7
M5 | 14 | 1 | 667.2322 | 667.2971 | 672 | 5 | 19.2604 | 20 | 3 | 649.7061 | 684.9026 | 638 | 706 | 0.28288 | 0.7
M5 | 15 | 1 | 667.1519 | 667.2155 | 672 | 5 | 19.2557 | 20 | 3 | 649.6412 | 684.8216 | 638 | 706 | 0.27383 | 0.7
M5 | 16 | 1 | 667.0373 | 667.095 | 672 | 5 | 19.2244 | 20 | 3 | 649.4997 | 684.7524 | 638 | 706 | 0.28872 | 0.7
M6 | 1 | 1 | 745.6092 | 745.552 | 746 | 2 | 13.504 | 15 | 2 | 733.6464 | 757.5811 | 721 | 771 | 0.2734 | 0.8
M6 | 2 | 1 | 745.7717 | 745.7103 | 746 | 2 | 13.5062 | 15 | 2 | 733.7592 | 757.7768 | 721 | 771 | 0.26196 | 0.8
M6 | 3 | 1 | 745.9097 | 745.8406 | 746 | 2 | 13.5087 | 15 | 2 | 733.8844 | 757.8924 | 721 | 771 | 0.26386 | 0.8
M6 | 4 | 1 | 746.0144 | 745.9374 | 746 | 2 | 13.5148 | 15 | 2 | 733.9794 | 757.9858 | 721 | 771 | 0.26388 | 0.8
M6 | 5 | 1 | 746.0909 | 746.0091 | 746 | 2 | 13.5179 | 15 | 2 | 734.0468 | 758.0562 | 721 | 771 | 0.26424 | 0.8
M6 | 6 | 1 | 746.1393 | 746.0541 | 746 | 2 | 13.5137 | 15 | 2 | 734.1044 | 758.0934 | 721 | 771 | 0.26661 | 0.8
M6 | 7 | 1 | 746.1564 | 746.0692 | 746 | 2 | 13.5104 | 15 | 2 | 734.1375 | 758.0945 | 721 | 771 | 0.26157 | 0.8
M6 | 8 | 1 | 746.1522 | 746.0635 | 746 | 2 | 13.5095 | 15 | 2 | 734.1342 | 758.0878 | 721 | 771 | 0.26054 | 0.8
M6 | 9 | 1 | 746.1373 | 746.0475 | 746 | 2 | 13.5051 | 15 | 2 | 734.1022 | 758.0888 | 721 | 771 | 0.26878 | 0.8
M6 | 10 | 1 | 746.1091 | 746.0176 | 746 | 2 | 13.5021 | 15 | 2 | 734.0656 | 758.0595 | 721 | 771 | 0.27153 | 0.8
M6 | 11 | 1 | 746.074 | 745.9822 | 746 | 2 | 13.5045 | 15 | 2 | 734.0186 | 758.0344 | 721 | 771 | 0.26486 | 0.8
M6 | 12 | 1 | 746.0245 | 745.9337 | 746 | 2 | 13.5031 | 15 | 2 | 733.9553 | 757.993 | 721 | 771 | 0.27148 | 0.8
M6 | 13 | 1 | 745.9669 | 745.8776 | 746 | 2 | 13.5021 | 15 | 2 | 733.8807 | 757.9627 | 721 | 771 | 0.26855 | 0.8
M6 | 14 | 1 | 745.9058 | 745.8198 | 746 | 2 | 13.5036 | 15 | 2 | 733.8136 | 757.9225 | 721 | 771 | 0.26986 | 0.8
M6 | 15 | 1 | 745.8373 | 745.755 | 746 | 2 | 13.504 | 15 | 2 | 733.7551 | 757.8407 | 721 | 771 | 0.26684 | 0.8
M6 | 16 | 1 | 745.7662 | 745.6865 | 746 | 2 | 13.5104 | 15 | 2 | 733.6657 | 757.814 | 721 | 771 | 0.27208 | 0.8
M7 | 1 | 1 | 866.7959 | 866.9595 | 865 | 8 | 36.0918 | 39 | 5 | 842.1199 | 892.0074 | 801 | 929 | 0.12158 | 0.7
M7 | 2 | 1 | 867.0135 | 867.1684 | 865 | 8 | 36.1193 | 39 | 5 | 842.3109 | 892.1813 | 801 | 929 | 0.11726 | 0.7
M7 | 3 | 1 | 867.2056 | 867.3544 | 865 | 8 | 36.1507 | 39 | 5 | 842.4998 | 892.3176 | 801 | 929 | 0.11333 | 0.7
M7 | 4 | 1 | 867.364 | 867.5079 | 865 | 8 | 36.1737 | 39 | 5 | 842.6847 | 892.4214 | 801 | 929 | 0.11367 | 0.7
M7 | 5 | 1 | 867.4822 | 867.6216 | 865 | 8 | 36.1945 | 39 | 5 | 842.8571 | 892.4714 | 801 | 929 | 0.15712 | 0.7
M7 | 6 | 1 | 867.573 | 867.7182 | 865 | 8 | 36.1939 | 39 | 5 | 843.0021 | 892.5202 | 801 | 929 | 0.1531 | 0.7
M7 | 7 | 1 | 867.6434 | 867.777 | 865 | 8 | 36.1939 | 39 | 5 | 843.1028 | 892.5459 | 801 | 929 | 0.15081 | 0.7
M7 | 8 | 1 | 867.6896 | 867.8189 | 865 | 8 | 36.2089 | 39 | 5 | 843.1909 | 892.5645 | 801 | 929 | 0.14881 | 0.7
M7 | 9 | 1 | 867.728 | 867.8549 | 865 | 8 | 36.2264 | 39 | 5 | 843.2079 | 892.5817 | 801 | 929 | 0.14758 | 0.7
M7 | 10 | 1 | 867.7413 | 867.8702 | 865 | 8 | 36.2314 | 39 | 5 | 843.2015 | 892.5878 | 801 | 929 | 0.14863 | 0.7
M7 | 11 | 1 | 867.7542 | 867.8847 | 865 | 8 | 36.2296 | 39 | 5 | 843.1991 | 892.5948 | 801 | 929 | 0.14955 | 0.7
M7 | 12 | 1 | 867.7465 | 867.8762 | 865 | 8 | 36.2239 | 39 | 5 | 843.1667 | 892.5885 | 801 | 929 | 0.14936 | 0.7
M7 | 13 | 1 | 867.7258 | 867.8562 | 865 | 8 | 36.2161 | 39 | 5 | 843.0742 | 892.5752 | 801 | 929 | 0.15227 | 0.7
M7 | 14 | 1 | 867.7077 | 867.8393 | 865 | 8 | 36.2268 | 39 | 5 | 842.9914 | 892.5701 | 801 | 929 | 0.1544 | 0.7
M7 | 15 | 1 | 867.692 | 867.8259 | 865 | 8 | 36.2255 | 39 | 5 | 842.9083 | 892.5637 | 801 | 929 | 0.15677 | 0.7
M7 | 16 | 1 | 867.6328 | 867.7765 | 865 | 8 | 36.2173 | 39 | 5 | 842.8104 | 892.5214 | 801 | 929 | 0.15958 | 0.7

## SMWIR
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source. | [8 empty] | Non-compliance with specification
3. The SMWIR bands were measured duirng TV testing at the nominal iinstrument temperature plateau.
4. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
5. Odd and even detector arrays were measured separately for the SMWIR bands.
6. Atmospheric correction applied to M13.
Band | Detector | Subsample | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
|  |  | Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
|  |  | Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
I3 | 1 | 1 | 1600.6602 | 1601.3285 | 1610 | 14 | 58.6332 | 60 | 9 | 1542.2968 | 1663.3218 | 1509 | 1709 | 0.46875 | 0.7
I3 | 2 | 1 | 1601.0645 | 1601.7146 | 1610 | 14 | 59.5114 | 60 | 9 | 1542.4252 | 1664.1949 | 1509 | 1709 | 0.45868 | 0.7
I3 | 3 | 1 | 1601.3037 | 1601.8478 | 1610 | 14 | 59.3296 | 60 | 9 | 1542.7733 | 1664.1014 | 1509 | 1709 | 0.46865 | 0.7
I3 | 4 | 1 | 1601.7067 | 1602.2599 | 1610 | 14 | 59.6514 | 60 | 9 | 1543.2065 | 1665.2158 | 1509 | 1709 | 0.5865 | 0.7
I3 | 5 | 1 | 1601.6995 | 1602.3147 | 1610 | 14 | 59.6896 | 60 | 9 | 1543.3098 | 1665.1567 | 1509 | 1709 | 0.46317 | 0.7
I3 | 6 | 1 | 1602.1818 | 1602.6947 | 1610 | 14 | 59.9412 | 60 | 9 | 1543.7982 | 1665.6641 | 1509 | 1709 | 0.4646 | 0.7
I3 | 7 | 1 | 1602.2212 | 1602.8696 | 1610 | 14 | 60.0068 | 60 | 9 | 1543.976 | 1665.9536 | 1509 | 1709 | 0.45209 | 0.7
I3 | 8 | 1 | 1602.6465 | 1603.2079 | 1610 | 14 | 60.0192 | 60 | 9 | 1544.3243 | 1666.536 | 1509 | 1709 | 0.45687 | 0.7
I3 | 9 | 1 | 1602.6326 | 1603.2959 | 1610 | 14 | 60.1042 | 60 | 9 | 1544.5151 | 1666.7241 | 1509 | 1709 | 0.45166 | 0.7
I3 | 10 | 1 | 1603.0009 | 1603.5648 | 1610 | 14 | 60.2319 | 60 | 9 | 1544.5856 | 1667.0375 | 1509 | 1709 | 0.45428 | 0.7
I3 | 11 | 1 | 1603.0165 | 1603.6703 | 1610 | 14 | 60.1492 | 60 | 9 | 1544.8093 | 1667.4783 | 1509 | 1709 | 0.44991 | 0.7
I3 | 12 | 1 | 1603.3323 | 1603.8882 | 1610 | 14 | 60.3419 | 60 | 9 | 1544.9597 | 1667.4796 | 1509 | 1709 | 0.45416 | 0.7
I3 | 13 | 1 | 1603.2932 | 1603.9597 | 1610 | 14 | 60.3608 | 60 | 9 | 1545.1669 | 1667.8207 | 1509 | 1709 | 0.45135 | 0.7
I3 | 14 | 1 | 1603.4465 | 1603.994 | 1610 | 14 | 60.3163 | 60 | 9 | 1545.033 | 1667.7673 | 1509 | 1709 | 0.45197 | 0.7
I3 | 15 | 1 | 1603.4015 | 1604.0973 | 1610 | 14 | 60.4434 | 60 | 9 | 1545.2816 | 1667.9663 | 1509 | 1709 | 0.45772 | 0.7
I3 | 16 | 1 | 1603.6489 | 1604.1925 | 1610 | 14 | 60.5427 | 60 | 9 | 1545.0392 | 1668.1532 | 1509 | 1709 | 0.44982 | 0.7
I3 | 17 | 1 | 1603.4944 | 1604.2009 | 1610 | 14 | 60.6077 | 60 | 9 | 1545.2682 | 1668.3553 | 1509 | 1709 | 0.4392 | 0.7
I3 | 18 | 1 | 1603.7524 | 1604.2938 | 1610 | 14 | 60.7307 | 60 | 9 | 1545.0696 | 1668.3265 | 1509 | 1709 | 0.44607 | 0.7
I3 | 19 | 1 | 1603.5847 | 1604.3071 | 1610 | 14 | 60.8896 | 60 | 9 | 1545.1957 | 1668.6193 | 1509 | 1709 | 0.44359 | 0.7
I3 | 20 | 1 | 1603.9551 | 1604.5134 | 1610 | 14 | 60.9139 | 60 | 9 | 1545.1887 | 1668.6542 | 1509 | 1709 | 0.44124 | 0.7
I3 | 21 | 1 | 1603.7327 | 1604.4536 | 1610 | 14 | 60.8679 | 60 | 9 | 1545.3333 | 1668.8282 | 1509 | 1709 | 0.43616 | 0.7
I3 | 22 | 1 | 1603.9028 | 1604.4751 | 1610 | 14 | 61.1226 | 60 | 9 | 1544.9991 | 1668.6112 | 1509 | 1709 | 0.43833 | 0.7
I3 | 23 | 1 | 1603.7875 | 1604.5154 | 1610 | 14 | 61.1404 | 60 | 9 | 1545.1471 | 1668.9929 | 1509 | 1709 | 0.4407 | 0.7
I3 | 24 | 1 | 1603.9189 | 1604.4966 | 1610 | 14 | 61.1954 | 60 | 9 | 1544.8636 | 1668.5818 | 1509 | 1709 | 0.44518 | 0.7
I3 | 25 | 1 | 1603.842 | 1604.5745 | 1610 | 14 | 61.0044 | 60 | 9 | 1545.0676 | 1668.9724 | 1509 | 1709 | 0.44971 | 0.7
I3 | 26 | 1 | 1604.0369 | 1604.6666 | 1610 | 14 | 61.0676 | 60 | 9 | 1544.8683 | 1668.6456 | 1509 | 1709 | 0.45174 | 0.7
I3 | 27 | 1 | 1603.8184 | 1604.5759 | 1610 | 14 | 61.0605 | 60 | 9 | 1545.0018 | 1668.7784 | 1509 | 1709 | 0.44283 | 0.7
I3 | 28 | 1 | 1604.0159 | 1604.666 | 1610 | 14 | 61.0127 | 60 | 9 | 1544.85 | 1668.5248 | 1509 | 1709 | 0.44856 | 0.7
I3 | 29 | 1 | 1603.8561 | 1604.6094 | 1610 | 14 | 60.9487 | 60 | 9 | 1544.8286 | 1668.7039 | 1509 | 1709 | 0.44632 | 0.7
I3 | 30 | 1 | 1603.943 | 1604.593 | 1610 | 14 | 60.9353 | 60 | 9 | 1544.5918 | 1668.1494 | 1509 | 1709 | 0.46185 | 0.7
I3 | 31 | 1 | 1603.8059 | 1604.5867 | 1610 | 14 | 60.8702 | 60 | 9 | 1544.7263 | 1668.2831 | 1509 | 1709 | 0.45969 | 0.7
I3 | 32 | 1 | 1603.9656 | 1604.6472 | 1610 | 14 | 60.9344 | 60 | 9 | 1544.6533 | 1668.1965 | 1509 | 1709 | 0.45909 | 0.7
I3 | 1 | 2 | 1600.66 | 1601.3285 | 1610 | 14 | 58.6335 | 60 | 9 | 1542.295 | 1663.3422 | 1509 | 1709 | 0.469 | 0.7
I3 | 2 | 2 | 1601.0647 | 1601.7146 | 1610 | 14 | 59.5104 | 60 | 9 | 1542.4434 | 1664.1917 | 1509 | 1709 | 0.45887 | 0.7
I3 | 3 | 2 | 1601.3033 | 1601.8478 | 1610 | 14 | 59.3296 | 60 | 9 | 1542.7871 | 1664.0903 | 1509 | 1709 | 0.46884 | 0.7
I3 | 4 | 2 | 1601.7081 | 1602.2572 | 1610 | 14 | 59.6653 | 60 | 9 | 1543.136 | 1666.3278 | 1509 | 1709 | 0.54947 | 0.7
I3 | 5 | 2 | 1601.6992 | 1602.314 | 1610 | 14 | 59.6895 | 60 | 9 | 1543.3015 | 1665.1385 | 1509 | 1709 | 0.46281 | 0.7
I3 | 6 | 2 | 1602.182 | 1602.6947 | 1610 | 14 | 59.9421 | 60 | 9 | 1543.788 | 1665.673 | 1509 | 1709 | 0.46522 | 0.7
I3 | 7 | 2 | 1602.2212 | 1602.8691 | 1610 | 14 | 60.0065 | 60 | 9 | 1543.99 | 1665.9325 | 1509 | 1709 | 0.45405 | 0.7
I3 | 8 | 2 | 1602.6467 | 1603.208 | 1610 | 14 | 60.0197 | 60 | 9 | 1544.3196 | 1666.5502 | 1509 | 1709 | 0.45753 | 0.7
I3 | 9 | 2 | 1602.6324 | 1603.2963 | 1610 | 14 | 60.104 | 60 | 9 | 1544.5089 | 1666.7109 | 1509 | 1709 | 0.45243 | 0.7
I3 | 10 | 2 | 1603.0005 | 1603.5641 | 1610 | 14 | 60.2316 | 60 | 9 | 1544.5835 | 1667.0441 | 1509 | 1709 | 0.45431 | 0.7
I3 | 11 | 2 | 1603.0164 | 1603.6705 | 1610 | 14 | 60.1503 | 60 | 9 | 1544.7998 | 1667.4774 | 1509 | 1709 | 0.44952 | 0.7
I3 | 12 | 2 | 1603.3325 | 1603.8871 | 1610 | 14 | 60.3412 | 60 | 9 | 1544.9539 | 1667.4817 | 1509 | 1709 | 0.45239 | 0.7
I3 | 13 | 2 | 1603.2935 | 1603.96 | 1610 | 14 | 60.3608 | 60 | 9 | 1545.168 | 1667.816 | 1509 | 1709 | 0.45284 | 0.7
I3 | 14 | 2 | 1603.4465 | 1603.9949 | 1610 | 14 | 60.3159 | 60 | 9 | 1545.0392 | 1667.7667 | 1509 | 1709 | 0.45108 | 0.7
I3 | 15 | 2 | 1603.4015 | 1604.0962 | 1610 | 14 | 60.4434 | 60 | 9 | 1545.2777 | 1667.9425 | 1509 | 1709 | 0.45688 | 0.7
I3 | 16 | 2 | 1603.6489 | 1604.1932 | 1610 | 14 | 60.5426 | 60 | 9 | 1545.0548 | 1668.1782 | 1509 | 1709 | 0.44909 | 0.7
I3 | 17 | 2 | 1603.4943 | 1604.2013 | 1610 | 14 | 60.6079 | 60 | 9 | 1545.2758 | 1668.3682 | 1509 | 1709 | 0.44039 | 0.7
I3 | 18 | 2 | 1603.7522 | 1604.2944 | 1610 | 14 | 60.7303 | 60 | 9 | 1545.0691 | 1668.2999 | 1509 | 1709 | 0.44793 | 0.7
I3 | 19 | 2 | 1603.5845 | 1604.3079 | 1610 | 14 | 60.8898 | 60 | 9 | 1545.2013 | 1668.624 | 1509 | 1709 | 0.44445 | 0.7
I3 | 20 | 2 | 1603.9548 | 1604.5137 | 1610 | 14 | 60.913 | 60 | 9 | 1545.1848 | 1668.676 | 1509 | 1709 | 0.44023 | 0.7
I3 | 21 | 2 | 1603.7327 | 1604.4535 | 1610 | 14 | 60.8677 | 60 | 9 | 1545.3284 | 1668.8308 | 1509 | 1709 | 0.4354 | 0.7
I3 | 22 | 2 | 1603.9027 | 1604.4734 | 1610 | 14 | 61.123 | 60 | 9 | 1544.9929 | 1668.6445 | 1509 | 1709 | 0.4387 | 0.7
I3 | 23 | 2 | 1603.7874 | 1604.5148 | 1610 | 14 | 61.1404 | 60 | 9 | 1545.1516 | 1668.9746 | 1509 | 1709 | 0.44123 | 0.7
I3 | 24 | 2 | 1603.9187 | 1604.4967 | 1610 | 14 | 61.1948 | 60 | 9 | 1544.8568 | 1668.6211 | 1509 | 1709 | 0.44527 | 0.7
I3 | 25 | 2 | 1603.842 | 1604.5754 | 1610 | 14 | 61.004 | 60 | 9 | 1545.0486 | 1668.949 | 1509 | 1709 | 0.44982 | 0.7
I3 | 26 | 2 | 1604.0366 | 1604.6661 | 1610 | 14 | 61.0685 | 60 | 9 | 1544.8577 | 1668.6857 | 1509 | 1709 | 0.45161 | 0.7
I3 | 27 | 2 | 1603.818 | 1604.5757 | 1610 | 14 | 61.0608 | 60 | 9 | 1544.9927 | 1668.7706 | 1509 | 1709 | 0.44392 | 0.7
I3 | 28 | 2 | 1604.0156 | 1604.6655 | 1610 | 14 | 61.0128 | 60 | 9 | 1544.865 | 1668.5344 | 1509 | 1709 | 0.44783 | 0.7
I3 | 29 | 2 | 1603.856 | 1604.6084 | 1610 | 14 | 60.9481 | 60 | 9 | 1544.8208 | 1668.6864 | 1509 | 1709 | 0.44635 | 0.7
I3 | 30 | 2 | 1603.9431 | 1604.5935 | 1610 | 14 | 60.9346 | 60 | 9 | 1544.6008 | 1668.1578 | 1509 | 1709 | 0.4603 | 0.7
I3 | 31 | 2 | 1603.8057 | 1604.5869 | 1610 | 14 | 60.8708 | 60 | 9 | 1544.7321 | 1668.2695 | 1509 | 1709 | 0.46074 | 0.7
I3 | 32 | 2 | 1603.9653 | 1604.6465 | 1610 | 14 | 60.9348 | 60 | 9 | 1544.6599 | 1668.1993 | 1509 | 1709 | 0.45803 | 0.7
I4 | 1 | 1 | 3742.981 | 3745.4983 | 3740 | 40 | 391.3481 | 380 | 30 | 3470.8015 | 4011.5354 | 3340 | 4140 | 0.19884 | 0.5
I4 | 2 | 1 | 3745.0015 | 3745.9739 | 3740 | 40 | 390.1045 | 380 | 30 | 3471.5576 | 4012.7375 | 3340 | 4140 | 0.18786 | 0.5
I4 | 3 | 1 | 3744.3757 | 3746.5161 | 3740 | 40 | 390.6655 | 380 | 30 | 3472.3853 | 4012.7112 | 3340 | 4140 | 0.1798 | 0.5
I4 | 4 | 1 | 3746.7136 | 3747.0461 | 3740 | 40 | 388.2397 | 380 | 30 | 3473.2195 | 4013.8074 | 3340 | 4140 | 0.17771 | 0.5
I4 | 5 | 1 | 3745.675 | 3748.0273 | 3740 | 40 | 390.0117 | 380 | 30 | 3473.3806 | 4014.074 | 3340 | 4140 | 0.19351 | 0.5
I4 | 6 | 1 | 3747.3379 | 3747.7371 | 3740 | 40 | 387.9846 | 380 | 30 | 3473.4673 | 4014.6079 | 3340 | 4140 | 0.18333 | 0.5
I4 | 7 | 1 | 3746.3906 | 3749.0293 | 3740 | 40 | 389.386 | 380 | 30 | 3474.2373 | 4014.4839 | 3340 | 4140 | 0.18275 | 0.5
I4 | 8 | 1 | 3748.1118 | 3748.4719 | 3740 | 40 | 386.9414 | 380 | 30 | 3474.2031 | 4014.8325 | 3340 | 4140 | 0.18249 | 0.5
I4 | 9 | 1 | 3746.9771 | 3748.9319 | 3740 | 40 | 389.2786 | 380 | 30 | 3474.2986 | 4014.7224 | 3340 | 4140 | 0.18242 | 0.5
I4 | 10 | 1 | 3748.3135 | 3747.7864 | 3740 | 40 | 386.885 | 380 | 30 | 3474.2195 | 4014.7104 | 3340 | 4140 | 0.17787 | 0.5
I4 | 11 | 1 | 3747.9634 | 3750.1375 | 3740 | 40 | 388.3208 | 380 | 30 | 3475.1353 | 4015.0059 | 3340 | 4140 | 0.18038 | 0.5
I4 | 12 | 1 | 3748.9097 | 3748.0786 | 3740 | 40 | 386.1145 | 380 | 30 | 3474.5361 | 4014.7219 | 3340 | 4140 | 0.18327 | 0.5
I4 | 13 | 1 | 3748.0293 | 3749.8074 | 3740 | 40 | 387.8325 | 380 | 30 | 3475.053 | 4014.9614 | 3340 | 4140 | 0.18141 | 0.5
I4 | 14 | 1 | 3749.2688 | 3748.1033 | 3740 | 40 | 384.6958 | 380 | 30 | 3474.8601 | 4014.3948 | 3340 | 4140 | 0.18328 | 0.5
I4 | 15 | 1 | 3748.1177 | 3750.2705 | 3740 | 40 | 387.5457 | 380 | 30 | 3475.3992 | 4014.8413 | 3340 | 4140 | 0.18003 | 0.5
I4 | 16 | 1 | 3749.1685 | 3748.0134 | 3740 | 40 | 384.6172 | 380 | 30 | 3474.7173 | 4014.1929 | 3340 | 4140 | 0.1798 | 0.5
I4 | 17 | 1 | 3747.8684 | 3749.7671 | 3740 | 40 | 387.4531 | 380 | 30 | 3474.8989 | 4014.5957 | 3340 | 4140 | 0.17864 | 0.5
I4 | 18 | 1 | 3748.8245 | 3747.4875 | 3740 | 40 | 384.8862 | 380 | 30 | 3474.3647 | 4014.1042 | 3340 | 4140 | 0.18257 | 0.5
I4 | 19 | 1 | 3747.8076 | 3749.9478 | 3740 | 40 | 387.2871 | 380 | 30 | 3474.6892 | 4014.4636 | 3340 | 4140 | 0.18141 | 0.5
I4 | 20 | 1 | 3748.6868 | 3747.509 | 3740 | 40 | 384.4961 | 380 | 30 | 3474.354 | 4013.7822 | 3340 | 4140 | 0.18458 | 0.5
I4 | 21 | 1 | 3747.6641 | 3750.4375 | 3740 | 40 | 387.0286 | 380 | 30 | 3474.657 | 4014.1479 | 3340 | 4140 | 0.1823 | 0.5
I4 | 22 | 1 | 3748.2346 | 3747.8862 | 3740 | 40 | 384.9385 | 380 | 30 | 3473.9128 | 4013.4897 | 3340 | 4140 | 0.18148 | 0.5
I4 | 23 | 1 | 3746.6567 | 3749.2583 | 3740 | 40 | 387.2 | 380 | 30 | 3473.7278 | 4013.2834 | 3340 | 4140 | 0.18604 | 0.5
I4 | 24 | 1 | 3747.5347 | 3747.3533 | 3740 | 40 | 384.907 | 380 | 30 | 3473.356 | 4012.7239 | 3340 | 4140 | 0.18138 | 0.5
I4 | 25 | 1 | 3745.853 | 3748.8577 | 3740 | 40 | 387.155 | 380 | 30 | 3472.9683 | 4012.5315 | 3340 | 4140 | 0.18834 | 0.5
I4 | 26 | 1 | 3746.5654 | 3746.7217 | 3740 | 40 | 384.8987 | 380 | 30 | 3472.4612 | 4011.5261 | 3340 | 4140 | 0.18647 | 0.5
I4 | 27 | 1 | 3745.1792 | 3748.4795 | 3740 | 40 | 387.1985 | 380 | 30 | 3472.3201 | 4011.9998 | 3340 | 4140 | 0.18778 | 0.5
I4 | 28 | 1 | 3745.9792 | 3746.8994 | 3740 | 40 | 384.9531 | 380 | 30 | 3471.8293 | 4011.1069 | 3340 | 4140 | 0.1881 | 0.5
I4 | 29 | 1 | 3744.5229 | 3748.2925 | 3740 | 40 | 387.4019 | 380 | 30 | 3471.5535 | 4011.5854 | 3340 | 4140 | 0.18709 | 0.5
I4 | 30 | 1 | 3745.1909 | 3746.7451 | 3740 | 40 | 385.7917 | 380 | 30 | 3470.9204 | 4010.824 | 3340 | 4140 | 0.1866 | 0.5
I4 | 31 | 1 | 3743.7285 | 3748.1145 | 3740 | 40 | 387.427 | 380 | 30 | 3471.1257 | 4011.1482 | 3340 | 4140 | 0.19807 | 0.5
I4 | 32 | 1 | 3744.4741 | 3747.0798 | 3740 | 40 | 385.6816 | 380 | 30 | 3470.196 | 4010.1824 | 3340 | 4140 | 0.19166 | 0.5
I4 | 1 | 2 | 3742.981 | 3745.4971 | 3740 | 40 | 391.3467 | 380 | 30 | 3470.8235 | 4011.5327 | 3340 | 4140 | 0.19734 | 0.5
I4 | 2 | 2 | 3745.001 | 3745.9751 | 3740 | 40 | 390.1067 | 380 | 30 | 3471.5422 | 4012.9109 | 3340 | 4140 | 0.18541 | 0.5
I4 | 3 | 2 | 3744.3765 | 3746.5156 | 3740 | 40 | 390.6643 | 380 | 30 | 3472.3354 | 4012.7258 | 3340 | 4140 | 0.17748 | 0.5
I4 | 4 | 2 | 3746.7134 | 3747.0493 | 3740 | 40 | 388.2393 | 380 | 30 | 3473.2209 | 4013.897 | 3340 | 4140 | 0.17649 | 0.5
I4 | 5 | 2 | 3745.6758 | 3748.0261 | 3740 | 40 | 390.0083 | 380 | 30 | 3473.4402 | 4014.0417 | 3340 | 4140 | 0.19385 | 0.5
I4 | 6 | 2 | 3747.3367 | 3747.7407 | 3740 | 40 | 387.9888 | 380 | 30 | 3473.4656 | 4014.6172 | 3340 | 4140 | 0.18165 | 0.5
I4 | 7 | 2 | 3746.3901 | 3749.0298 | 3740 | 40 | 389.3899 | 380 | 30 | 3474.2087 | 4014.5264 | 3340 | 4140 | 0.18058 | 0.5
I4 | 8 | 2 | 3748.1125 | 3748.4727 | 3740 | 40 | 386.9385 | 380 | 30 | 3474.2009 | 4014.9067 | 3340 | 4140 | 0.18302 | 0.5
I4 | 9 | 2 | 3746.9766 | 3748.9292 | 3740 | 40 | 389.283 | 380 | 30 | 3474.2939 | 4014.7588 | 3340 | 4140 | 0.18127 | 0.5
I4 | 10 | 2 | 3748.3135 | 3747.7903 | 3740 | 40 | 386.885 | 380 | 30 | 3474.2158 | 4014.7917 | 3340 | 4140 | 0.17695 | 0.5
I4 | 11 | 2 | 3747.9619 | 3750.1323 | 3740 | 40 | 388.3225 | 380 | 30 | 3475.0908 | 4014.9214 | 3340 | 4140 | 0.17908 | 0.5
I4 | 12 | 2 | 3748.9099 | 3748.0845 | 3740 | 40 | 386.1182 | 380 | 30 | 3474.5498 | 4014.7444 | 3340 | 4140 | 0.18292 | 0.5
I4 | 13 | 2 | 3748.0283 | 3749.8062 | 3740 | 40 | 387.8347 | 380 | 30 | 3475.0107 | 4014.9434 | 3340 | 4140 | 0.18037 | 0.5
I4 | 14 | 2 | 3749.269 | 3748.1091 | 3740 | 40 | 384.6946 | 380 | 30 | 3474.8621 | 4014.5222 | 3340 | 4140 | 0.1825 | 0.5
I4 | 15 | 2 | 3748.1182 | 3750.2703 | 3740 | 40 | 387.5466 | 380 | 30 | 3475.3328 | 4014.9478 | 3340 | 4140 | 0.18016 | 0.5
I4 | 16 | 2 | 3749.1675 | 3748.0137 | 3740 | 40 | 384.6199 | 380 | 30 | 3474.7173 | 4014.2651 | 3340 | 4140 | 0.18007 | 0.5
I4 | 17 | 2 | 3747.8677 | 3749.7639 | 3740 | 40 | 387.4553 | 380 | 30 | 3474.8376 | 4014.6453 | 3340 | 4140 | 0.17886 | 0.5
I4 | 18 | 2 | 3748.8242 | 3747.4902 | 3740 | 40 | 384.8853 | 380 | 30 | 3474.3911 | 4014.0898 | 3340 | 4140 | 0.18154 | 0.5
I4 | 19 | 2 | 3747.8066 | 3749.9424 | 3740 | 40 | 387.2903 | 380 | 30 | 3474.6467 | 4014.4792 | 3340 | 4140 | 0.17945 | 0.5
I4 | 20 | 2 | 3748.686 | 3747.5156 | 3740 | 40 | 384.4966 | 380 | 30 | 3474.3713 | 4013.8572 | 3340 | 4140 | 0.18433 | 0.5
I4 | 21 | 2 | 3747.6646 | 3750.4375 | 3740 | 40 | 387.0266 | 380 | 30 | 3474.6726 | 4014.1636 | 3340 | 4140 | 0.18268 | 0.5
I4 | 22 | 2 | 3748.2329 | 3747.8892 | 3740 | 40 | 384.9377 | 380 | 30 | 3473.8992 | 4013.5415 | 3340 | 4140 | 0.18088 | 0.5
I4 | 23 | 2 | 3746.6553 | 3749.2585 | 3740 | 40 | 387.2039 | 380 | 30 | 3473.6387 | 4013.2703 | 3340 | 4140 | 0.1848 | 0.5
I4 | 24 | 2 | 3747.5337 | 3747.3577 | 3740 | 40 | 384.9062 | 380 | 30 | 3473.4397 | 4012.7593 | 3340 | 4140 | 0.18207 | 0.5
I4 | 25 | 2 | 3745.854 | 3748.8574 | 3740 | 40 | 387.1582 | 380 | 30 | 3472.9075 | 4012.5581 | 3340 | 4140 | 0.18655 | 0.5
I4 | 26 | 2 | 3746.5669 | 3746.7292 | 3740 | 40 | 384.897 | 380 | 30 | 3472.4905 | 4011.5903 | 3340 | 4140 | 0.1853 | 0.5
I4 | 27 | 2 | 3745.1797 | 3748.4768 | 3740 | 40 | 387.1956 | 380 | 30 | 3472.3772 | 4012.0234 | 3340 | 4140 | 0.18722 | 0.5
I4 | 28 | 2 | 3745.9785 | 3746.8999 | 3740 | 40 | 384.9583 | 380 | 30 | 3471.8105 | 4011.0981 | 3340 | 4140 | 0.18854 | 0.5
I4 | 29 | 2 | 3744.5229 | 3748.2864 | 3740 | 40 | 387.3997 | 380 | 30 | 3471.531 | 4011.5562 | 3340 | 4140 | 0.18623 | 0.5
I4 | 30 | 2 | 3745.1912 | 3746.7458 | 3740 | 40 | 385.7881 | 380 | 30 | 3470.9011 | 4010.8379 | 3340 | 4140 | 0.18461 | 0.5
I4 | 31 | 2 | 3743.7266 | 3748.1138 | 3740 | 40 | 387.4314 | 380 | 30 | 3471.0298 | 4011.1484 | 3340 | 4140 | 0.19661 | 0.5
I4 | 32 | 2 | 3744.4746 | 3747.0813 | 3740 | 40 | 385.6829 | 380 | 30 | 3470.1558 | 4010.21 | 3340 | 4140 | 0.18834 | 0.5
M8 | 1 | 1 | 1237.9185 | 1238.0586 | 1240 | 5 | 26.2942 | 20 | 4 | 1213.2947 | 1264.8636 | 1205 | 1275 | 0.53384 | 0.8
M8 | 2 | 1 | 1238.0924 | 1238.2272 | 1240 | 5 | 26.2705 | 20 | 4 | 1213.4189 | 1265.0027 | 1205 | 1275 | 0.53423 | 0.8
M8 | 3 | 1 | 1238.347 | 1238.4939 | 1240 | 5 | 26.1331 | 20 | 4 | 1213.8884 | 1264.9769 | 1205 | 1275 | 0.53633 | 0.8
M8 | 4 | 1 | 1238.4951 | 1238.642 | 1240 | 5 | 26.1593 | 20 | 4 | 1214.0535 | 1265.0863 | 1205 | 1275 | 0.5121 | 0.8
M8 | 5 | 1 | 1238.5981 | 1238.7479 | 1240 | 5 | 26.0197 | 20 | 4 | 1214.3369 | 1265.0142 | 1205 | 1275 | 0.50536 | 0.8
M8 | 6 | 1 | 1238.6477 | 1238.7979 | 1240 | 5 | 26.0607 | 20 | 4 | 1214.3522 | 1265.1154 | 1205 | 1275 | 0.48798 | 0.8
M8 | 7 | 1 | 1238.6697 | 1238.8202 | 1240 | 5 | 25.9479 | 20 | 4 | 1214.4739 | 1265.1298 | 1205 | 1275 | 0.49698 | 0.8
M8 | 8 | 1 | 1238.6725 | 1238.8131 | 1240 | 5 | 26.1084 | 20 | 4 | 1214.292 | 1265.1967 | 1205 | 1275 | 0.4964 | 0.8
M8 | 9 | 1 | 1238.5781 | 1238.7318 | 1240 | 5 | 25.9725 | 20 | 4 | 1214.3661 | 1264.9923 | 1205 | 1275 | 0.47458 | 0.8
M8 | 10 | 1 | 1238.5618 | 1238.7039 | 1240 | 5 | 26.1216 | 20 | 4 | 1214.28 | 1265.0166 | 1205 | 1275 | 0.47811 | 0.8
M8 | 11 | 1 | 1238.4155 | 1238.5581 | 1240 | 5 | 26.0568 | 20 | 4 | 1214.1521 | 1264.8813 | 1205 | 1275 | 0.47368 | 0.8
M8 | 12 | 1 | 1238.3936 | 1238.5319 | 1240 | 5 | 26.1515 | 20 | 4 | 1214.0563 | 1264.9764 | 1205 | 1275 | 0.48867 | 0.8
M8 | 13 | 1 | 1238.1709 | 1238.3141 | 1240 | 5 | 26.1129 | 20 | 4 | 1213.8939 | 1264.7897 | 1205 | 1275 | 0.48236 | 0.8
M8 | 14 | 1 | 1238.1613 | 1238.2898 | 1240 | 5 | 26.1897 | 20 | 4 | 1213.8638 | 1264.8383 | 1205 | 1275 | 0.48335 | 0.8
M8 | 15 | 1 | 1237.8672 | 1238.0092 | 1240 | 5 | 26.1862 | 20 | 4 | 1213.6934 | 1264.3945 | 1205 | 1275 | 0.47193 | 0.8
M8 | 16 | 1 | 1237.8702 | 1237.984 | 1240 | 5 | 26.2603 | 20 | 4 | 1213.6489 | 1264.4966 | 1205 | 1275 | 0.47885 | 0.8
M9 | 1 | 1 | 1374.3567 | 1374.8162 | 1378 | 4 | 13.8153 | 15 | 3 | 1360.9027 | 1388.6504 | 1351 | 1405 | 0.55789 | 1
M9 | 2 | 1 | 1374.3484 | 1374.7412 | 1378 | 4 | 13.9293 | 15 | 3 | 1360.7483 | 1388.728 | 1351 | 1405 | 0.49332 | 1
M9 | 3 | 1 | 1375.5012 | 1375.2428 | 1378 | 4 | 13.8873 | 15 | 3 | 1362.0153 | 1388.7236 | 1351 | 1405 | 0.55085 | 1
M9 | 4 | 1 | 1375.501 | 1375.1921 | 1378 | 4 | 13.9606 | 15 | 3 | 1361.873 | 1388.739 | 1351 | 1405 | 0.51019 | 1
M9 | 5 | 1 | 1375.5999 | 1375.486 | 1378 | 4 | 13.8398 | 15 | 3 | 1362.4271 | 1389.5637 | 1351 | 1405 | 0.46332 | 1
M9 | 6 | 1 | 1375.6877 | 1375.5336 | 1378 | 4 | 13.829 | 15 | 3 | 1362.3652 | 1389.6296 | 1351 | 1405 | 0.45151 | 1
M9 | 7 | 1 | 1375.6494 | 1375.6177 | 1378 | 4 | 13.6826 | 15 | 3 | 1362.5465 | 1389.6908 | 1351 | 1405 | 0.46065 | 1
M9 | 8 | 1 | 1375.726 | 1375.6829 | 1378 | 4 | 13.7146 | 15 | 3 | 1362.4545 | 1389.8652 | 1351 | 1405 | 0.43942 | 1
M9 | 9 | 1 | 1375.6121 | 1375.6241 | 1378 | 4 | 13.6212 | 15 | 3 | 1362.4993 | 1389.7262 | 1351 | 1405 | 0.47382 | 1
M9 | 10 | 1 | 1375.6362 | 1375.6202 | 1378 | 4 | 13.6449 | 15 | 3 | 1362.4043 | 1389.7861 | 1351 | 1405 | 0.44546 | 1
M9 | 11 | 1 | 1375.4795 | 1375.5546 | 1378 | 4 | 13.6477 | 15 | 3 | 1362.3282 | 1389.7598 | 1351 | 1405 | 0.47304 | 1
M9 | 12 | 1 | 1375.4913 | 1375.4998 | 1378 | 4 | 13.6711 | 15 | 3 | 1362.121 | 1389.7996 | 1351 | 1405 | 0.45939 | 1
M9 | 13 | 1 | 1375.3717 | 1375.4435 | 1378 | 4 | 13.5769 | 15 | 3 | 1362.1466 | 1389.7462 | 1351 | 1405 | 0.50973 | 1
M9 | 14 | 1 | 1375.3568 | 1375.3606 | 1378 | 4 | 13.6492 | 15 | 3 | 1361.9607 | 1389.7562 | 1351 | 1405 | 0.47606 | 1
M9 | 15 | 1 | 1375.2997 | 1375.3151 | 1378 | 4 | 13.5278 | 15 | 3 | 1362.0809 | 1389.5646 | 1351 | 1405 | 0.49689 | 1
M9 | 16 | 1 | 1375.2596 | 1375.1578 | 1378 | 4 | 13.5842 | 15 | 3 | 1361.7893 | 1389.571 | 1351 | 1405 | 0.4897 | 1
M10 | 1 | 1 | 1601.769 | 1602.5161 | 1610 | 14 | 59.0892 | 60 | 9 | 1544.0995 | 1664.8958 | 1509 | 1709 | 0.451 | 0.7
M10 | 2 | 1 | 1602.7091 | 1603.4044 | 1610 | 14 | 59.3818 | 60 | 9 | 1544.8408 | 1666.3917 | 1509 | 1709 | 0.46147 | 0.7
M10 | 3 | 1 | 1602.557 | 1603.2874 | 1610 | 14 | 59.3762 | 60 | 9 | 1544.8529 | 1666.1151 | 1509 | 1709 | 0.44684 | 0.7
M10 | 4 | 1 | 1603.3317 | 1604.0308 | 1610 | 14 | 59.7544 | 60 | 9 | 1545.7041 | 1667.1637 | 1509 | 1709 | 0.44929 | 0.7
M10 | 5 | 1 | 1603.0861 | 1603.807 | 1610 | 14 | 59.7554 | 60 | 9 | 1545.6093 | 1666.6471 | 1509 | 1709 | 0.43129 | 0.7
M10 | 6 | 1 | 1603.7621 | 1604.4409 | 1610 | 14 | 59.9514 | 60 | 9 | 1546.1724 | 1667.605 | 1509 | 1709 | 0.45425 | 0.7
M10 | 7 | 1 | 1603.4758 | 1604.1848 | 1610 | 14 | 59.9196 | 60 | 9 | 1546.0175 | 1667.016 | 1509 | 1709 | 0.43909 | 0.7
M10 | 8 | 1 | 1603.9908 | 1604.6857 | 1610 | 14 | 60.2573 | 60 | 9 | 1546.2686 | 1667.9829 | 1509 | 1709 | 0.44519 | 0.7
M10 | 9 | 1 | 1603.7224 | 1604.4376 | 1610 | 14 | 60.1803 | 60 | 9 | 1545.8989 | 1667.6726 | 1509 | 1709 | 0.43164 | 0.7
M10 | 10 | 1 | 1604.2864 | 1604.9857 | 1610 | 14 | 60.2701 | 60 | 9 | 1546.5045 | 1668.1259 | 1509 | 1709 | 0.4429 | 0.7
M10 | 11 | 1 | 1603.9799 | 1604.73 | 1610 | 14 | 60.2534 | 60 | 9 | 1546.1504 | 1667.7394 | 1509 | 1709 | 0.43983 | 0.7
M10 | 12 | 1 | 1604.5459 | 1605.2854 | 1610 | 14 | 60.3646 | 60 | 9 | 1546.6212 | 1668.457 | 1509 | 1709 | 0.43236 | 0.7
M10 | 13 | 1 | 1604.2778 | 1605.0599 | 1610 | 14 | 60.2992 | 60 | 9 | 1546.1135 | 1668.3754 | 1509 | 1709 | 0.45828 | 0.7
M10 | 14 | 1 | 1604.634 | 1605.4059 | 1610 | 14 | 60.4403 | 60 | 9 | 1546.2653 | 1668.88 | 1509 | 1709 | 0.45017 | 0.7
M10 | 15 | 1 | 1604.4617 | 1605.256 | 1610 | 14 | 60.3389 | 60 | 9 | 1546.1249 | 1668.3807 | 1509 | 1709 | 0.43881 | 0.7
M10 | 16 | 1 | 1604.8367 | 1605.6271 | 1610 | 14 | 60.4283 | 60 | 9 | 1546.4893 | 1668.5997 | 1509 | 1709 | 0.43071 | 0.7
M11 | 1 | 1 | 2258.6206 | 2259.4229 | 2250 | 13 | 51.615 | 50 | 6 | 2209.71 | 2315.8391 | 2167 | 2333 | 0.388 | 1
M11 | 2 | 1 | 2258.5469 | 2259.0703 | 2250 | 13 | 52.2263 | 50 | 6 | 2209.4214 | 2315.0391 | 2167 | 2333 | 0.36708 | 1
M11 | 3 | 1 | 2259.0527 | 2259.8352 | 2250 | 13 | 51.7224 | 50 | 6 | 2210.1899 | 2316.1372 | 2167 | 2333 | 0.38747 | 1
M11 | 4 | 1 | 2259.0142 | 2259.7061 | 2250 | 13 | 51.9226 | 50 | 6 | 2210.229 | 2315.9229 | 2167 | 2333 | 0.37135 | 1
M11 | 5 | 1 | 2258.6558 | 2259.1765 | 2250 | 13 | 52.0078 | 50 | 6 | 2209.9382 | 2314.1992 | 2167 | 2333 | 0.37257 | 1
M11 | 6 | 1 | 2258.6851 | 2259.1365 | 2250 | 13 | 52.1895 | 50 | 6 | 2209.8013 | 2314.3008 | 2167 | 2333 | 0.36554 | 1
M11 | 7 | 1 | 2258.9102 | 2259.4172 | 2250 | 13 | 51.9666 | 50 | 6 | 2210.3242 | 2314.4949 | 2167 | 2333 | 0.37349 | 1
M11 | 8 | 1 | 2259.0107 | 2259.7461 | 2250 | 13 | 51.9023 | 50 | 6 | 2210.3096 | 2315.5542 | 2167 | 2333 | 0.37221 | 1
M11 | 9 | 1 | 2258.5732 | 2259.1528 | 2250 | 13 | 51.9739 | 50 | 6 | 2210.0342 | 2314.2363 | 2167 | 2333 | 0.37574 | 1
M11 | 10 | 1 | 2258.4692 | 2259.0447 | 2250 | 13 | 52.0652 | 50 | 6 | 2209.8625 | 2314.51 | 2167 | 2333 | 0.37411 | 1
M11 | 11 | 1 | 2257.8726 | 2258.4409 | 2250 | 13 | 51.8367 | 50 | 6 | 2209.3308 | 2313.3794 | 2167 | 2333 | 0.37879 | 1
M11 | 12 | 1 | 2257.4805 | 2258.022 | 2250 | 13 | 52.0459 | 50 | 6 | 2208.6462 | 2313.2585 | 2167 | 2333 | 0.36601 | 1
M11 | 13 | 1 | 2257.2065 | 2257.7488 | 2250 | 13 | 52.0803 | 50 | 6 | 2208.3921 | 2312.8274 | 2167 | 2333 | 0.37553 | 1
M11 | 14 | 1 | 2257.158 | 2257.7236 | 2250 | 13 | 52.2461 | 50 | 6 | 2208.1465 | 2313.2725 | 2167 | 2333 | 0.36436 | 1
M11 | 15 | 1 | 2256.856 | 2257.5005 | 2250 | 13 | 52.011 | 50 | 6 | 2208.2305 | 2312.7847 | 2167 | 2333 | 0.36434 | 1
M11 | 16 | 1 | 2256.9275 | 2257.606 | 2250 | 13 | 52.0576 | 50 | 6 | 2208.3108 | 2313.4292 | 2167 | 2333 | 0.36408 | 1
M12 | 1 | 1 | 3697.0283 | 3698.4153 | 3700 | 32 | 196.1765 | 180 | 20 | 3516.8345 | 3893.3418 | 3410 | 3990 | 0.37653 | 1.1
M12 | 2 | 1 | 3698.0942 | 3699.3574 | 3700 | 32 | 195.9897 | 180 | 20 | 3517.1724 | 3894.3411 | 3410 | 3990 | 0.36325 | 1.1
M12 | 3 | 1 | 3697.7937 | 3699.2024 | 3700 | 32 | 195.8032 | 180 | 20 | 3517.5659 | 3893.9512 | 3410 | 3990 | 0.37344 | 1.1
M12 | 4 | 1 | 3698.5212 | 3699.7725 | 3700 | 32 | 195.7534 | 180 | 20 | 3518.0972 | 3894.1843 | 3410 | 3990 | 0.35885 | 1.1
M12 | 5 | 1 | 3698.1497 | 3699.5698 | 3700 | 32 | 195.5063 | 180 | 20 | 3518.8242 | 3893.0156 | 3410 | 3990 | 0.36516 | 1.1
M12 | 6 | 1 | 3698.647 | 3700.2324 | 3700 | 32 | 195.4045 | 180 | 20 | 3519.6887 | 3893.5239 | 3410 | 3990 | 0.35476 | 1.1
M12 | 7 | 1 | 3698.1379 | 3699.8506 | 3700 | 32 | 194.8042 | 180 | 20 | 3519.6328 | 3892.9705 | 3410 | 3990 | 0.37659 | 1.1
M12 | 8 | 1 | 3698.9702 | 3700.2834 | 3700 | 32 | 195.4783 | 180 | 20 | 3518.6204 | 3893.9868 | 3410 | 3990 | 0.35386 | 1.1
M12 | 9 | 1 | 3697.8672 | 3699.6252 | 3700 | 32 | 194.873 | 180 | 20 | 3519.5852 | 3892.429 | 3410 | 3990 | 0.37704 | 1.1
M12 | 10 | 1 | 3698.5151 | 3699.8774 | 3700 | 32 | 195.075 | 180 | 20 | 3518.0586 | 3893.4888 | 3410 | 3990 | 0.35108 | 1.1
M12 | 11 | 1 | 3697.2573 | 3699.2373 | 3700 | 32 | 194.5486 | 180 | 20 | 3519.303 | 3891.8525 | 3410 | 3990 | 0.3787 | 1.1
M12 | 12 | 1 | 3698.0063 | 3699.3508 | 3700 | 32 | 195.3367 | 180 | 20 | 3517.6633 | 3892.7114 | 3410 | 3990 | 0.35166 | 1.1
M12 | 13 | 1 | 3696.7749 | 3698.5835 | 3700 | 32 | 194.7278 | 180 | 20 | 3517.5935 | 3891.6748 | 3410 | 3990 | 0.3714 | 1.1
M12 | 14 | 1 | 3696.9761 | 3698.7681 | 3700 | 32 | 195.3167 | 180 | 20 | 3516.6626 | 3892.6624 | 3410 | 3990 | 0.35686 | 1.1
M12 | 15 | 1 | 3695.8049 | 3697.856 | 3700 | 32 | 194.8081 | 180 | 20 | 3516.8416 | 3891.0425 | 3410 | 3990 | 0.38526 | 1.1
M12 | 16 | 1 | 3696.2476 | 3698.0059 | 3700 | 32 | 195.5491 | 180 | 20 | 3515.8513 | 3892.2795 | 3410 | 3990 | 0.36573 | 1.1
M13 | 1 | 1 | 4070.8801 | 4070.9961 | 4050 | 34 | 155.6606 | 155 | 20 | 3908.2373 | 4217.6714 | 3790 | 4310 | 0.76271 | 1.3
M13 | 2 | 1 | 4070.874 | 4070.9875 | 4050 | 34 | 155.2278 | 155 | 20 | 3908.3875 | 4218.1753 | 3790 | 4310 | 0.72944 | 1.3
M13 | 3 | 1 | 4071.6396 | 4071.803 | 4050 | 34 | 155.3145 | 155 | 20 | 3909.2791 | 4217.4297 | 3790 | 4310 | 0.78198 | 1.3
M13 | 4 | 1 | 4071.0708 | 4070.9778 | 4050 | 34 | 155.135 | 155 | 20 | 3908.9087 | 4217.5049 | 3790 | 4310 | 0.71431 | 1.3
M13 | 5 | 1 | 4072.7798 | 4072.8286 | 4050 | 34 | 154.8604 | 155 | 20 | 3910.9558 | 4216.7622 | 3790 | 4310 | 0.79301 | 1.3
M13 | 6 | 1 | 4070.6555 | 4070.6936 | 4050 | 34 | 154.269 | 155 | 20 | 3908.9304 | 4216.5068 | 3790 | 4310 | 0.70256 | 1.3
M13 | 7 | 1 | 4071.1284 | 4071.2119 | 4050 | 34 | 154.4788 | 155 | 20 | 3909.2463 | 4215.5312 | 3790 | 4310 | 0.76159 | 1.3
M13 | 8 | 1 | 4069.9536 | 4069.9583 | 4050 | 34 | 153.9092 | 155 | 20 | 3908.428 | 4215.6685 | 3790 | 4310 | 0.69908 | 1.3
M13 | 9 | 1 | 4070.4597 | 4070.4739 | 4050 | 34 | 154.3882 | 155 | 20 | 3908.438 | 4214.9189 | 3790 | 4310 | 0.87978 | 1.3
M13 | 10 | 1 | 4069.1724 | 4069.084 | 4050 | 34 | 153.8379 | 155 | 20 | 3907.6816 | 4214.6709 | 3790 | 4310 | 0.76759 | 1.3
M13 | 11 | 1 | 4069.354 | 4069.3845 | 4050 | 34 | 154.2568 | 155 | 20 | 3907.6072 | 4214.3013 | 3790 | 4310 | 0.88247 | 1.3
M13 | 12 | 1 | 4068.2966 | 4068.0354 | 4050 | 34 | 154.3013 | 155 | 20 | 3906.5105 | 4213.9351 | 3790 | 4310 | 0.75446 | 1.3
M13 | 13 | 1 | 4068.6807 | 4068.4099 | 4050 | 34 | 154.7498 | 155 | 20 | 3906.3994 | 4213.5796 | 3790 | 4310 | 0.82466 | 1.3
M13 | 14 | 1 | 4067.2795 | 4066.9866 | 4050 | 34 | 154.1919 | 155 | 20 | 3905.1377 | 4213.5576 | 3790 | 4310 | 0.75052 | 1.3
M13 | 15 | 1 | 4067.2922 | 4067.2654 | 4050 | 34 | 154.9263 | 155 | 20 | 3905.0818 | 4212.7847 | 3790 | 4310 | 0.86558 | 1.3
M13 | 16 | 1 | 4066.3755 | 4066.0881 | 4050 | 34 | 154.6799 | 155 | 20 | 3904.1262 | 4213.1128 | 3790 | 4310 | 0.7397 | 1.3

## LWIR
1. This table provides key J1 VIIRS spectral performance metrics (band median).
2. The measurements were taken during FP-15 and FP-16 testing using the SpMA source. | [8 empty] | Non-compliance with specification
3. The LWIR bands were measured duirng TV testing at the nominal iinstrument temperature plateau.
4. Center, bandwidth, and 1% limits are taken form FP-15 measurements.
5. Odd and even detector arrays were measured separately for the LWIR bands.
Band | Detector | Subsample | Center Wavelength (nm) |  |  |  | Bandwidth (nm) |  |  | 1 % Limits (nm) |  |  |  | IOOB (%)
|  |  | Measurement |  | Specification |  | Measurement | Specification |  | Measurement |  | Specification |  | Measurement | Specification
|  |  | Center | Centroid | Center | Tolerance |  | Bandwidth | Tolerance | Lower Limit | Upper Limit | Lower Limit | Upper Limit
I5 | 1 | 1 | 11460.7178 | 11469.2344 | 11450 | 125 | 1885.2363 | 1900 | 100 | 10099.8135 | 13105.2979 | 9900 | 12900 | 0.28681 | 0.4
I5 | 2 | 1 | 11471.084 | 11466.6816 | 11450 | 125 | 1821.709 | 1900 | 100 | 10209.3066 | 13048.2637 | 9900 | 12900 | 0.35865 | 0.4
I5 | 3 | 1 | 11461.5371 | 11472.7998 | 11450 | 125 | 1881.542 | 1900 | 100 | 10122.5312 | 13117.6338 | 9900 | 12900 | 0.27848 | 0.4
I5 | 4 | 1 | 11475.7344 | 11470.5811 | 11450 | 125 | 1827.7939 | 1900 | 100 | 10207.0127 | 13062.7217 | 9900 | 12900 | 0.31281 | 0.4
I5 | 5 | 1 | 11468.5352 | 11478.0752 | 11450 | 125 | 1889.4248 | 1900 | 100 | 10114.9805 | 13113.7676 | 9900 | 12900 | 0.25373 | 0.4
I5 | 6 | 1 | 11483.3076 | 11475.7881 | 11450 | 125 | 1839.0293 | 1900 | 100 | 10203.5645 | 13026.793 | 9900 | 12900 | 0.31444 | 0.4
I5 | 7 | 1 | 11475.2988 | 11481.7031 | 11450 | 125 | 1894.3955 | 1900 | 100 | 10116.4531 | 13111.5693 | 9900 | 12900 | 0.28911 | 0.4
I5 | 8 | 1 | 11484.6152 | 11476.8291 | 11450 | 125 | 1834.5127 | 1900 | 100 | 10205.4053 | 13031.9209 | 9900 | 12900 | 0.32673 | 0.4
I5 | 9 | 1 | 11473.5859 | 11481.9971 | 11450 | 125 | 1883.4941 | 1900 | 100 | 10135.7402 | 13113.5459 | 9900 | 12900 | 0.43171 | 0.4
I5 | 10 | 1 | 11477.9941 | 11474.4043 | 11450 | 125 | 1817.0635 | 1900 | 100 | 10209.4268 | 13019.7598 | 9900 | 12900 | 0.40039 | 0.4
I5 | 11 | 1 | 11474.3359 | 11481.9277 | 11450 | 125 | 1885.2979 | 1900 | 100 | 10129.8369 | 13109.918 | 9900 | 12900 | 0.26247 | 0.4
I5 | 12 | 1 | 11481.0273 | 11476.9541 | 11450 | 125 | 1821.9971 | 1900 | 100 | 10211.5527 | 13045.3418 | 9900 | 12900 | 0.28378 | 0.4
I5 | 13 | 1 | 11483.2715 | 11487.502 | 11450 | 125 | 1903.8965 | 1900 | 100 | 10115.3555 | 13111.9375 | 9900 | 12900 | 0.28579 | 0.4
I5 | 14 | 1 | 11485.1328 | 11479.7744 | 11450 | 125 | 1832.0244 | 1900 | 100 | 10206.7178 | 13050.9268 | 9900 | 12900 | 0.27244 | 0.4
I5 | 15 | 1 | 11488.8848 | 11490.7812 | 11450 | 125 | 1917.1924 | 1900 | 100 | 10120.9912 | 13102.4922 | 9900 | 12900 | 0.23793 | 0.4
I5 | 16 | 1 | 11488.8564 | 11483.5322 | 11450 | 125 | 1840.125 | 1900 | 100 | 10214.9209 | 13060.5176 | 9900 | 12900 | 0.37122 | 0.4
I5 | 17 | 1 | 11490.3027 | 11491.6445 | 11450 | 125 | 1922.1123 | 1900 | 100 | 10123.3604 | 13112.6885 | 9900 | 12900 | 0.27068 | 0.4
I5 | 18 | 1 | 11489.2578 | 11484.8477 | 11450 | 125 | 1838.4932 | 1900 | 100 | 10211.0049 | 13038.2559 | 9900 | 12900 | 0.24352 | 0.4
I5 | 19 | 1 | 11487.9023 | 11495.2207 | 11450 | 125 | 1912.4111 | 1900 | 100 | 10117.7324 | 13133.7881 | 9900 | 12900 | 0.2956 | 0.4
I5 | 20 | 1 | 11493.3613 | 11489.0312 | 11450 | 125 | 1849.5186 | 1900 | 100 | 10209.7568 | 13043.5566 | 9900 | 12900 | 0.24175 | 0.4
I5 | 21 | 1 | 11486.5449 | 11494.0312 | 11450 | 125 | 1912.1201 | 1900 | 100 | 10119.6289 | 13114.3633 | 9900 | 12900 | 0.25768 | 0.4
I5 | 22 | 1 | 11489.5283 | 11485.9092 | 11450 | 125 | 1846.0781 | 1900 | 100 | 10210.7588 | 13041.0938 | 9900 | 12900 | 0.2811 | 0.4
I5 | 23 | 1 | 11485.9824 | 11492.2383 | 11450 | 125 | 1917.3271 | 1900 | 100 | 10119.8135 | 13117.6416 | 9900 | 12900 | 0.23626 | 0.4
I5 | 24 | 1 | 11487.4844 | 11484.2773 | 11450 | 125 | 1844.3096 | 1900 | 100 | 10208.665 | 13024.9609 | 9900 | 12900 | 0.44577 | 0.4
I5 | 25 | 1 | 11484.6885 | 11490.3691 | 11450 | 125 | 1916.8047 | 1900 | 100 | 10116.7637 | 13107.415 | 9900 | 12900 | 0.26537 | 0.4
I5 | 26 | 1 | 11485.6406 | 11482.1084 | 11450 | 125 | 1847.3896 | 1900 | 100 | 10208.9736 | 13041.6895 | 9900 | 12900 | 0.2789 | 0.4
I5 | 27 | 1 | 11485.8555 | 11488.6055 | 11450 | 125 | 1930.502 | 1900 | 100 | 10118.0664 | 13124.3262 | 9900 | 12900 | 0.26241 | 0.4
I5 | 28 | 1 | 11485.0625 | 11481.7178 | 11450 | 125 | 1848.4814 | 1900 | 100 | 10209.8418 | 13070.6211 | 9900 | 12900 | 0.39703 | 0.4
I5 | 29 | 1 | 11480.8301 | 11485.9824 | 11450 | 125 | 1921.251 | 1900 | 100 | 10111.6699 | 13113.3486 | 9900 | 12900 | 0.36674 | 0.4
I5 | 30 | 1 | 11484.4414 | 11478.4502 | 11450 | 125 | 1853.7822 | 1900 | 100 | 10199.1455 | 13021.3135 | 9900 | 12900 | 0.27978 | 0.4
I5 | 31 | 1 | 11482.2666 | 11481.4658 | 11450 | 125 | 1929.5156 | 1900 | 100 | 10107.71 | 13106.9785 | 9900 | 12900 | 0.30775 | 0.4
I5 | 32 | 1 | 11481.8643 | 11474.4932 | 11450 | 125 | 1850.8457 | 1900 | 100 | 10202.5625 | 13055.7197 | 9900 | 12900 | 0.34815 | 0.4
I5 | 1 | 2 | 11460.6582 | 11469.3008 | 11450 | 125 | 1885.1338 | 1900 | 100 | 10097.25 | 13106.2314 | 9900 | 12900 | 0.28788 | 0.4
I5 | 2 | 2 | 11471.2051 | 11466.7812 | 11450 | 125 | 1821.9971 | 1900 | 100 | 10208.8721 | 13049.6289 | 9900 | 12900 | 0.359 | 0.4
I5 | 3 | 2 | 11461.5566 | 11472.8838 | 11450 | 125 | 1881.5557 | 1900 | 100 | 10120.5732 | 13117.1865 | 9900 | 12900 | 0.2758 | 0.4
I5 | 4 | 2 | 11475.6953 | 11470.6943 | 11450 | 125 | 1827.6924 | 1900 | 100 | 10206.9209 | 13063.2217 | 9900 | 12900 | 0.31588 | 0.4
I5 | 5 | 2 | 11468.498 | 11478.0088 | 11450 | 125 | 1889.3525 | 1900 | 100 | 10112.3125 | 13113.7119 | 9900 | 12900 | 0.25492 | 0.4
I5 | 6 | 2 | 11483.2422 | 11475.8457 | 11450 | 125 | 1838.8916 | 1900 | 100 | 10203.7803 | 13027.5107 | 9900 | 12900 | 0.31426 | 0.4
I5 | 7 | 2 | 11475.3535 | 11481.6914 | 11450 | 125 | 1894.499 | 1900 | 100 | 10115.1523 | 13109.9209 | 9900 | 12900 | 0.28803 | 0.4
I5 | 8 | 2 | 11484.5586 | 11476.9609 | 11450 | 125 | 1834.4658 | 1900 | 100 | 10205.5176 | 13033.9014 | 9900 | 12900 | 0.32061 | 0.4
I5 | 9 | 2 | 11473.502 | 11482.0576 | 11450 | 125 | 1883.3096 | 1900 | 100 | 10135.2305 | 13115.1416 | 9900 | 12900 | 0.42909 | 0.4
I5 | 10 | 2 | 11478.0449 | 11474.4268 | 11450 | 125 | 1817.1611 | 1900 | 100 | 10209.2783 | 13019.8662 | 9900 | 12900 | 0.40366 | 0.4
I5 | 11 | 2 | 11474.334 | 11481.9424 | 11450 | 125 | 1885.249 | 1900 | 100 | 10129.1318 | 13111.9268 | 9900 | 12900 | 0.25863 | 0.4
I5 | 12 | 2 | 11481.0527 | 11477.1484 | 11450 | 125 | 1822.0518 | 1900 | 100 | 10211.4893 | 13049.3164 | 9900 | 12900 | 0.28399 | 0.4
I5 | 13 | 2 | 11483.2285 | 11487.5127 | 11450 | 125 | 1903.8896 | 1900 | 100 | 10115.9258 | 13112.0771 | 9900 | 12900 | 0.28649 | 0.4
I5 | 14 | 2 | 11485.1768 | 11479.752 | 11450 | 125 | 1832.1562 | 1900 | 100 | 10206.6826 | 13050.126 | 9900 | 12900 | 0.27742 | 0.4
I5 | 15 | 2 | 11488.9805 | 11490.8721 | 11450 | 125 | 1917.4424 | 1900 | 100 | 10120.4248 | 13103.1621 | 9900 | 12900 | 0.24035 | 0.4
I5 | 16 | 2 | 11488.957 | 11483.4785 | 11450 | 125 | 1840.4346 | 1900 | 100 | 10214.4746 | 13061.0947 | 9900 | 12900 | 0.37227 | 0.4
I5 | 17 | 2 | 11490.1992 | 11491.7129 | 11450 | 125 | 1921.8857 | 1900 | 100 | 10122.5146 | 13113.9043 | 9900 | 12900 | 0.26943 | 0.4
I5 | 18 | 2 | 11489.3047 | 11484.8174 | 11450 | 125 | 1838.6133 | 1900 | 100 | 10211.0762 | 13037.5107 | 9900 | 12900 | 0.2454 | 0.4
I5 | 19 | 2 | 11487.9082 | 11495.3193 | 11450 | 125 | 1912.377 | 1900 | 100 | 10118.084 | 13136.3223 | 9900 | 12900 | 0.29628 | 0.4
I5 | 20 | 2 | 11493.3574 | 11488.9873 | 11450 | 125 | 1849.4951 | 1900 | 100 | 10208.8809 | 13043.1572 | 9900 | 12900 | 0.2418 | 0.4
I5 | 21 | 2 | 11486.4844 | 11494.2246 | 11450 | 125 | 1911.9502 | 1900 | 100 | 10118.6074 | 13116.8096 | 9900 | 12900 | 0.26638 | 0.4
I5 | 22 | 2 | 11489.5508 | 11485.8994 | 11450 | 125 | 1846.166 | 1900 | 100 | 10210.9092 | 13040.2871 | 9900 | 12900 | 0.28223 | 0.4
I5 | 23 | 2 | 11486.0547 | 11492.3574 | 11450 | 125 | 1917.501 | 1900 | 100 | 10120.0088 | 13118.3916 | 9900 | 12900 | 0.23684 | 0.4
I5 | 24 | 2 | 11487.5469 | 11484.3057 | 11450 | 125 | 1844.4746 | 1900 | 100 | 10208.7051 | 13024.5518 | 9900 | 12900 | 0.44599 | 0.4
I5 | 25 | 2 | 11484.6699 | 11490.4775 | 11450 | 125 | 1916.7256 | 1900 | 100 | 10116.3555 | 13107.4785 | 9900 | 12900 | 0.2619 | 0.4
I5 | 26 | 2 | 11485.6777 | 11482.1084 | 11450 | 125 | 1847.4639 | 1900 | 100 | 10209.4053 | 13042.21 | 9900 | 12900 | 0.28107 | 0.4
I5 | 27 | 2 | 11485.7344 | 11488.6924 | 11450 | 125 | 1930.2178 | 1900 | 100 | 10117.3164 | 13124.7959 | 9900 | 12900 | 0.26342 | 0.4
I5 | 28 | 2 | 11485.1348 | 11481.7266 | 11450 | 125 | 1848.624 | 1900 | 100 | 10210.5938 | 13067.2393 | 9900 | 12900 | 0.40075 | 0.4
I5 | 29 | 2 | 11480.6123 | 11485.9814 | 11450 | 125 | 1920.7227 | 1900 | 100 | 10109.8252 | 13112.6826 | 9900 | 12900 | 0.3654 | 0.4
I5 | 30 | 2 | 11484.5176 | 11478.4551 | 11450 | 125 | 1853.915 | 1900 | 100 | 10199.2041 | 13021.5928 | 9900 | 12900 | 0.27852 | 0.4
I5 | 31 | 2 | 11482.2246 | 11481.377 | 11450 | 125 | 1929.4326 | 1900 | 100 | 10106.0908 | 13106.3184 | 9900 | 12900 | 0.30391 | 0.4
I5 | 32 | 2 | 11481.9775 | 11474.6611 | 11450 | 125 | 1851.043 | 1900 | 100 | 10201.8291 | 13057.4756 | 9900 | 12900 | 0.35093 | 0.4
M14 | 1 | 1 | 8580.2578 | 8583.8096 | 8550 | 70 | 338.9014 | 300 | 40 | 8337.2373 | 8878.792 | 8050 | 9050 | 0.23891 | 0.9
M14 | 2 | 1 | 8576.3086 | 8579.6357 | 8550 | 70 | 337.8936 | 300 | 40 | 8335.4316 | 8873.957 | 8050 | 9050 | 0.24805 | 0.9
M14 | 3 | 1 | 8583.2852 | 8586.5244 | 8550 | 70 | 338.915 | 300 | 40 | 8338.7676 | 8881.417 | 8050 | 9050 | 0.22902 | 0.9
M14 | 4 | 1 | 8578.832 | 8582.2988 | 8550 | 70 | 338.8076 | 300 | 40 | 8336.5371 | 8876.6016 | 8050 | 9050 | 0.25444 | 0.9
M14 | 5 | 1 | 8585.0625 | 8588.0791 | 8550 | 70 | 339.46 | 300 | 40 | 8339.5332 | 8883.7676 | 8050 | 9050 | 0.22546 | 0.9
M14 | 6 | 1 | 8580.1328 | 8583.6826 | 8550 | 70 | 339.4736 | 300 | 40 | 8337.0176 | 8878.0049 | 8050 | 9050 | 0.23879 | 0.9
M14 | 7 | 1 | 8585.3018 | 8588.376 | 8550 | 70 | 339.5293 | 300 | 40 | 8339.6299 | 8884.9395 | 8050 | 9050 | 0.22098 | 0.9
M14 | 8 | 1 | 8580.002 | 8583.4229 | 8550 | 70 | 339.3779 | 300 | 40 | 8336.6855 | 8877.1826 | 8050 | 9050 | 0.24384 | 0.9
M14 | 9 | 1 | 8584.7861 | 8587.8994 | 8550 | 70 | 339.9434 | 300 | 40 | 8339.2051 | 8884.8105 | 8050 | 9050 | 0.24879 | 0.9
M14 | 10 | 1 | 8579.3057 | 8582.707 | 8550 | 70 | 340.5586 | 300 | 40 | 8335.7598 | 8876.9404 | 8050 | 9050 | 0.24119 | 0.9
M14 | 11 | 1 | 8583.3789 | 8586.6846 | 8550 | 70 | 339.9014 | 300 | 40 | 8337.9053 | 8883.999 | 8050 | 9050 | 0.22534 | 0.9
M14 | 12 | 1 | 8577.6465 | 8581.1025 | 8550 | 70 | 340.6035 | 300 | 40 | 8334.7383 | 8875.7324 | 8050 | 9050 | 0.24531 | 0.9
M14 | 13 | 1 | 8581.1514 | 8584.2686 | 8550 | 70 | 339.0293 | 300 | 40 | 8336.833 | 8880.7334 | 8050 | 9050 | 0.23597 | 0.9
M14 | 14 | 1 | 8575.2656 | 8578.4746 | 8550 | 70 | 339.1016 | 300 | 40 | 8332.5605 | 8873.6836 | 8050 | 9050 | 0.2564 | 0.9
M14 | 15 | 1 | 8578.4326 | 8581.5938 | 8550 | 70 | 338.8652 | 300 | 40 | 8335.3662 | 8876.7646 | 8050 | 9050 | 0.23518 | 0.9
M14 | 16 | 1 | 8572.5703 | 8575.5635 | 8550 | 70 | 338.5605 | 300 | 40 | 8328.8955 | 8870.4893 | 8050 | 9050 | 0.27235 | 0.9
M15 | 1 | 1 | 10723.4531 | 10711.1787 | 10763 | 113 | 1014.1494 | 1000 | 100 | 9887.7354 | 11634.0479 | 9700 | 11740 | 0.42061 | 0.4
M15 | 2 | 1 | 10727.6641 | 10715.4824 | 10763 | 113 | 1012.417 | 1000 | 100 | 9892.2637 | 11624.1865 | 9700 | 11740 | 0.41861 | 0.4
M15 | 3 | 1 | 10728.248 | 10717.0322 | 10763 | 113 | 1010.126 | 1000 | 100 | 9902.2852 | 11639.835 | 9700 | 11740 | 0.42884 | 0.4
M15 | 4 | 1 | 10731.5449 | 10720.1973 | 10763 | 113 | 1007.6504 | 1000 | 100 | 9909.7607 | 11631.1641 | 9700 | 11740 | 0.40988 | 0.4
M15 | 5 | 1 | 10730.8711 | 10720.7607 | 10763 | 113 | 998.6689 | 1000 | 100 | 9919.4756 | 11644.4336 | 9700 | 11740 | 0.424 | 0.4
M15 | 6 | 1 | 10733.4639 | 10722.5234 | 10763 | 113 | 999.8008 | 1000 | 100 | 9922.9102 | 11634.5771 | 9700 | 11740 | 0.41906 | 0.4
M15 | 7 | 1 | 10733.7695 | 10724.0254 | 10763 | 113 | 1002.8975 | 1000 | 100 | 9923.0654 | 11645.4521 | 9700 | 11740 | 0.42471 | 0.4
M15 | 8 | 1 | 10735.8848 | 10725.4209 | 10763 | 113 | 1004.2412 | 1000 | 100 | 9925.6328 | 11636.0176 | 9700 | 11740 | 0.41759 | 0.4
M15 | 9 | 1 | 10734.3184 | 10724.5166 | 10763 | 113 | 1002.249 | 1000 | 100 | 9924.0039 | 11645.3838 | 9700 | 11740 | 0.42011 | 0.4
M15 | 10 | 1 | 10735.3086 | 10724.9062 | 10763 | 113 | 998.3291 | 1000 | 100 | 9926.041 | 11636.7715 | 9700 | 11740 | 0.41424 | 0.4
M15 | 11 | 1 | 10733.1836 | 10723.2803 | 10763 | 113 | 999.2695 | 1000 | 100 | 9922.5039 | 11644.8379 | 9700 | 11740 | 0.42151 | 0.4
M15 | 12 | 1 | 10733.9746 | 10723.5225 | 10763 | 113 | 997.9268 | 1000 | 100 | 9923.915 | 11634.6631 | 9700 | 11740 | 0.4134 | 0.4
M15 | 13 | 1 | 10729.6914 | 10723.6006 | 10763 | 113 | 990.1074 | 1000 | 100 | 9918.7871 | 11646.7559 | 9700 | 11740 | 0.42298 | 0.4
M15 | 14 | 1 | 10730.1406 | 10722.6543 | 10763 | 113 | 988.0713 | 1000 | 100 | 9920.2598 | 11635.4102 | 9700 | 11740 | 0.41962 | 0.4
M15 | 15 | 1 | 10726.2236 | 10721.4004 | 10763 | 113 | 991.4375 | 1000 | 100 | 9909.0449 | 11644.2246 | 9700 | 11740 | 0.43328 | 0.4
M15 | 16 | 1 | 10726.7207 | 10720.8467 | 10763 | 113 | 988.4453 | 1000 | 100 | 9912.5127 | 11631.0518 | 9700 | 11740 | 0.42367 | 0.4
M16A | 1 | 1 | 11875.3633 | 11870.0742 | 12013 | 88 | 915.126 | 950 | 50 | 11101.4414 | 12683.7012 | 11060 | 13050 | 0.66609 | 0.4
M16A | 2 | 1 | 11876.2695 | 11870.5215 | 12013 | 88 | 923.5762 | 950 | 50 | 11098.2402 | 12684.4951 | 11060 | 13050 | 0.63403 | 0.4
M16A | 3 | 1 | 11880.8945 | 11875.4287 | 12013 | 88 | 914.4248 | 950 | 50 | 11105.0322 | 12688.4893 | 11060 | 13050 | 0.55786 | 0.4
M16A | 4 | 1 | 11879.6211 | 11873.9004 | 12013 | 88 | 922.9053 | 950 | 50 | 11099.7705 | 12690.1211 | 11060 | 13050 | 0.55418 | 0.4
M16A | 5 | 1 | 11881.1299 | 11874.6514 | 12013 | 88 | 900.3535 | 950 | 50 | 11109.999 | 12689.5879 | 11060 | 13050 | 0.61393 | 0.4
M16A | 6 | 1 | 11883.3691 | 11878.5576 | 12013 | 88 | 922.8203 | 950 | 50 | 11101.166 | 12692.1533 | 11060 | 13050 | 0.57517 | 0.4
M16A | 7 | 1 | 11883.5527 | 11876.4404 | 12013 | 88 | 899 | 950 | 50 | 11112.0137 | 12696.7949 | 11060 | 13050 | 0.54083 | 0.4
M16A | 8 | 1 | 11885.3242 | 11881.207 | 12013 | 88 | 914.7158 | 950 | 50 | 11103.8545 | 12696.082 | 11060 | 13050 | 0.56067 | 0.4
M16A | 9 | 1 | 11886.2656 | 11880.4717 | 12013 | 88 | 900.5771 | 950 | 50 | 11110.9941 | 12697.7139 | 11060 | 13050 | 0.5372 | 0.4
M16A | 10 | 1 | 11886.0625 | 11881.9795 | 12013 | 88 | 927.4102 | 950 | 50 | 11102.1299 | 12698.2041 | 11060 | 13050 | 0.53636 | 0.4
M16A | 11 | 1 | 11888.4414 | 11884.498 | 12013 | 88 | 911.9561 | 950 | 50 | 11108.6016 | 12696.0703 | 11060 | 13050 | 0.55282 | 0.4
M16A | 12 | 1 | 11884.7207 | 11880.4932 | 12013 | 88 | 928.5283 | 950 | 50 | 11098.6396 | 12695.3906 | 11060 | 13050 | 0.62551 | 0.4
M16A | 13 | 1 | 11883.832 | 11878.5557 | 12013 | 88 | 902.5166 | 950 | 50 | 11107.8125 | 12693.4717 | 11060 | 13050 | 0.56714 | 0.4
M16A | 14 | 1 | 11882.2471 | 11877.6826 | 12013 | 88 | 932.9824 | 950 | 50 | 11099.2715 | 12693.1934 | 11060 | 13050 | 0.56545 | 0.4
M16A | 15 | 1 | 11879.2109 | 11873.1934 | 12013 | 88 | 904.3232 | 950 | 50 | 11105.4541 | 12691.5176 | 11060 | 13050 | 0.60449 | 0.4
M16A | 16 | 1 | 11875.873 | 11871.8438 | 12013 | 88 | 916.4492 | 950 | 50 | 11097.9971 | 12684.0264 | 11060 | 13050 | 0.62926 | 0.4
M16B | 1 | 1 | 11875.0566 | 11868.0986 | 12013 | 88 | 934.6709 | 950 | 50 | 11097.4609 | 12687.8682 | 11060 | 13050 | 0.6201 | 0.4
M16B | 2 | 1 | 11875.5605 | 11864.2598 | 12013 | 88 | 933.6846 | 950 | 50 | 11099.2676 | 12698.9365 | 11060 | 13050 | 0.59608 | 0.4
M16B | 3 | 1 | 11879.6699 | 11871.7793 | 12013 | 88 | 941.5801 | 950 | 50 | 11099.3291 | 12694.5479 | 11060 | 13050 | 0.58057 | 0.4
M16B | 4 | 1 | 11880.3789 | 11870.0381 | 12013 | 88 | 931.627 | 950 | 50 | 11102.165 | 12699.4854 | 11060 | 13050 | 0.59864 | 0.4
M16B | 5 | 1 | 11883.9805 | 11877.0771 | 12013 | 88 | 935.6748 | 950 | 50 | 11098.0156 | 12694.2715 | 11060 | 13050 | 0.92991 | 0.4
M16B | 6 | 1 | 11884.3828 | 11874.7227 | 12013 | 88 | 935.7979 | 950 | 50 | 11104.3291 | 12701.2773 | 11060 | 13050 | 0.54034 | 0.4
M16B | 7 | 1 | 11887.2422 | 11882.3594 | 12013 | 88 | 929.7129 | 950 | 50 | 11102.1201 | 12700.3115 | 11060 | 13050 | 0.52296 | 0.4
M16B | 8 | 1 | 11887.1094 | 11880.375 | 12013 | 88 | 941.2959 | 950 | 50 | 11103.7822 | 12703.7266 | 11060 | 13050 | 0.55418 | 0.4
M16B | 9 | 1 | 11887.6016 | 11883.3223 | 12013 | 88 | 928.6084 | 950 | 50 | 11103.2539 | 12699.6572 | 11060 | 13050 | 0.55178 | 0.4
M16B | 10 | 1 | 11885.8213 | 11878.29 | 12013 | 88 | 942.8223 | 950 | 50 | 11104.0566 | 12704.3145 | 11060 | 13050 | 0.52986 | 0.4
M16B | 11 | 1 | 11886.1582 | 11879.957 | 12013 | 88 | 940.4873 | 950 | 50 | 11101.9629 | 12702.3643 | 11060 | 13050 | 0.53031 | 0.4
M16B | 12 | 1 | 11884.4277 | 11876.4619 | 12013 | 88 | 941.5713 | 950 | 50 | 11103.4707 | 12704.1475 | 11060 | 13050 | 0.64146 | 0.4
M16B | 13 | 1 | 11884.3086 | 11878.9844 | 12013 | 88 | 937.5508 | 950 | 50 | 11101.0996 | 12697.3867 | 11060 | 13050 | 0.60411 | 0.4
M16B | 14 | 1 | 11881.7695 | 11875.7588 | 12013 | 88 | 940.5342 | 950 | 50 | 11099.9932 | 12698.3955 | 11060 | 13050 | 0.59611 | 0.4
M16B | 15 | 1 | 11880.373 | 11875.874 | 12013 | 88 | 927.9131 | 950 | 50 | 11099.0264 | 12692.5176 | 11060 | 13050 | 0.60623 | 0.4
M16B | 16 | 1 | 11875.2295 | 11870.3242 | 12013 | 88 | 913.2852 | 950 | 50 | 11098.0244 | 12689.0508 | 11060 | 13050 | 0.63339 | 0.4