## Instrument Improvements
VIIRS Instrument Improvements for J2+ (Note: items not in priority order)
Date of Latest Update: |  | 2012-08-10
Improvement ID | Title | Description | Benefit
HW-1 | NFR Reduction | Reduce optical scattering by 2x or more. May be achievable through combination of improved mirror surface quality, improved mirror coating (IAD), and improved contamination control. | Improved SDR quality in highly structured scenes for which optical scatter is a dominant source of radiometric error. Improved Ocean Color and Land Surface Temperature performance in presence of clouds and near ocean/land boundaries. Improved Cloud Optical Properties performance.
HW-2 | Produce Science Data at Full Resolution | Eliminate onboard aggregation, bowtie deletion and use of predictor bands to provide science data at full spatial resolution capability and simplify instrument functionality. Implement any hardware modifications required to accommodate increased data rate (est. 2X data rate increase). | Improved fire products, cloud mask, subpixel cloud detection and ocean products in coastal zones. High SNR margins allow trade-off of SNR for enhanced resolution to benefit many EDRs. Provides users with greater flexibility to optimize aggregation at product level.
HW-3 | Implement Dual Gain in M15 | Extend measurement range to higher radiances for fire products. | Requested by fire science team to determine fire power (i.e. T of fire).
HW-4 | Add WV Bands | Add 6.7 micron band for water vapor and polar winds/numerical weather forecasting applications. | Provides water vapor and polar winds. This H2O band is useful for cloud masks, especially over polar regions.
HW-5 | Add Band at 13.3 microns | Add band at 13.3 microns for cloud top pressure and altitude retrieval. | Provides improved forecasting and cloud top pressure and altitude. Note by improving cloud top pressure it also improves the other cloud top parameters as well. There are refereed papers with quantitative measurements of these improvements. This band also has benefits to the cloud mask. Important to NASA for (1) MODIS data record continuity of cloud detection (cloud mask) and optical property (cloud-top pressure and temperature used as input to retrieval algorithms) products , (2) cloud top pressure and emissivity data record continuity from MODIS (though this may be doable with coarser spatial resolution CrIS products). However, that single band isn't sufficient for MODIS continuity. At a minimum, the 13.64 and 13.94 µm are required in addition to the 13.3. Not sure how important the final MODIS CO2 band is (14.2 µm).
HW-6 | Additional Mirror Witness Samples | Acquire witness samples from every flight coating run. Mount HAM witness sample in HAM cavity. Expose witness samples to same environment as flight surfaces, to extent feasible, and peridically characterize these samples. | Reduces risk of flight mirror degradation that goes undetected prior to launch. Mirror degradation during prelaunch testing and storage can significantly increase optical scatter and adversely impact radiometric performance.
HW-7 | Redesign Spectralon Witness Sample for Accessibility | Current spectralon witness sample cannot be accessed for removal and characterization. | Allows trending of spectralon witness sample optical characteristics.
HW-8 | Implement Dual Gain in M6 | Extend measurement range to higher radiances to avoid saturation on clouds and desert plumes. | Supports greater range of aerosol loading over ocean. Could support far-field stray light corrections from for ocean color. For oceans, this is a low priority. If designed correctly, would address M6 rollover.
HW-9 | Eliminate Instrument Response Rollover | Redesign Readout Integrated Circuit (ROIC) chips to eliminate response rollover back into measurement range. This effect can produce spurious measured radiances that actually correspond to much higher observed scene radiances. | Eliminates possibility of spurious response due to rollover effect (eg. VIIRS F1 band M6)
HW-10 | Ghosting Fix | Eliminate ghosting artifacts in NFR caused by multiple reflections between FPA filters and dewar cold window. Ghosting artifacts increase scattered light contribution to radiometric error by 1% to 2%. Requires redesign to tilt, wedge and/or reposition dewar window. | Very slight improvement in data product quality in highly structured scenes given current level of optical scattering. Benefit would increase modestly if NFR were improved by 2x and would increase significantly if NFR were reduced by 10x.
HW-11 | Fix Dual Gain Anomaly | The dual gain anomaly refers to excessive noise and nonlinearity in a narrow region of the meaurement range (approx. 1% of range) occuring about 10% below the switchpoint between high and low gain states in dual gain bands. | Provides minor benefit to land and aerosol products. Anomaly is generally out of range for ocean products.
HW-12 | Remove I2/M7 degeneracy | I2 and M7 share a common filter. Redesign filter to provide different bandpass for M7. | Allows additional VisNIR band without block change. Aggregated I2 can replace M7. I2 is spec'ed radiometrically to be a general imaging band. So, aggregation may well reduce the noise, but does it have the necessary sensitivity that ocean color got from M7?
HW-13 | Add SWIR band to SDSM | Add detectors in SWIR band(s) to SDSM. Requires addition of a TE cooler. | Improves tracking of SD changes over time in SWIR bands. Current strategy is to extrapolate from VisNIR, but such extrapolation may not be valid.
HW-14 | M15 unaggregated | Can be done on F2 via SW change but is included in HW list due to possible s/c hardware implications. Delta DR is 0.7 Mbps | Improves fire product. Allows identification and filtering out of saturated samples.
HW-15 | Improve Reliability of Moving Mechanical Assemblies | Assuming JPSS will require mission life of 10 years instead of current 7 years, reliability will need to be improved, particularly for moving mechanical assemblies such as RTA , HAM and SDSM motors. | Increased reliability and lifetime.
HW-16 | Reduction of DNB stray light | Significant stray light seen in DNB over certain parts of SNPP orbit. Stray light model predicts do not match on-orbit measurement situation. | Performance of DNB in parts of orbit not affected by stray light is superior to previous sensors. DNB science interest is increasing.
HW-17 | Employ filters in SDSM from same filter runs as focal plane filters, with minimal oob/scattering effects
HW-18 | Ensure the SNPP VIIRS roll, pitch and yaw manevuers will be performed on J2+ | Critical for radiometric trending, system level SD/SDS/SDSM performance, and determination of TEB RVS | Provide critical instrument calibration and characterization data needed to continue key operational and research data sets
HW-19 | Improve the RTA scan speed | The scan speed can be improved by burning 6 or 8 bits in the SCE FPGA up from the current 4 bits. | Improves knowledge and precision of instrument scan angle to improve the determination of pointing.
HW-20 | Add thermistors to the OBC-BB | More thermistors on the BB provides a better characterization of BB gradients and matches the 12 thermistors on the heritage MODIS instruments, | Better determination of/accounting for effect of BB T gradients in VIIRS TEB calibration
HW-21 | Make the angular sample interval (ASI) ±5 to ±10% of the ASI equivalent | The M band detector sizes in the scan direction are ~20% wider than the ASI equivalent and the I band detector sized are 30 to 40% narrower than the ASI equivalent. It is recommended that the requirement for ASI be ±5 to ±10% of the ASI equivalent. SNPP VIIRS failed the MTF requirement for many M bands on the VisNIR FPA partly due to this. | GIFOV to match the sampling interval.

## Test Improvements
VIIRS Testing Improvements for J2+ (Note: items not in priority order)
Date of Latest Update: |  | 2012-08-10
Improvement ID | Title | Description | Benefit
TEST-1 | E2E Calibrtion Test for Reflective Bands | There is no end-to-end calibration test of the reflective bands in the baseline test program. A pathfinder test has been designed and implemented on F1 at the SC level. Result from this test thus far suggest that a test redesign is needed for improved collimated source irradiance uniformity and knowledge. | VIIRS is a radiometer employing data from onboard sources for calibration. As such, full system level testing in which the radiometric path is stimulated as it is on orbit is considered essential for design and for reliable compliance verification. This tests confirms measurements of the fundamental RSB observing mode and must be verified for each unit.
TEST-2 | Improved SD and SDS BRDF Characterization | Improve characterization of Solar Diffuser (SD) Bidirectional Reflectance Distribution Function (BRDF) and SDS BTDF over more angles to achieve significantly lower uncertainties, e.g., less than 1%. Perform measurements throughout the reflective band spectral region, including the SWIR to approximately 2500 nm. The current characterization technique provides values with uncertainties as high as 1.6% in some bands and characterization extends only to 1700 nm. | SD BRDF uncertainty is by far the dominant error contributor in the reflective band calibration uncertainty tree. Better SD BRDF characterization would provide higher quality SDRs in all reflective bands and improve EDR quality for all EDRs utilizing these bands.
TEST-3 | Maintain Calibration of OGSE Between Instrument Characterizations | All optical GSE used to stimulate VIIRS in test program should be calibrated with sufficient frequency to ensure adequate knowledge of source characteristics during testing. Calibrations of some key OGSE, in particular the BCS, have not been regularly performed. These tests must be documented in the OGSE maintanence manual. | Protects integrity of test program and avoids potentially enormous cost and schedule impacts of repeat testing if OGSE found to be out of calibration subsequent to use.
TEST-4 | Laser Driven Source for Reflective Band RSR Characterization | Replace SpMA with laser driven source for characterization of Relative Spectral Response (RSR) in the reflective bands. | Provides unpolarized illumination, reduced wavelength uncertainty, uniform illumination of entire focal plane or single band, and no spectral smile. Can provide absolute spectral response in appropriately designed test configuration, e.g., when coupled to integrating sphere with radiance monitors viewing exit port (demonstrated in SC level testing). Relative Spectral Output (RSO) better known and more repeatable than SpMA RSO.
TEST-5 | FTIR Measurement of Emissive Band RSRs | Replace SpMA with Fourier Transform Infrared Spectrometer (FTIR) for characterization of RSR in the emissive bands. Provides spectral coverage from approximately 2 microns to 13+ microns. | Greatly reduced data acquisition time. In-band (IB) and out-of-band (OOB) data can be collected at same time, potentially in less than one hour per band. Provides spectral resolution approximately 2x greater than SpMA. SNRs expected to be comparable to those obtained from SpMA. Stitching of IB and OOB data is simplified. Fast acquisitions provide opportunity for repeat collects to better estimate uncertainties.
TEST-6 | AOA Spectral Characterizaiton using FTIR | Same spectral measurement technique as in TEST-5 but performed at AOA level. | Provides demonstration of TEST-5 capability at lower integration level, with focal planes driven by laboratory readout electronics (not VIIRS EM) not constrained to operation only within VIIRS scan sectors.
TEST-7 | Acquire Additional Data for IB/OOB RSR Stitching | Increase spectral range in which IB and OOB RSR measurements overlap. Increase sampling in this overlap region. Alternative is to move to T-SIRCUS and use Argon laser sources for better SNR in region below 450 nm. | Improved stitching reduces uncertainty in OOB RSR and Integrated OOB (IOOB) response.
TEST-8 | Improve Radiometric Characterization of M13 LG | Characterize M13 LG radiometric performance throughout its measurement range. Current characterization does not reach maximum scene temperature. Less benefit unless dual gain is implemented in M15. Higher brightness source needed | Benefit to fire products. Reduced uncertainty in inferred fire temperature.
TEST-9 | Improve Stray Light Test | In current test configuration VIIRS is mounted with scan plane parallel to floor. Studio lamp sources are arranged to provide sampling within a quadrant above the scan plane and a small range of off-axis angles in the opposite quadrant. Desired improvement would provide sampling over entire hemisphere with line of sight to VIIRS nadir cavity. | Provides more comprehensive test that eliminates reliance on symmetry argument to estimate stray light impact in non-tested directions.
TEST-10 | Improve NFR Test | Modify test to provide along-track as well as in-track information. | Provide additional characterization information for assessing NFR impact on data products.
TEST-11 | Characterize AOA polarization sensitivity | Run same testing at AOA level for polarization characterization as is done at system level. This test will complement the system level test and is not considered a substitute for the system level testing of polarization characterization. | On-orbit changes in characterization cannot be tracked directly, but must be made as potential updates to an analytic sensor polarization sensitivy model. We do not now have a useful mode of this nature, and testing of polarization at AOA level is expected to improve our analytic understanding of how sensor design influences polarization characteristics.
TEST-12 | Trend Stray Light Performance | Perform stray light test twice, as early in ambient phase as possible, and again after TVAC testing just prior to shipment. Objective is to detect any degradation in mirror scatter, and as such, the post-TVAC test could be abbreviated to include only the inner lamp positions that directly illuminate the RTA primary (and possibly with RTA at only one [e.g., nadir] cross-track scan position). | Can detect long-term degradation in mirror scatter and allows extrapolation of any observed degradation over ground storage period.
TEST-13 | Improved RSB Calibration Source | Replace SIS-1 with SIS-2. | Spectral radiance at M1 increases significantly at least a factor of 2. Additional increase is available if SIS-2 exit port is reduced similar to SIS-1. SIS-2 has an auxiliary integrating sphere with lamp with a variable aperture. This enables DNB testing with the same spectral radiance shape across the dynamic range.
TEST-14 | Improved Characterization of M1-M3/LG | Replace ND filter technique for adjustment of bright SIS/TMC radiance levels with method that does not change source uniformity. Characterize source uniformity with and without attenuator. M1-M3/LG characterization requires source brighter than SIS-1 or SIS-2. | Improved source radiance uniformity and uniformity knowledge for bright SIS/TMC/LAA used for M1-M3 LG radiometric characterization will allow much more accurate inference of calibration coefficients.
TEST-15 | Characterize SIS-X Uniformity | Characterize uniformity of SIS-1 or SIS-2 used for radiometric characterization of reflective bands. | Knowledge of source nonuniformity allows more accurate characterization of calibrated response uniformity and additional processing to improve calibrated response uniformity.
TEST-16 | Improve VIIRS STOP model using actual TV test results | Compare STOP predictions with actual TV test results for pointing and band to band registration as function of T. This involves an overall improvement in STOP cal/val processes, including (1) utilizing cross-hair testing during all T plateau transitions, (2) add LSF simulations for edge detectors and not just center detectors, (3) monitor stress/strain on strategic structures during T transitions, (4) incorporate measured instrument Ts to tune the model. | Critical for design of J2+ and for knowledge of on-orbit performance
TEST-17 | Improve slit spacing in LSF testing | In F1, the reticles used in LSF testing did not have optimal slit spacing. For F2+ LSF testing, it is recommended that a new reticle be used with 9.1 sampling intervals apart for the M-band scan LSF and 8.3 sampling intervals apart for the I-band scan LSF thru-focus testing | Addresses undersampling/oversampling issues in constructing LSF experienced in F1 testing.
TEST-18 | Implement T/V testing of band to band registration of cold and warm focal planes | Part of TEST-16 above
TEST-19 | Utilize mosaic polarizer | Currently use 2 polarizing sheets in VIIRS characterization. Would be able to do this test with one polarizer. | Saves test time and schedule.
TEST-20 | M15 unaggregated | Can be done on F2 via SW change. Is also included in HW list due to possible s/c hardware implications. Delta DR is 0.7 Mbps | Improves fire product. Allows identification and filtering out of saturated samples.

## Software Improvements
VIIRS Software Improvements for J2+ (Note: items not in priority order)
Date of Latest Update: |  | 2012-08-10
Improvement ID | Title | Description | Benefit
SW-1 | Implement far field registration in cal sectors, particularly the SVS and BB | Improves the analysis of free Moon observations. Eliminates need to perform EV sector rotation before LRMs. Will enlarge SV without increasing scan angles. Since there appears to be a small gradient on the BB, all bands would use the same combination of thermistors, instead of possible band dependence. | Makes lunar roll maneuvers much easier to do and provides 1:1 aggregation. Makes BB measurements more consistent wrt T.
SW-2 | Acquiring more frames in SV. | Narrow SV causes issues with lunar intrusion, a larger SV can be used to determine offset directly as in MODIS, rather than searching for other scans. BB is large enough to accommodate more samples | Improves quality of TEB band data
SW-3 | Test the capability of operating instrument in diagnostic mode and processing data on the ground. | Currently IDPS or any science/engineering team has the capability to do this. | Capability untested on SNPP VIIRS, yet the instrument is designed to be able to do this.
SW-4 | Increase SV to 96 M-band samples along with sample drop done in EV | The SV sample size could be increased to 96 M-samples along with sample drop as the EV is doing. Sample size increase would (1) need to be studied and (2) would require APID changes and downlink bandwidth changes. Sample drop would be an easy FSW change.
SW-5 | Sample and transmit the attitude data at a higher frequency (eg. 5 to 10 Hz sampling frequency) | For pointing accuracy, the attitude data are currently sampled at 1 Hz. That produces 0.5 Hz Nyquist frequency. The SNPP attitude data occasionally exhibits some higher frequency (>0.5 Hz). In order to capture this higher frequency, it is necessary to sample and transmit the atitude data at a higher frequency (i.e. 5 to 10 Hz). | Provides improvement in the determination of pointing accuracy.
SW-6 | Collect HG/LG OBC data (BB and SV in particular) in each scan | Allows processing data on every scan for RSB and TEB. Half of available samples will be in high gain and the other half in low gain.
SW-7 | Make the angular sample interval (ASI) ±5 to ±10% of the ASI equivalent | The M band detector sizes in the scan direction are ~20% wider than the ASI equivalent and the I band detector sized are 30 to 40% narrower than the ASI equivalent. It is recommended that the requirement for ASI be ±5 to ±10% of the ASI equivalent. SNPP VIIRS failed the MTF requirement for many M bands on the VisNIR FPA partly due to this. | GIFOV to match the sampling interval.
SW-8 | M15 unaggregated | Can be done on J1 via SW change but is included in HW list due to possible s/c hardware implications. Delta DR is 0.7 Mbps | Improves fire product. Allows identification and filtering out of saturated samples.