## Sheet1
1. This table provides key performance of FU1 TEB radiometric parameters.
2. The Bands I4 - I5, M12 - M15 are from run2 (ambient phase II) and M16A/B from run 1 due to unavailability of M16A/B in run 2.
3. Band 13 low-gain only provides radiometric coefficients because of LABB temperature is limited to 345K and TMC data is less reliable.
4. The dn@Lmax provides an estimation to the dynamic range for single-gain bands.
5. The test was set to aggregated mode. For convinence to compare with the specification, the SNR and NEdT at Ltyp are converted to equivalent diagnostic mode.
The data are calculated based on the SNR@Ltyp divided by 3^0.5. These data are listed in columns M - N (YELLOW BOX).
6. The columns O - R are the part of radiometric uncertainty of the measurements. The values in YELLOW BOX are for diagnostic mode.
Band | E_side | HAM | Detector | Subsample | Calculated based upon the measurement data
[5 empty] | Coef(0) | Coef(1),Gain
I4 | A | 1 | 1 | 1 | -39.8232 | 1128.9562 | 1137.0694
I4 | A | 1 | 1 | 2 | -39.8768 | 1129.2112 | 1137.1311
I4 | A | 1 | 2 | 1 | -39.6988 | 1125.1471
I4 | A | 1 | 2 | 2 | -39.7246 | 1125.0359
I4 | A | 1 | 3 | 1 | -39.8277 | 1130.4686
I4 | A | 1 | 3 | 2 | -39.8797 | 1130.8182
I4 | A | 1 | 4 | 1 | -39.7469 | 1134.1924
I4 | A | 1 | 4 | 2 | -39.6556 | 1133.9419
I4 | A | 1 | 5 | 1 | -39.8336 | 1135.9407
I4 | A | 1 | 5 | 2 | -39.8124 | 1136.1539
I4 | A | 1 | 6 | 1 | -39.5063 | 1107.0109
I4 | A | 1 | 6 | 2 | -39.4955 | 1107.0643
I4 | A | 1 | 7 | 1 | -39.9708 | 1145.1351
I4 | A | 1 | 7 | 2 | -39.9999 | 1145.4683
I4 | A | 1 | 8 | 1 | -39.9563 | 1132.0924
I4 | A | 1 | 8 | 2 | -39.9572 | 1131.9927
I4 | A | 1 | 9 | 1 | -40.0434 | 1142.0508
I4 | A | 1 | 9 | 2 | -40.0711 | 1142.3652
I4 | A | 1 | 10 | 1 | -39.797 | 1129.5804
I4 | A | 1 | 10 | 2 | -39.7901 | 1129.4104
I4 | A | 1 | 11 | 1 | -39.9561 | 1145.4224
I4 | A | 1 | 11 | 2 | -40.0004 | 1145.673
I4 | A | 1 | 12 | 1 | -39.6605 | 1115.9139
I4 | A | 1 | 12 | 2 | -39.6518 | 1115.7961
I4 | A | 1 | 13 | 1 | -39.8493 | 1125.4454
I4 | A | 1 | 13 | 2 | -39.8735 | 1125.6082
I4 | A | 1 | 14 | 1 | -39.8782 | 1124.9298
I4 | A | 1 | 14 | 2 | -39.8125 | 1124.7494
I4 | A | 1 | 15 | 1 | -39.4072 | 1093.4807
I4 | A | 1 | 15 | 2 | -39.4223 | 1093.5118
I4 | A | 1 | 16 | 1 | -39.8539 | 1124.5295
I4 | A | 1 | 16 | 2 | -39.7769 | 1124.4167
I4 | A | 1 | 17 | 1 | -40.2031 | 1149.4795
I4 | A | 1 | 17 | 2 | -40.255 | 1149.6821
I4 | A | 1 | 18 | 1 | -40.1543 | 1145.4818
I4 | A | 1 | 18 | 2 | -40.1297 | 1145.3198
I4 | A | 1 | 19 | 1 | -40.1197 | 1145.9042
I4 | A | 1 | 19 | 2 | -40.1455 | 1146.0638
I4 | A | 1 | 20 | 1 | -40.2965 | 1158.925
I4 | A | 1 | 20 | 2 | -40.2997 | 1158.8983
I4 | A | 1 | 21 | 1 | -40.2934 | 1156.4589
I4 | A | 1 | 21 | 2 | -40.2824 | 1156.6794
I4 | A | 1 | 22 | 1 | -40.3471 | 1165.4047
I4 | A | 1 | 22 | 2 | -40.3453 | 1165.3264
I4 | A | 1 | 23 | 1 | -40.1118 | 1138.625
I4 | A | 1 | 23 | 2 | -40.1566 | 1138.8781
I4 | A | 1 | 24 | 1 | -40.121 | 1139.741
I4 | A | 1 | 24 | 2 | -40.0972 | 1139.6083
I4 | A | 1 | 25 | 1 | -39.8638 | 1123.0646
I4 | A | 1 | 25 | 2 | -39.9492 | 1123.3575
I4 | A | 1 | 26 | 1 | -39.9589 | 1133.4556
I4 | A | 1 | 26 | 2 | -40.0069 | 1133.5454
I4 | A | 1 | 27 | 1 | -40.4835 | 1165.1895
I4 | A | 1 | 27 | 2 | -40.4238 | 1165.2216
I4 | A | 1 | 28 | 1 | -40.2191 | 1145.4976
I4 | A | 1 | 28 | 2 | -40.2435 | 1145.4749
I4 | A | 1 | 29 | 1 | -40.0334 | 1128.1558
I4 | A | 1 | 29 | 2 | -40.0754 | 1128.4296
I4 | A | 1 | 30 | 1 | -40.5055 | 1156.3674
I4 | A | 1 | 30 | 2 | -40.4796 | 1156.2146
I4 | A | 1 | 31 | 1 | -40.1467 | 1145.1218
I4 | A | 1 | 31 | 2 | -40.186 | 1145.3593
I4 | A | 1 | 32 | 1 | -40.3201 | 1148.0748
I4 | A | 1 | 32 | 2 | -40.2689 | 1147.8929
I4 | A | 2 | 1 | 1 | -39.7149 | 1127.6135 | 1135.6894
I4 | A | 2 | 1 | 2 | -39.7891 | 1128.0685 | 1135.7171
I4 | A | 2 | 2 | 1 | -39.7523 | 1123.7345
I4 | A | 2 | 2 | 2 | -39.7267 | 1123.6399
I4 | A | 2 | 3 | 1 | -39.7362 | 1128.7767
I4 | A | 2 | 3 | 2 | -39.7406 | 1129.0266
I4 | A | 2 | 4 | 1 | -39.7944 | 1132.8199
I4 | A | 2 | 4 | 2 | -39.7121 | 1132.5858
I4 | A | 2 | 5 | 1 | -39.8439 | 1134.5887
I4 | A | 2 | 5 | 2 | -39.8601 | 1134.7844
I4 | A | 2 | 6 | 1 | -39.572 | 1105.874
I4 | A | 2 | 6 | 2 | -39.5172 | 1105.7844
I4 | A | 2 | 7 | 1 | -39.9323 | 1143.6027
I4 | A | 2 | 7 | 2 | -39.9903 | 1143.8665
I4 | A | 2 | 8 | 1 | -39.929 | 1130.562
I4 | A | 2 | 8 | 2 | -39.9057 | 1130.3525
I4 | A | 2 | 9 | 1 | -39.9019 | 1140.3812
I4 | A | 2 | 9 | 2 | -39.9109 | 1140.6782
I4 | A | 2 | 10 | 1 | -39.8528 | 1128.3074
I4 | A | 2 | 10 | 2 | -39.6661 | 1127.8235
I4 | A | 2 | 11 | 1 | -40.0028 | 1144.1423
I4 | A | 2 | 11 | 2 | -40.0417 | 1144.5608
I4 | A | 2 | 12 | 1 | -39.7132 | 1114.6151
I4 | A | 2 | 12 | 2 | -39.6825 | 1114.395
I4 | A | 2 | 13 | 1 | -39.8465 | 1124.1965
I4 | A | 2 | 13 | 2 | -39.8827 | 1124.3208
I4 | A | 2 | 14 | 1 | -39.8896 | 1123.5378
I4 | A | 2 | 14 | 2 | -39.7359 | 1123.1594
I4 | A | 2 | 15 | 1 | -39.392 | 1092.2416
I4 | A | 2 | 15 | 2 | -39.5014 | 1092.5153
I4 | A | 2 | 16 | 1 | -39.8291 | 1122.947
I4 | A | 2 | 16 | 2 | -39.7292 | 1122.9158
I4 | A | 2 | 17 | 1 | -40.0769 | 1147.9994
I4 | A | 2 | 17 | 2 | -40.1565 | 1148.2136
I4 | A | 2 | 18 | 1 | -40.1591 | 1144.0485
I4 | A | 2 | 18 | 2 | -40.0893 | 1143.9525
I4 | A | 2 | 19 | 1 | -40.0738 | 1144.2365
I4 | A | 2 | 19 | 2 | -40.1196 | 1144.642
I4 | A | 2 | 20 | 1 | -40.266 | 1157.6581
I4 | A | 2 | 20 | 2 | -40.1947 | 1157.2534
I4 | A | 2 | 21 | 1 | -40.2323 | 1154.9347
I4 | A | 2 | 21 | 2 | -40.2727 | 1155.2307
I4 | A | 2 | 22 | 1 | -40.3783 | 1163.9857
I4 | A | 2 | 22 | 2 | -40.3874 | 1163.8369
I4 | A | 2 | 23 | 1 | -40.0429 | 1137.0826
I4 | A | 2 | 23 | 2 | -40.1594 | 1137.4285
I4 | A | 2 | 24 | 1 | -40.1674 | 1138.5145
I4 | A | 2 | 24 | 2 | -40.0873 | 1138.4164
I4 | A | 2 | 25 | 1 | -39.859 | 1121.9784
I4 | A | 2 | 25 | 2 | -39.8807 | 1122.0817
I4 | A | 2 | 26 | 1 | -40.1061 | 1132.6827
I4 | A | 2 | 26 | 2 | -40.0519 | 1132.4481
I4 | A | 2 | 27 | 1 | -40.4728 | 1163.9486
I4 | A | 2 | 27 | 2 | -40.474 | 1164.2383
I4 | A | 2 | 28 | 1 | -40.3095 | 1144.3765
I4 | A | 2 | 28 | 2 | -40.3376 | 1144.2145
I4 | A | 2 | 29 | 1 | -40.0003 | 1126.8052
I4 | A | 2 | 29 | 2 | -40.0241 | 1126.9767
I4 | A | 2 | 30 | 1 | -40.3603 | 1154.4598
I4 | A | 2 | 30 | 2 | -40.3334 | 1154.3771
I4 | A | 2 | 31 | 1 | -40.1962 | 1143.6837
I4 | A | 2 | 31 | 2 | -40.2243 | 1143.9019
I4 | A | 2 | 32 | 1 | -40.2999 | 1146.71
I4 | A | 2 | 32 | 2 | -40.2483 | 1146.6293
I5 | A | 1 | 1 | 1 | -30.3643 | 121.1401 | 113.3869
I5 | A | 1 | 1 | 2 | -33.6792 | 119.4494 | 113.3925
I5 | A | 1 | 2 | 1 | -29.6284 | 114.0503
I5 | A | 1 | 2 | 2 | -26.251 | 115.8563
I5 | A | 1 | 3 | 1 | -30.4308 | 122.4244
I5 | A | 1 | 3 | 2 | -33.774 | 120.6609
I5 | A | 1 | 4 | 1 | -29.7532 | 114.5467
I5 | A | 1 | 4 | 2 | -26.441 | 116.4062
I5 | A | 1 | 5 | 1 | -30.5162 | 123.7432
I5 | A | 1 | 5 | 2 | -33.8686 | 121.96
I5 | A | 1 | 6 | 1 | -29.923 | 115.9032
I5 | A | 1 | 6 | 2 | -26.6941 | 117.6488
I5 | A | 1 | 7 | 1 | -30.3703 | 122.2311
I5 | A | 1 | 7 | 2 | -33.6959 | 120.4523
I5 | A | 1 | 8 | 1 | -29.8182 | 114.3754
I5 | A | 1 | 8 | 2 | -26.4842 | 116.1077
I5 | A | 1 | 9 | 1 | -30.3651 | 120.6231
I5 | A | 1 | 9 | 2 | -33.7863 | 118.844
I5 | A | 1 | 10 | 1 | -29.5234 | 111.1999
I5 | A | 1 | 10 | 2 | -25.9745 | 112.9997
I5 | A | 1 | 11 | 1 | -30.2491 | 119.4819
I5 | A | 1 | 11 | 2 | -33.6486 | 117.6896
I5 | A | 1 | 12 | 1 | -29.3449 | 109.4807
I5 | A | 1 | 12 | 2 | -25.8145 | 111.3257
I5 | A | 1 | 13 | 1 | -30.1741 | 119.0014
I5 | A | 1 | 13 | 2 | -33.5782 | 117.1921
I5 | A | 1 | 14 | 1 | -29.1594 | 107.8775
I5 | A | 1 | 14 | 2 | -25.5088 | 109.8442
I5 | A | 1 | 15 | 1 | -29.9113 | 115.7133
I5 | A | 1 | 15 | 2 | -33.6283 | 113.8087
I5 | A | 1 | 16 | 1 | -28.9654 | 105.122
I5 | A | 1 | 16 | 2 | -25.0306 | 107.195
I5 | A | 1 | 17 | 1 | -29.9382 | 116.6247
I5 | A | 1 | 17 | 2 | -33.5819 | 114.6829
I5 | A | 1 | 18 | 1 | -29.2909 | 109.8662
I5 | A | 1 | 18 | 2 | -25.8493 | 111.6895
I5 | A | 1 | 19 | 1 | -29.882 | 115.1581
I5 | A | 1 | 19 | 2 | -33.6012 | 113.2318
I5 | A | 1 | 20 | 1 | -29.1754 | 107.7575
I5 | A | 1 | 20 | 2 | -25.6503 | 109.5927
I5 | A | 1 | 21 | 1 | -30.0213 | 116.284
I5 | A | 1 | 21 | 2 | -33.9628 | 114.2027
I5 | A | 1 | 22 | 1 | -29.2081 | 107.5485
I5 | A | 1 | 22 | 2 | -25.2022 | 109.599
I5 | A | 1 | 23 | 1 | -29.8054 | 114.9179
I5 | A | 1 | 23 | 2 | -33.2864 | 113.1358
I5 | A | 1 | 24 | 1 | -29.2627 | 107.7839
I5 | A | 1 | 24 | 2 | -25.9582 | 109.5181
I5 | A | 1 | 25 | 1 | -29.6269 | 113.4969
I5 | A | 1 | 25 | 2 | -32.9548 | 111.7414
I5 | A | 1 | 26 | 1 | -29.0824 | 105.5659
I5 | A | 1 | 26 | 2 | -25.6052 | 107.359
I5 | A | 1 | 27 | 1 | -29.7875 | 113.7881
I5 | A | 1 | 27 | 2 | -33.3524 | 111.9211
I5 | A | 1 | 28 | 1 | -29.0326 | 106.4066
I5 | A | 1 | 28 | 2 | -25.6094 | 108.2482
I5 | A | 1 | 29 | 1 | -29.7576 | 113.0811
I5 | A | 1 | 29 | 2 | -33.2048 | 111.2967
I5 | A | 1 | 30 | 1 | -29.0081 | 104.8288
I5 | A | 1 | 30 | 2 | -25.4656 | 106.6367
I5 | A | 1 | 31 | 1 | -29.3543 | 106.763
I5 | A | 1 | 31 | 2 | -32.8475 | 104.9083
I5 | A | 1 | 32 | 1 | -28.8132 | 102.904
I5 | A | 1 | 32 | 2 | -25.2383 | 104.8355
I5 | A | 2 | 1 | 1 | -30.3757 | 120.6689 | 112.8665
I5 | A | 2 | 1 | 2 | -33.6378 | 118.9682 | 112.8751
I5 | A | 2 | 2 | 1 | -29.5812 | 113.5127
I5 | A | 2 | 2 | 2 | -26.1737 | 115.3035
I5 | A | 2 | 3 | 1 | -30.4351 | 121.9536
I5 | A | 2 | 3 | 2 | -33.7742 | 120.209
I5 | A | 2 | 4 | 1 | -29.7722 | 114.1039
I5 | A | 2 | 4 | 2 | -26.3833 | 115.916
I5 | A | 2 | 5 | 1 | -30.5047 | 123.1104
I5 | A | 2 | 5 | 2 | -33.8842 | 121.3922
I5 | A | 2 | 6 | 1 | -29.8865 | 115.4016
I5 | A | 2 | 6 | 2 | -26.6204 | 117.1281
I5 | A | 2 | 7 | 1 | -30.4078 | 121.667
I5 | A | 2 | 7 | 2 | -33.7644 | 119.9442
I5 | A | 2 | 8 | 1 | -29.717 | 113.7833
I5 | A | 2 | 8 | 2 | -26.3776 | 115.5315
I5 | A | 2 | 9 | 1 | -30.2829 | 120.0052
I5 | A | 2 | 9 | 2 | -33.6723 | 118.2285
I5 | A | 2 | 10 | 1 | -29.462 | 110.6556
I5 | A | 2 | 10 | 2 | -25.9087 | 112.4782
I5 | A | 2 | 11 | 1 | -30.2522 | 118.8983
I5 | A | 2 | 11 | 2 | -33.653 | 117.1184
I5 | A | 2 | 12 | 1 | -29.276 | 108.9524
I5 | A | 2 | 12 | 2 | -25.7531 | 110.7863
I5 | A | 2 | 13 | 1 | -30.2192 | 118.4637
I5 | A | 2 | 13 | 2 | -33.6336 | 116.6634
I5 | A | 2 | 14 | 1 | -29.2406 | 107.4303
I5 | A | 2 | 14 | 2 | -25.5452 | 109.412
I5 | A | 2 | 15 | 1 | -29.8899 | 115.1325
I5 | A | 2 | 15 | 2 | -33.5529 | 113.2015
I5 | A | 2 | 16 | 1 | -28.9041 | 104.5874
I5 | A | 2 | 16 | 2 | -24.963 | 106.6608
I5 | A | 2 | 17 | 1 | -29.984 | 116.1356
I5 | A | 2 | 17 | 2 | -33.5941 | 114.1933
I5 | A | 2 | 18 | 1 | -29.3713 | 109.4379
I5 | A | 2 | 18 | 2 | -25.8744 | 111.2071
I5 | A | 2 | 19 | 1 | -29.7844 | 114.5615
I5 | A | 2 | 19 | 2 | -33.5399 | 112.671
I5 | A | 2 | 20 | 1 | -29.2056 | 107.3025
I5 | A | 2 | 20 | 2 | -25.656 | 109.1152
I5 | A | 2 | 21 | 1 | -30.1225 | 115.8431
I5 | A | 2 | 21 | 2 | -34.1038 | 113.7948
I5 | A | 2 | 22 | 1 | -29.1814 | 106.9979
I5 | A | 2 | 22 | 2 | -25.2288 | 109.0707
I5 | A | 2 | 23 | 1 | -29.8162 | 114.3497
I5 | A | 2 | 23 | 2 | -33.2526 | 112.5476
I5 | A | 2 | 24 | 1 | -29.2873 | 107.3362
I5 | A | 2 | 24 | 2 | -25.9937 | 109.0812
I5 | A | 2 | 25 | 1 | -29.6619 | 113.0469
I5 | A | 2 | 25 | 2 | -32.9735 | 111.2681
I5 | A | 2 | 26 | 1 | -28.9069 | 105.0667
I5 | A | 2 | 26 | 2 | -25.3957 | 106.8402
I5 | A | 2 | 27 | 1 | -29.7406 | 113.2357
I5 | A | 2 | 27 | 2 | -33.2806 | 111.3929
I5 | A | 2 | 28 | 1 | -29.0832 | 105.9581
I5 | A | 2 | 28 | 2 | -25.5715 | 107.7778
I5 | A | 2 | 29 | 1 | -29.7022 | 112.4978
I5 | A | 2 | 29 | 2 | -33.1407 | 110.7301
I5 | A | 2 | 30 | 1 | -28.9557 | 104.3466
I5 | A | 2 | 30 | 2 | -25.429 | 106.142
I5 | A | 2 | 31 | 1 | -29.023 | 106.1649
I5 | A | 2 | 31 | 2 | -32.5395 | 104.3474
I5 | A | 2 | 32 | 1 | -28.6358 | 102.3281
I5 | A | 2 | 32 | 2 | -25.0581 | 104.2638
M12 | A | 1 | 1 | 1 | -32.4699 | 745.9375 | 1159.7459
M12 | A | 1 | 2 | 1 | -37.0928 | 1156.4301
M12 | A | 1 | 3 | 1 | -37.2976 | 1175.2598
M12 | A | 1 | 4 | 1 | -37.0783 | 1150.7885
M12 | A | 1 | 5 | 1 | -37.5346 | 1191.0283
M12 | A | 1 | 6 | 1 | -36.972 | 1134.2567
M12 | A | 1 | 7 | 1 | -37.3497 | 1177.4764
M12 | A | 1 | 8 | 1 | -37.107 | 1139.5532
M12 | A | 1 | 9 | 1 | -37.5239 | 1184.6412
M12 | A | 1 | 10 | 1 | -36.983 | 1130.3379
M12 | A | 1 | 11 | 1 | -37.612 | 1181.7693
M12 | A | 1 | 12 | 1 | -36.7746 | 1110.8546
M12 | A | 1 | 13 | 1 | -37.6499 | 1183.7758
M12 | A | 1 | 14 | 1 | -37.0093 | 1130.5453
M12 | A | 1 | 15 | 1 | -37.8099 | 1189.7262
M12 | A | 1 | 16 | 1 | -37.2365 | 1135.3373
M12 | A | 2 | 1 | 1 | -32.4697 | 745.2924 | 1158.2876
M12 | A | 2 | 2 | 1 | -37.0653 | 1154.8691
M12 | A | 2 | 3 | 1 | -37.3018 | 1173.7739
M12 | A | 2 | 4 | 1 | -37.0548 | 1149.3555
M12 | A | 2 | 5 | 1 | -37.5705 | 1189.6047
M12 | A | 2 | 6 | 1 | -36.9973 | 1132.8651
M12 | A | 2 | 7 | 1 | -37.3719 | 1176.1051
M12 | A | 2 | 8 | 1 | -37.031 | 1137.9457
M12 | A | 2 | 9 | 1 | -37.4995 | 1183.0281
M12 | A | 2 | 10 | 1 | -36.9815 | 1129.0632
M12 | A | 2 | 11 | 1 | -37.6092 | 1180.2126
M12 | A | 2 | 12 | 1 | -36.7519 | 1109.5457
M12 | A | 2 | 13 | 1 | -37.6522 | 1182.3124
M12 | A | 2 | 14 | 1 | -37.0246 | 1129.293
M12 | A | 2 | 15 | 1 | -37.7201 | 1188.0522
M12 | A | 2 | 16 | 1 | -37.1798 | 1133.6945
M13 | A | 1 | 1 | 1 | -33.7934 | 595.3817 | 589.5056
M13 | A | 1 | 2 | 1 | -33.4523 | 576.5771
M13 | A | 1 | 3 | 1 | -33.9265 | 600.0939
M13 | A | 1 | 4 | 1 | -33.5111 | 579.1065
M13 | A | 1 | 5 | 1 | -34.0966 | 603.2935
M13 | A | 1 | 6 | 1 | -33.506 | 580.3857
M13 | A | 1 | 7 | 1 | -34.0213 | 600.1388
M13 | A | 1 | 8 | 1 | -33.3763 | 572.9642
M13 | A | 1 | 9 | 1 | -34.084 | 599.4692
M13 | A | 1 | 10 | 1 | -33.4316 | 580.2411
M13 | A | 1 | 11 | 1 | -34.1578 | 602.9363
M13 | A | 1 | 12 | 1 | -33.4614 | 579.25
M13 | A | 1 | 13 | 1 | -34.1686 | 599.9185
M13 | A | 1 | 14 | 1 | -33.6764 | 581.2133
M13 | A | 1 | 15 | 1 | -34.1416 | 597.4906
M13 | A | 1 | 16 | 1 | -33.6413 | 577.8801
M13 | A | 2 | 1 | 1 | -33.7823 | 594.8004 | 588.8501
M13 | A | 2 | 2 | 1 | -33.4398 | 576.0004
M13 | A | 2 | 3 | 1 | -33.9367 | 599.3817
M13 | A | 2 | 4 | 1 | -33.4944 | 578.4727
M13 | A | 2 | 5 | 1 | -34.086 | 602.5377
M13 | A | 2 | 6 | 1 | -33.4799 | 579.751
M13 | A | 2 | 7 | 1 | -34.0318 | 599.4731
M13 | A | 2 | 8 | 1 | -33.3839 | 572.287
M13 | A | 2 | 9 | 1 | -34.08 | 598.7606
M13 | A | 2 | 10 | 1 | -33.4217 | 579.6312
M13 | A | 2 | 11 | 1 | -34.1718 | 602.2525
M13 | A | 2 | 12 | 1 | -33.4242 | 578.608
M13 | A | 2 | 13 | 1 | -34.1584 | 599.2671
M13 | A | 2 | 14 | 1 | -33.6745 | 580.6212
M13 | A | 2 | 15 | 1 | -34.1601 | 596.857
M13 | A | 2 | 16 | 1 | -33.6584 | 577.2029
M14 | A | 1 | 1 | 1 | -58.4991 | 183.1142 | 177.3215
M14 | A | 1 | 2 | 1 | -58.7135 | 183.3896
M14 | A | 1 | 3 | 1 | -58.0637 | 178.6222
M14 | A | 1 | 4 | 1 | -58.4422 | 180.7651
M14 | A | 1 | 5 | 1 | -58.5031 | 181.138
M14 | A | 1 | 6 | 1 | -58.4094 | 180.4113
M14 | A | 1 | 7 | 1 | -58.2736 | 179.3192
M14 | A | 1 | 8 | 1 | -57.884 | 176.3333
M14 | A | 1 | 9 | 1 | -57.7989 | 176.4773
M14 | A | 1 | 10 | 1 | -57.7127 | 174.9074
M14 | A | 1 | 11 | 1 | -57.5824 | 173.2622
M14 | A | 1 | 12 | 1 | -57.7794 | 174.838
M14 | A | 1 | 13 | 1 | -58.1879 | 177.5127
M14 | A | 1 | 14 | 1 | -57.4476 | 171.2088
M14 | A | 1 | 15 | 1 | -57.693 | 174.3165
M14 | A | 1 | 16 | 1 | -58.3551 | 178.3602
M14 | A | 2 | 1 | 1 | -58.5025 | 181.7726 | 175.8957
M14 | A | 2 | 2 | 1 | -58.5916 | 181.9877
M14 | A | 2 | 3 | 1 | -57.9912 | 177.2272
M14 | A | 2 | 4 | 1 | -58.2833 | 179.3277
M14 | A | 2 | 5 | 1 | -58.3824 | 179.6985
M14 | A | 2 | 6 | 1 | -58.2611 | 179.0002
M14 | A | 2 | 7 | 1 | -58.2222 | 177.9211
M14 | A | 2 | 8 | 1 | -57.758 | 174.9407
M14 | A | 2 | 9 | 1 | -57.7253 | 175.0476
M14 | A | 2 | 10 | 1 | -57.6198 | 173.5211
M14 | A | 2 | 11 | 1 | -57.4242 | 171.7997
M14 | A | 2 | 12 | 1 | -57.6663 | 173.4169
M14 | A | 2 | 13 | 1 | -58.0982 | 176.0458
M14 | A | 2 | 14 | 1 | -57.2681 | 169.7692
M14 | A | 2 | 15 | 1 | -57.589 | 172.8369
M14 | A | 2 | 16 | 1 | -58.1342 | 176.7594
M15 | A | 1 | 1 | 1 | -41.5492 | 170.2234 | 160.3275
M15 | A | 1 | 2 | 1 | -41.2209 | 166.9371
M15 | A | 1 | 3 | 1 | -41.5251 | 169.637
M15 | A | 1 | 4 | 1 | -41.0313 | 165.2525
M15 | A | 1 | 5 | 1 | -41.0646 | 165.2833
M15 | A | 1 | 6 | 1 | -40.5427 | 160.1359
M15 | A | 1 | 7 | 1 | -40.57 | 160.3592
M15 | A | 1 | 8 | 1 | -40.6129 | 160.6762
M15 | A | 1 | 9 | 1 | -40.7994 | 162.3161
M15 | A | 1 | 10 | 1 | -40.3772 | 158.4274
M15 | A | 1 | 11 | 1 | -40.5283 | 159.1021
M15 | A | 1 | 12 | 1 | -39.9804 | 154.2666
M15 | A | 1 | 13 | 1 | -40.2333 | 156.3112
M15 | A | 1 | 14 | 1 | -39.8648 | 152.3883
M15 | A | 1 | 15 | 1 | -39.9755 | 153.4927
M15 | A | 1 | 16 | 1 | -39.656 | 150.1751
M15 | A | 2 | 1 | 1 | -41.469 | 169.1621 | 159.2648
M15 | A | 2 | 2 | 1 | -41.1738 | 165.8972
M15 | A | 2 | 3 | 1 | -41.4744 | 168.5553
M15 | A | 2 | 4 | 1 | -41.0109 | 164.1743
M15 | A | 2 | 5 | 1 | -41.0499 | 164.2112
M15 | A | 2 | 6 | 1 | -40.4938 | 159.1098
M15 | A | 2 | 7 | 1 | -40.5511 | 159.3295
M15 | A | 2 | 8 | 1 | -40.5613 | 159.6223
M15 | A | 2 | 9 | 1 | -40.7592 | 161.2599
M15 | A | 2 | 10 | 1 | -40.3466 | 157.3949
M15 | A | 2 | 11 | 1 | -40.4128 | 158.0161
M15 | A | 2 | 12 | 1 | -39.925 | 153.2239
M15 | A | 2 | 13 | 1 | -40.1829 | 155.2412
M15 | A | 2 | 14 | 1 | -39.7595 | 151.2945
M15 | A | 2 | 15 | 1 | -39.8749 | 152.3766
M15 | A | 2 | 16 | 1 | -39.6005 | 149.0835
M16A | A | 1 | 1 | 1 | -71.031 | 201.06 | 203.57
M16A | A | 1 | 2 | 1 | -72.216 | 210.65
M16A | A | 1 | 3 | 1 | -67.134 | 211.08
M16A | A | 1 | 4 | 1 | -59.016 | 203.91
M16A | A | 1 | 5 | 1 | -69.426 | 209.42
M16A | A | 1 | 6 | 1 | -63.878 | 208.71
M16A | A | 1 | 7 | 1 | -68.633 | 206.57
M16A | A | 1 | 8 | 1 | -63.053 | 192.22
M16A | A | 1 | 9 | 1 | -61.315 | 202.85
M16A | A | 1 | 10 | 1 | -60.048 | 199.53
M16A | A | 1 | 11 | 1 | -68.46 | 204.64
M16A | A | 1 | 12 | 1 | -78.372 | 202.26
M16A | A | 1 | 13 | 1 | -76.171 | 200.89
M16A | A | 1 | 14 | 1 | -77.482 | 201.19
M16A | A | 1 | 15 | 1 | -73.341 | 196.06
M16A | A | 1 | 16 | 1 | -63.056 | 201.23
M16B | A | 1 | 1 | 1 | -64.687 | 200.67 | 203.5214
M16B | A | 1 | 2 | 1 | -69.285 | 208.95
M16B | A | 1 | 3 | 1 | -74.478 | 209.15
M16B | A | 1 | 4 | 1 | -81.212 | 211.81
M16B | A | 1 | 5 | 1 | -70.944 | 206.98
M16B | A | 1 | 6 | 1 | -75.603 | 205.9
M16B | A | 1 | 7 | 1 | -70.845 | 207.95
M16B | A | 1 | 8 | 1 | -67.083 | 194.44
M16B | A | 1 | 9 | 1 | -75.851 | 204.63
M16B | A | 1 | 10 | 1 | -75.748 | 204.28
M16B | A | 1 | 11 | 1 | -67.639 | 200.73
M16B | A | 1 | 12 | 1 | -57.134 | 201.35
M16B | A | 1 | 13 | 1 | -58.427 | 199.73
M16B | A | 1 | 14 | 1 | -57.233 | 200.09
M16B | A | 1 | 15 | 1 | -57.236 | 193.31
M16B | A | 1 | 16 | 1 | -69.542 | 194.41
M16A | A | 2 | 1 | 1 | -70.907 | 200.52 | 202.8714
M16A | A | 2 | 2 | 1 | -71.941 | 209.9
M16A | A | 2 | 3 | 1 | -67.265 | 210.28
M16A | A | 2 | 4 | 1 | -58.868 | 203.3
M16A | A | 2 | 5 | 1 | -67.854 | 207.67
M16A | A | 2 | 6 | 1 | -63.715 | 208.09
M16A | A | 2 | 7 | 1 | -68.55 | 205.96
M16A | A | 2 | 8 | 1 | -62.044 | 191.04
M16A | A | 2 | 9 | 1 | -68.952 | 203.27
M16A | A | 2 | 10 | 1 | -59.826 | 198.95
M16A | A | 2 | 11 | 1 | -68.334 | 204.04
M16A | A | 2 | 12 | 1 | -77.187 | 200.84
M16A | A | 2 | 13 | 1 | -75.836 | 200.88
M16A | A | 2 | 14 | 1 | -76.787 | 200.56
M16A | A | 2 | 15 | 1 | -73.631 | 195.42
M16A | A | 2 | 16 | 1 | -62.663 | 200.49
M16B | A | 2 | 1 | 1 | -64.479 | 200.09 | 202.99
M16B | A | 2 | 2 | 1 | -69.184 | 208.58
M16B | A | 2 | 3 | 1 | -74.164 | 208.73
M16B | A | 2 | 4 | 1 | -81.092 | 211.25
M16B | A | 2 | 5 | 1 | -72.013 | 207.54
M16B | A | 2 | 6 | 1 | -75.338 | 205.27
M16B | A | 2 | 7 | 1 | -70.718 | 207.3
M16B | A | 2 | 8 | 1 | -67.664 | 194.3
M16B | A | 2 | 9 | 1 | -67.647 | 202.86
M16B | A | 2 | 10 | 1 | -75.537 | 203.68
M16B | A | 2 | 11 | 1 | -67.448 | 200.1
M16B | A | 2 | 12 | 1 | -57.975 | 201.52
M16B | A | 2 | 13 | 1 | -58.439 | 198.56
M16B | A | 2 | 14 | 1 | -57.691 | 199.48
M16B | A | 2 | 15 | 1 | -56.635 | 192.69
M16B | A | 2 | 16 | 1 | -69.357 | 193.74
I4 | B | 1 | 1 | 1 | -36.7782 | 1120.1433
I4 | B | 1 | 1 | 2 | -37.1527 | 1119.6034
I4 | B | 1 | 2 | 1 | -36.7711 | 1116.2875
I4 | B | 1 | 2 | 2 | -36.4048 | 1117.127
I4 | B | 1 | 3 | 1 | -36.7848 | 1121.6512
I4 | B | 1 | 3 | 2 | -37.2235 | 1121.1677
I4 | B | 1 | 4 | 1 | -36.7916 | 1125.348
I4 | B | 1 | 4 | 2 | -36.3992 | 1125.9045
I4 | B | 1 | 5 | 1 | -36.8624 | 1127.4719
I4 | B | 1 | 5 | 2 | -37.2313 | 1126.8328
I4 | B | 1 | 6 | 1 | -36.6649 | 1098.3864
I4 | B | 1 | 6 | 2 | -36.3394 | 1099.3755
I4 | B | 1 | 7 | 1 | -37.0415 | 1136.7551
I4 | B | 1 | 7 | 2 | -37.3585 | 1136.0967
I4 | B | 1 | 8 | 1 | -36.9276 | 1123.0999
I4 | B | 1 | 8 | 2 | -36.5902 | 1123.9934
I4 | B | 1 | 9 | 1 | -36.9446 | 1133.6074
I4 | B | 1 | 9 | 2 | -37.2748 | 1133.0311
I4 | B | 1 | 10 | 1 | -36.8383 | 1120.9153
I4 | B | 1 | 10 | 2 | -36.5439 | 1121.8998
I4 | B | 1 | 11 | 1 | -36.9795 | 1137.5219
I4 | B | 1 | 11 | 2 | -37.3623 | 1137.0135
I4 | B | 1 | 12 | 1 | -36.7938 | 1107.5374
I4 | B | 1 | 12 | 2 | -36.4111 | 1108.2595
I4 | B | 1 | 13 | 1 | -36.8944 | 1117.2559
I4 | B | 1 | 13 | 2 | -37.3157 | 1116.6416
I4 | B | 1 | 14 | 1 | -36.8247 | 1116.4142
I4 | B | 1 | 14 | 2 | -36.3778 | 1117.2623
I4 | B | 1 | 15 | 1 | -36.6294 | 1085.3473
I4 | B | 1 | 15 | 2 | -37.0738 | 1084.7209
I4 | B | 1 | 16 | 1 | -36.9107 | 1116.2115
I4 | B | 1 | 16 | 2 | -36.4627 | 1116.905
I4 | B | 1 | 17 | 1 | -37.1055 | 1141.2451
I4 | B | 1 | 17 | 2 | -37.4726 | 1140.759
I4 | B | 1 | 18 | 1 | -37.1224 | 1137.5367
I4 | B | 1 | 18 | 2 | -36.8092 | 1138.3716
I4 | B | 1 | 19 | 1 | -37.2179 | 1138.2306
I4 | B | 1 | 19 | 2 | -37.498 | 1137.4341
I4 | B | 1 | 20 | 1 | -37.2723 | 1150.9421
I4 | B | 1 | 20 | 2 | -36.8859 | 1151.496
I4 | B | 1 | 21 | 1 | -37.3277 | 1148.8253
I4 | B | 1 | 21 | 2 | -37.6554 | 1148.1376
I4 | B | 1 | 22 | 1 | -37.4018 | 1157.6473
I4 | B | 1 | 22 | 2 | -37.0899 | 1158.3376
I4 | B | 1 | 23 | 1 | -37.1871 | 1130.9607
I4 | B | 1 | 23 | 2 | -37.5567 | 1130.2797
I4 | B | 1 | 24 | 1 | -37.2323 | 1131.8582
I4 | B | 1 | 24 | 2 | -36.8594 | 1132.7474
I4 | B | 1 | 25 | 1 | -37.0962 | 1115.8594
I4 | B | 1 | 25 | 2 | -37.3979 | 1115.2341
I4 | B | 1 | 26 | 1 | -37.168 | 1126.0586
I4 | B | 1 | 26 | 2 | -36.8309 | 1126.8795
I4 | B | 1 | 27 | 1 | -37.6421 | 1158.1964
I4 | B | 1 | 27 | 2 | -37.9254 | 1157.5929
I4 | B | 1 | 28 | 1 | -37.3233 | 1137.98
I4 | B | 1 | 28 | 2 | -36.9287 | 1138.5836
I4 | B | 1 | 29 | 1 | -37.1657 | 1120.9795
I4 | B | 1 | 29 | 2 | -37.5116 | 1120.3256
I4 | B | 1 | 30 | 1 | -37.4449 | 1148.5876
I4 | B | 1 | 30 | 2 | -37.1081 | 1149.3783
I4 | B | 1 | 31 | 1 | -37.2478 | 1138.0793
I4 | B | 1 | 31 | 2 | -37.6138 | 1137.5127
I4 | B | 1 | 32 | 1 | -37.4761 | 1140.7316
I4 | B | 1 | 32 | 2 | -37.1113 | 1141.2457
I4 | B | 2 | 1 | 1 | -36.7783 | 1119.4231
I4 | B | 2 | 1 | 2 | -37.1389 | 1118.7944
I4 | B | 2 | 2 | 1 | -36.6656 | 1114.6442
I4 | B | 2 | 2 | 2 | -36.2827 | 1115.4783
I4 | B | 2 | 3 | 1 | -36.7403 | 1120.439
I4 | B | 2 | 3 | 2 | -37.1352 | 1119.8472
I4 | B | 2 | 4 | 1 | -36.7709 | 1123.7061
I4 | B | 2 | 4 | 2 | -36.3827 | 1124.4481
I4 | B | 2 | 5 | 1 | -36.8507 | 1126.1848
I4 | B | 2 | 5 | 2 | -37.2715 | 1125.7157
I4 | B | 2 | 6 | 1 | -36.7082 | 1097.1755
I4 | B | 2 | 6 | 2 | -36.2592 | 1097.698
I4 | B | 2 | 7 | 1 | -36.944 | 1135.2289
I4 | B | 2 | 7 | 2 | -37.2928 | 1134.6696
I4 | B | 2 | 8 | 1 | -36.8482 | 1121.8709
I4 | B | 2 | 8 | 2 | -36.469 | 1122.5874
I4 | B | 2 | 9 | 1 | -36.9491 | 1132.2534
I4 | B | 2 | 9 | 2 | -37.2765 | 1131.4443
I4 | B | 2 | 10 | 1 | -36.8165 | 1119.6625
I4 | B | 2 | 10 | 2 | -36.4688 | 1120.4381
I4 | B | 2 | 11 | 1 | -36.861 | 1135.6764
I4 | B | 2 | 11 | 2 | -37.2505 | 1135.1975
I4 | B | 2 | 12 | 1 | -36.7139 | 1106.0859
I4 | B | 2 | 12 | 2 | -36.3729 | 1106.7604
I4 | B | 2 | 13 | 1 | -36.8946 | 1116.0253
I4 | B | 2 | 13 | 2 | -37.2858 | 1115.3922
I4 | B | 2 | 14 | 1 | -36.8741 | 1115.0396
I4 | B | 2 | 14 | 2 | -36.3226 | 1115.6047
I4 | B | 2 | 15 | 1 | -36.5801 | 1083.8964
I4 | B | 2 | 15 | 2 | -37.064 | 1083.3088
I4 | B | 2 | 16 | 1 | -36.9046 | 1114.6534
I4 | B | 2 | 16 | 2 | -36.5323 | 1115.6128
I4 | B | 2 | 17 | 1 | -37.143 | 1139.7943
I4 | B | 2 | 17 | 2 | -37.5645 | 1139.3464
I4 | B | 2 | 18 | 1 | -37.1281 | 1136.0044
I4 | B | 2 | 18 | 2 | -36.6906 | 1136.6005
I4 | B | 2 | 19 | 1 | -37.0001 | 1136.3311
I4 | B | 2 | 19 | 2 | -37.3105 | 1135.8585
I4 | B | 2 | 20 | 1 | -37.1436 | 1149.3424
I4 | B | 2 | 20 | 2 | -36.856 | 1150.2487
I4 | B | 2 | 21 | 1 | -37.26 | 1147.3265
I4 | B | 2 | 21 | 2 | -37.6996 | 1147.0098
I4 | B | 2 | 22 | 1 | -37.3333 | 1156.095
I4 | B | 2 | 22 | 2 | -36.9307 | 1156.7476
I4 | B | 2 | 23 | 1 | -37.1664 | 1129.5474
I4 | B | 2 | 23 | 2 | -37.5936 | 1129.1134
I4 | B | 2 | 24 | 1 | -37.1391 | 1130.4164
I4 | B | 2 | 24 | 2 | -36.7164 | 1131.0383
I4 | B | 2 | 25 | 1 | -37.0127 | 1114.3462
I4 | B | 2 | 25 | 2 | -37.3749 | 1113.6627
I4 | B | 2 | 26 | 1 | -37.2214 | 1124.8447
I4 | B | 2 | 26 | 2 | -36.9147 | 1125.5647
I4 | B | 2 | 27 | 1 | -37.4463 | 1156.3943
I4 | B | 2 | 27 | 2 | -37.7932 | 1155.8453
I4 | B | 2 | 28 | 1 | -37.2861 | 1136.4071
I4 | B | 2 | 28 | 2 | -36.8654 | 1136.9763
I4 | B | 2 | 29 | 1 | -37.2015 | 1119.66
I4 | B | 2 | 29 | 2 | -37.4824 | 1118.9651
I4 | B | 2 | 30 | 1 | -37.4217 | 1146.9668
I4 | B | 2 | 30 | 2 | -37.1173 | 1147.6761
I4 | B | 2 | 31 | 1 | -37.3488 | 1136.9624
I4 | B | 2 | 31 | 2 | -37.7221 | 1136.4944
I4 | B | 2 | 32 | 1 | -37.449 | 1139.1622
I4 | B | 2 | 32 | 2 | -37.2098 | 1139.9205
I5 | B | 1 | 1 | 1 | -27.928 | 120.1928
I5 | B | 1 | 1 | 2 | -31.1069 | 118.419
I5 | B | 1 | 2 | 1 | -27.5359 | 112.994
I5 | B | 1 | 2 | 2 | -24.2958 | 114.8761
I5 | B | 1 | 3 | 1 | -28.0667 | 121.6297
I5 | B | 1 | 3 | 2 | -31.2202 | 119.7476
I5 | B | 1 | 4 | 1 | -27.5099 | 113.5774
I5 | B | 1 | 4 | 2 | -24.2569 | 115.45
I5 | B | 1 | 5 | 1 | -28.1227 | 122.7457
I5 | B | 1 | 5 | 2 | -31.2982 | 120.8905
I5 | B | 1 | 6 | 1 | -27.6593 | 114.8577
I5 | B | 1 | 6 | 2 | -24.5242 | 116.6837
I5 | B | 1 | 7 | 1 | -28.2308 | 121.3219
I5 | B | 1 | 7 | 2 | -31.4405 | 119.5006
I5 | B | 1 | 8 | 1 | -27.5283 | 113.2807
I5 | B | 1 | 8 | 2 | -24.3838 | 115.1348
I5 | B | 1 | 9 | 1 | -27.9295 | 119.6149
I5 | B | 1 | 9 | 2 | -31.1937 | 117.7667
I5 | B | 1 | 10 | 1 | -27.4023 | 110.1593
I5 | B | 1 | 10 | 2 | -24.0393 | 112.0782
I5 | B | 1 | 11 | 1 | -27.844 | 118.4273
I5 | B | 1 | 11 | 2 | -31.1409 | 116.5865
I5 | B | 1 | 12 | 1 | -27.2022 | 108.3858
I5 | B | 1 | 12 | 2 | -23.9156 | 110.317
I5 | B | 1 | 13 | 1 | -27.9531 | 118.0308
I5 | B | 1 | 13 | 2 | -31.2077 | 116.1346
I5 | B | 1 | 14 | 1 | -27.1122 | 106.749
I5 | B | 1 | 14 | 2 | -23.5269 | 108.7825
I5 | B | 1 | 15 | 1 | -27.7544 | 114.6112
I5 | B | 1 | 15 | 2 | -31.25 | 112.6248
I5 | B | 1 | 16 | 1 | -27.025 | 103.9521
I5 | B | 1 | 16 | 2 | -23.2321 | 106.0821
I5 | B | 1 | 17 | 1 | -27.7449 | 115.5258
I5 | B | 1 | 17 | 2 | -31.2696 | 113.5414
I5 | B | 1 | 18 | 1 | -27.3547 | 108.89
I5 | B | 1 | 18 | 2 | -24.0315 | 110.7436
I5 | B | 1 | 19 | 1 | -27.7548 | 114.1287
I5 | B | 1 | 19 | 2 | -31.2918 | 112.1274
I5 | B | 1 | 20 | 1 | -27.1375 | 106.7515
I5 | B | 1 | 20 | 2 | -23.7592 | 108.6527
I5 | B | 1 | 21 | 1 | -27.8111 | 115.3253
I5 | B | 1 | 21 | 2 | -31.6215 | 113.1561
I5 | B | 1 | 22 | 1 | -27.157 | 106.4643
I5 | B | 1 | 22 | 2 | -23.3264 | 108.6017
I5 | B | 1 | 23 | 1 | -27.6608 | 113.9051
I5 | B | 1 | 23 | 2 | -30.9645 | 112.0368
I5 | B | 1 | 24 | 1 | -27.1908 | 106.7732
I5 | B | 1 | 24 | 2 | -24.051 | 108.5985
I5 | B | 1 | 25 | 1 | -27.6892 | 112.7094
I5 | B | 1 | 25 | 2 | -30.8686 | 110.8612
I5 | B | 1 | 26 | 1 | -27.081 | 104.4817
I5 | B | 1 | 26 | 2 | -23.8519 | 106.3826
I5 | B | 1 | 27 | 1 | -27.683 | 112.8377
I5 | B | 1 | 27 | 2 | -31.0625 | 110.8829
I5 | B | 1 | 28 | 1 | -27.1215 | 105.488
I5 | B | 1 | 28 | 2 | -23.7833 | 107.3721
I5 | B | 1 | 29 | 1 | -27.6617 | 112.0409
I5 | B | 1 | 29 | 2 | -30.9729 | 110.1444
I5 | B | 1 | 30 | 1 | -27.0117 | 103.81
I5 | B | 1 | 30 | 2 | -23.6817 | 105.7408
I5 | B | 1 | 31 | 1 | -27.0914 | 105.5842
I5 | B | 1 | 31 | 2 | -30.3962 | 103.6464
I5 | B | 1 | 32 | 1 | -26.9021 | 101.8412
I5 | B | 1 | 32 | 2 | -23.3623 | 103.8234
I5 | B | 2 | 1 | 1 | -28.0798 | 119.8553
I5 | B | 2 | 1 | 2 | -31.2591 | 118.0543
I5 | B | 2 | 2 | 1 | -27.5593 | 112.5366
I5 | B | 2 | 2 | 2 | -24.3429 | 114.443
I5 | B | 2 | 3 | 1 | -28.1004 | 121.0515
I5 | B | 2 | 3 | 2 | -31.2774 | 119.1982
I5 | B | 2 | 4 | 1 | -27.563 | 113.125
I5 | B | 2 | 4 | 2 | -24.3105 | 114.9985
I5 | B | 2 | 5 | 1 | -28.0565 | 122.2208
I5 | B | 2 | 5 | 2 | -31.2844 | 120.3672
I5 | B | 2 | 6 | 1 | -27.6292 | 114.3257
I5 | B | 2 | 6 | 2 | -24.4788 | 116.1272
I5 | B | 2 | 7 | 1 | -28.1578 | 120.848
I5 | B | 2 | 7 | 2 | -31.3445 | 119.0166
I5 | B | 2 | 8 | 1 | -27.6088 | 112.7717
I5 | B | 2 | 8 | 2 | -24.4757 | 114.605
I5 | B | 2 | 9 | 1 | -27.9909 | 118.9865
I5 | B | 2 | 9 | 2 | -31.2183 | 117.1485
I5 | B | 2 | 10 | 1 | -27.2665 | 109.5719
I5 | B | 2 | 10 | 2 | -23.8796 | 111.4807
I5 | B | 2 | 11 | 1 | -28.0298 | 117.9529
I5 | B | 2 | 11 | 2 | -31.2832 | 116.0842
I5 | B | 2 | 12 | 1 | -27.2187 | 107.9053
I5 | B | 2 | 12 | 2 | -23.9162 | 109.8296
I5 | B | 2 | 13 | 1 | -27.8853 | 117.4244
I5 | B | 2 | 13 | 2 | -31.1656 | 115.5444
I5 | B | 2 | 14 | 1 | -27.2403 | 106.3996
I5 | B | 2 | 14 | 2 | -23.6531 | 108.4089
I5 | B | 2 | 15 | 1 | -27.7253 | 114.0396
I5 | B | 2 | 15 | 2 | -31.2458 | 112.0534
I5 | B | 2 | 16 | 1 | -26.9981 | 103.4394
I5 | B | 2 | 16 | 2 | -23.2199 | 105.5883
I5 | B | 2 | 17 | 1 | -27.8076 | 115.0453
I5 | B | 2 | 17 | 2 | -31.2831 | 113.0101
I5 | B | 2 | 18 | 1 | -27.3819 | 108.4017
I5 | B | 2 | 18 | 2 | -24.1112 | 110.2892
I5 | B | 2 | 19 | 1 | -27.7665 | 113.6608
I5 | B | 2 | 19 | 2 | -31.3466 | 111.6745
I5 | B | 2 | 20 | 1 | -27.1833 | 106.2264
I5 | B | 2 | 20 | 2 | -23.8371 | 108.1282
I5 | B | 2 | 21 | 1 | -27.7399 | 114.6879
I5 | B | 2 | 21 | 2 | -31.531 | 112.5253
I5 | B | 2 | 22 | 1 | -27.1493 | 105.9374
I5 | B | 2 | 22 | 2 | -23.3619 | 108.1355
I5 | B | 2 | 23 | 1 | -27.7579 | 113.4123
I5 | B | 2 | 23 | 2 | -31.0135 | 111.5118
I5 | B | 2 | 24 | 1 | -27.2032 | 106.314
I5 | B | 2 | 24 | 2 | -24.0187 | 108.1149
I5 | B | 2 | 25 | 1 | -27.71 | 112.1328
I5 | B | 2 | 25 | 2 | -30.8522 | 110.2778
I5 | B | 2 | 26 | 1 | -27.1417 | 104.024
I5 | B | 2 | 26 | 2 | -23.8275 | 105.8766
I5 | B | 2 | 27 | 1 | -27.6062 | 112.267
I5 | B | 2 | 27 | 2 | -30.9742 | 110.2971
I5 | B | 2 | 28 | 1 | -27.1531 | 105.0628
I5 | B | 2 | 28 | 2 | -23.817 | 106.9742
I5 | B | 2 | 29 | 1 | -27.5826 | 111.5029
I5 | B | 2 | 29 | 2 | -30.9029 | 109.6216
I5 | B | 2 | 30 | 1 | -27.0593 | 103.2917
I5 | B | 2 | 30 | 2 | -23.7726 | 105.2659
I5 | B | 2 | 31 | 1 | -27.2105 | 105.1375
I5 | B | 2 | 31 | 2 | -30.5685 | 103.2381
I5 | B | 2 | 32 | 1 | -26.9321 | 101.3733
I5 | B | 2 | 32 | 2 | -23.4755 | 103.3882
M12 | B | 1 | 1 | 1 | -31.4122 | 752.1047
M12 | B | 1 | 2 | 1 | -34.5447 | 1170.5544
M12 | B | 1 | 3 | 1 | -34.7034 | 1189.6683
M12 | B | 1 | 4 | 1 | -34.5879 | 1164.8315
M12 | B | 1 | 5 | 1 | -34.8357 | 1205.7439
M12 | B | 1 | 6 | 1 | -34.4697 | 1147.6815
M12 | B | 1 | 7 | 1 | -34.7712 | 1192.3652
M12 | B | 1 | 8 | 1 | -34.5647 | 1153.2363
M12 | B | 1 | 9 | 1 | -34.9594 | 1199.7111
M12 | B | 1 | 10 | 1 | -34.5149 | 1144.3942
M12 | B | 1 | 11 | 1 | -35.0207 | 1196.855
M12 | B | 1 | 12 | 1 | -34.4505 | 1124.8451
M12 | B | 1 | 13 | 1 | -35.1122 | 1199.1901
M12 | B | 1 | 14 | 1 | -34.6878 | 1145.0515
M12 | B | 1 | 15 | 1 | -35.2126 | 1205.4437
M12 | B | 1 | 16 | 1 | -34.8416 | 1150.062
M12 | B | 2 | 1 | 1 | -31.3967 | 751.2662
M12 | B | 2 | 2 | 1 | -34.4844 | 1168.8931
M12 | B | 2 | 3 | 1 | -34.6801 | 1187.9214
M12 | B | 2 | 4 | 1 | -34.5137 | 1163.1462
M12 | B | 2 | 5 | 1 | -34.8409 | 1204.1764
M12 | B | 2 | 6 | 1 | -34.4872 | 1146.2914
M12 | B | 2 | 7 | 1 | -34.7484 | 1190.6891
M12 | B | 2 | 8 | 1 | -34.5712 | 1151.7599
M12 | B | 2 | 9 | 1 | -34.9614 | 1198.0797
M12 | B | 2 | 10 | 1 | -34.4946 | 1142.8026
M12 | B | 2 | 11 | 1 | -35.0494 | 1195.3569
M12 | B | 2 | 12 | 1 | -34.4649 | 1123.4689
M12 | B | 2 | 13 | 1 | -35.1242 | 1197.8584
M12 | B | 2 | 14 | 1 | -34.6918 | 1143.4977
M12 | B | 2 | 15 | 1 | -35.1623 | 1203.7415
M12 | B | 2 | 16 | 1 | -34.8023 | 1148.4189
M13 | B | 1 | 1 | 1 | -31.3412 | 603.0335
M13 | B | 1 | 2 | 1 | -30.9611 | 583.7075
M13 | B | 1 | 3 | 1 | -31.4632 | 607.8647
M13 | B | 1 | 4 | 1 | -31.0277 | 586.4572
M13 | B | 1 | 5 | 1 | -31.5212 | 611.0825
M13 | B | 1 | 6 | 1 | -31.0932 | 587.9059
M13 | B | 1 | 7 | 1 | -31.49 | 608.0892
M13 | B | 1 | 8 | 1 | -31.0352 | 580.6669
M13 | B | 1 | 9 | 1 | -31.5815 | 607.3298
M13 | B | 1 | 10 | 1 | -31.3959 | 588.8161
M13 | B | 1 | 11 | 1 | -31.6313 | 610.9218
M13 | B | 1 | 12 | 1 | -31.2749 | 587.6672
M13 | B | 1 | 13 | 1 | -31.6866 | 607.9882
M13 | B | 1 | 14 | 1 | -31.4877 | 589.296
M13 | B | 1 | 15 | 1 | -31.734 | 605.6603
M13 | B | 1 | 16 | 1 | -31.4645 | 585.7847
M13 | B | 2 | 1 | 1 | -31.3242 | 602.397
M13 | B | 2 | 2 | 1 | -30.9392 | 582.998
M13 | B | 2 | 3 | 1 | -31.4346 | 607.0803
M13 | B | 2 | 4 | 1 | -31.0203 | 585.7809
M13 | B | 2 | 5 | 1 | -31.4709 | 610.2278
M13 | B | 2 | 6 | 1 | -31.087 | 587.1647
M13 | B | 2 | 7 | 1 | -31.4922 | 607.3029
M13 | B | 2 | 8 | 1 | -31.0365 | 579.9844
M13 | B | 2 | 9 | 1 | -31.5736 | 606.5844
M13 | B | 2 | 10 | 1 | -31.3655 | 588.1218
M13 | B | 2 | 11 | 1 | -31.654 | 610.2268
M13 | B | 2 | 12 | 1 | -31.303 | 587.0511
M13 | B | 2 | 13 | 1 | -31.6487 | 607.2008
M13 | B | 2 | 14 | 1 | -31.512 | 588.7314
M13 | B | 2 | 15 | 1 | -31.6873 | 604.8758
M13 | B | 2 | 16 | 1 | -31.4466 | 585.0582
M14 | B | 1 | 1 | 1 | -53.9468 | 180.9991
M14 | B | 1 | 2 | 1 | -54.3991 | 181.469
M14 | B | 1 | 3 | 1 | -53.7855 | 176.5046
M14 | B | 1 | 4 | 1 | -54.0539 | 178.6885
M14 | B | 1 | 5 | 1 | -54.1898 | 179.0739
M14 | B | 1 | 6 | 1 | -54.0302 | 178.4838
M14 | B | 1 | 7 | 1 | -53.949 | 177.3337
M14 | B | 1 | 8 | 1 | -53.7049 | 174.2797
M14 | B | 1 | 9 | 1 | -54.0012 | 174.7276
M14 | B | 1 | 10 | 1 | -53.6526 | 172.9474
M14 | B | 1 | 11 | 1 | -53.7787 | 171.2954
M14 | B | 1 | 12 | 1 | -53.7033 | 172.9551
M14 | B | 1 | 13 | 1 | -54.091 | 175.6912
M14 | B | 1 | 14 | 1 | -53.5095 | 169.3129
M14 | B | 1 | 15 | 1 | -53.8616 | 172.5492
M14 | B | 1 | 16 | 1 | -54.0209 | 176.4656
M14 | B | 2 | 1 | 1 | -53.8295 | 179.5766
M14 | B | 2 | 2 | 1 | -54.1949 | 179.9683
M14 | B | 2 | 3 | 1 | -53.6706 | 175.0537
M14 | B | 2 | 4 | 1 | -53.9267 | 177.2497
M14 | B | 2 | 5 | 1 | -54.1523 | 177.6436
M14 | B | 2 | 6 | 1 | -53.9527 | 177.0368
M14 | B | 2 | 7 | 1 | -53.9474 | 175.9171
M14 | B | 2 | 8 | 1 | -53.5962 | 172.8392
M14 | B | 2 | 9 | 1 | -53.8705 | 173.2274
M14 | B | 2 | 10 | 1 | -53.6473 | 171.5636
M14 | B | 2 | 11 | 1 | -53.5917 | 169.8197
M14 | B | 2 | 12 | 1 | -53.5991 | 171.4653
M14 | B | 2 | 13 | 1 | -53.9016 | 174.1465
M14 | B | 2 | 14 | 1 | -53.4484 | 167.9073
M14 | B | 2 | 15 | 1 | -53.7233 | 170.9967
M14 | B | 2 | 16 | 1 | -53.8192 | 174.9454
M15 | B | 1 | 1 | 1 | -24.7813 | 163.6953
M15 | B | 1 | 2 | 1 | -24.6303 | 160.4512
M15 | B | 1 | 3 | 1 | -24.8057 | 163.1087
M15 | B | 1 | 4 | 1 | -24.5679 | 158.7882
M15 | B | 1 | 5 | 1 | -24.5851 | 158.7263
M15 | B | 1 | 6 | 1 | -24.3201 | 153.7172
M15 | B | 1 | 7 | 1 | -24.3516 | 153.9265
M15 | B | 1 | 8 | 1 | -24.3361 | 154.2554
M15 | B | 1 | 9 | 1 | -24.4758 | 155.8459
M15 | B | 1 | 10 | 1 | -24.3067 | 152.045
M15 | B | 1 | 11 | 1 | -24.3574 | 152.6883
M15 | B | 1 | 12 | 1 | -24.1186 | 147.9234
M15 | B | 1 | 13 | 1 | -24.2031 | 149.935
M15 | B | 1 | 14 | 1 | -24.0716 | 146.0589
M15 | B | 1 | 15 | 1 | -24.0834 | 147.1836
M15 | B | 1 | 16 | 1 | -24.0367 | 143.8817
M15 | B | 2 | 1 | 1 | -24.7168 | 162.6024
M15 | B | 2 | 2 | 1 | -24.5893 | 159.3376
M15 | B | 2 | 3 | 1 | -24.7396 | 161.9699
M15 | B | 2 | 4 | 1 | -24.5485 | 157.6999
M15 | B | 2 | 5 | 1 | -24.5575 | 157.6482
M15 | B | 2 | 6 | 1 | -24.295 | 152.6975
M15 | B | 2 | 7 | 1 | -24.3431 | 152.8928
M15 | B | 2 | 8 | 1 | -24.3439 | 153.2099
M15 | B | 2 | 9 | 1 | -24.4327 | 154.7434
M15 | B | 2 | 10 | 1 | -24.2429 | 150.9448
M15 | B | 2 | 11 | 1 | -24.3029 | 151.6128
M15 | B | 2 | 12 | 1 | -24.1332 | 146.8621
M15 | B | 2 | 13 | 1 | -24.1973 | 148.8479
M15 | B | 2 | 14 | 1 | -24.0631 | 144.9819
M15 | B | 2 | 15 | 1 | -24.1104 | 146.0881
M15 | B | 2 | 16 | 1 | -23.9891 | 142.7845

## Sheet2
Average BCS data (BCS=300 K)
UAID | Band | HAM | Subsample | dn | std | gain | NEdT
3103277 | I4 | 1 | 17 | 1 | 1 | 519.3614 | 2.3093 | 1272.1484 | 17.3497 |  | 0.32584 | 3103276 | I4 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I4 | 1 | 17 | 2 | 1 | 519.3629 | 2.3102 | 1271.7794 | 17.3373 |  | 0.32607 | 3103276 | I4 | 1 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | I4 | 2 | 17 | 1 | 1 | 518.9291 | 2.3093 | 1271.1982 | 17.2983 |  | 0.32609 | 3103276 | I4 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I4 | 2 | 17 | 2 | 1 | 518.9319 | 2.3093 | 1270.7711 | 17.2936 |  | 0.32619 | 3103276 | I4 | 2 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 1 | 17 | 1 | 1 | 1573.4361 | 2.1686 | 185.0368 | 8.4177 |  | 0.26894 | 3103276 | I5 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 1 | 17 | 2 | 1 | 1573.4139 | 2.1677 | 184.956 | 8.3854 |  | 0.26895 | 3103276 | I5 | 1 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 2 | 17 | 1 | 1 | 1564.1527 | 2.1688 | 184.5076 | 8.4197 |  | 0.26973 | 3103276 | I5 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 2 | 17 | 2 | 1 | 1564.1298 | 2.1679 | 184.4501 | 8.3748 |  | 0.26971 | 3103276 | I5 | 2 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | M12 | 1 | 17 | 1 | 1 | 495.8663 | 0.75607 | 1323.8891 | 28.4766 |  | 0.11219 | 3103276 | M12 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M12 | 2 | 17 | 1 | 1 | 495.464 | 0.75595 | 1322.9391 | 28.4925 |  | 0.11225 | 3103276 | M12 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M13 | 1 | 17 | 1 | 1 | 494.0626 | 0.90755 | 672.288 | 12.734 |  | 0.043471 | 3103276 | M13 | 1 | 17 | 1 | 25.927 | 0.1577 | 278.9856 | 216.918 | 1.6075
3103277 | M13 | 2 | 17 | 1 | 1 | 493.7023 | 0.90722 | 671.873 | 12.7112 |  | 0.043482 | 3103276 | M13 | 2 | 17 | 1 | 25.9068 | 0.15759 | 278.8381 | 216.8077 | 1.6081
3103277 | M14 | 1 | 17 | 1 | 1 | 1748.1545 | 0.9706 | 209.18 | 7.0114 |  | 0.039177 | 3103276 | M14 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M14 | 2 | 17 | 1 | 1 | 1728.7729 | 0.97121 | 207.8148 | 7.0802 |  | 0.039461 | 3103276 | M14 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M15 | 1 | 17 | 1 | 1 | 1667.9107 | 0.5996 | 190.6126 | 6.4627 |  | 0.021654 | 3103276 | M15 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M15 | 2 | 17 | 1 | 1 | 1653.1947 | 0.59916 | 189.8149 | 6.4804 |  | 0.02173 | 3103276 | M15 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16AB | 1 | 17 | 1 | 1 | 1809.2112 | 0.57169 | 217.4821 | 7.6104 |  | 0.021701 | 3103276 | M16A | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16AB | 2 | 17 | 1 | 1 | 1801.6762 | 0.57156 | 216.6339 | 7.7154 |  | 0.021783 | 3103276 | M16A | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
[12 empty] | 3103276 | M16B | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
Average OBC data (OBC=292 K) | [11 empty] | 3103276 | M16B | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I4 | 1 | 17 | 1 | 1 | 373.5611 | 3.1751 | 1279.5963 | 17.2794 | 0.44536
3103277 | I4 | 1 | 17 | 2 | 1 | 373.5145 | 3.1561 | 1279.0914 | 17.3352 | 0.44289
3103277 | I4 | 2 | 17 | 1 | 1 | 373.1777 | 3.1701 | 1278.5308 | 17.273 | 0.44503 |  | 3103276 | I4 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I4 | 2 | 17 | 2 | 1 | 373.2692 | 3.1715 | 1278.4995 | 17.2696 | 0.44525 |  | 3103276 | I4 | 1 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 1 | 17 | 1 | 1 | 1406.8831 | 2.9393 | 185.6589 | 8.4367 | 0.36336 |  | 3103276 | I4 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 1 | 17 | 2 | 1 | 1402.3718 | 2.9723 | 184.9813 | 8.4049 | 0.36888 |  | 3103276 | I4 | 2 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 2 | 17 | 1 | 1 | 1397.5903 | 2.9424 | 185.1409 | 8.442 | 0.36478 |  | 3103276 | I5 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | I5 | 2 | 17 | 2 | 1 | 1393.1138 | 2.9721 | 184.49 | 8.3986 | 0.36983 |  | 3103276 | I5 | 1 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | M12 | 1 | 17 | 1 | 1 | 354.6213 | 0.93347 | 1328.7607 | 28.5665 | 0.13805 |  | 3103276 | I5 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M12 | 2 | 17 | 1 | 1 | 354.2989 | 0.93385 | 1327.8896 | 28.5971 | 0.13819 |  | 3103276 | I5 | 2 | 17 | 2 | 0 | 0 | 0 | 0 | 0
3103277 | M13 | 1 | 17 | 1 | 0 | 4.2895 | 0.21737 | 7.9551 | 0.16438 | 0.88037 |  | 3103276 | M12 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M13 | 1 | 17 | 1 | 1 | 364.2323 | 0.82832 | 675.4875 | 12.8126 | 0.039485 |  | 3103276 | M12 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M13 | 2 | 17 | 1 | 0 | 4.2851 | 0.21738 | 7.9489 | 0.15866 | 0.88105 |  | 3103276 | M13 | 1 | 17 | 1 | 4.2385 | 0.19684 | 7.867 | 0.19932 | 0.80592
3103277 | M13 | 2 | 17 | 1 | 1 | 363.9464 | 0.82657 | 675.1279 | 12.7907 | 0.039422 |  | 3103276 | M13 | 2 | 17 | 1 | 4.2412 | 0.1962 | 7.8733 | 0.18068 | 0.80256
3103277 | M14 | 1 | 17 | 1 | 1 | 1497.3256 | 1.2244 | 210.0872 | 7.0521 | 0.04921 |  | 3103276 | M14 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M14 | 2 | 17 | 1 | 1 | 1477.8692 | 1.2251 | 208.6842 | 7.1275 | 0.049575 |  | 3103276 | M14 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M15 | 1 | 17 | 1 | 1 | 1475.8761 | 0.75861 | 190.8749 | 6.4591 | 0.027359 |  | 3103276 | M15 | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M15 | 2 | 17 | 1 | 1 | 1461.0849 | 0.75815 | 190.0847 | 6.479 | 0.027456 |  | 3103276 | M15 | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16A | 1 | 17 | 1 | 1 | 1622.523 | 0.91251 | 220.2132 | 8.1154 | 0.03425 |  | 3103276 | M16A | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16A | 2 | 17 | 1 | 1 | 1614.9202 | 0.91223 | 219.7529 | 8.1351 | 0.034312 |  | 3103276 | M16A | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16B | 1 | 17 | 1 | 1 | 1623.2982 | 0.90272 | 220.2111 | 6.7206 | 0.033834 |  | 3103276 | M16B | 1 | 17 | 1 | 0 | 0 | 0 | 0 | 0
3103277 | M16B | 2 | 17 | 1 | 1 | 1615.8584 | 0.90299 | 219.7672 | 6.7314 | 0.033913 |  | 3103276 | M16B | 2 | 17 | 1 | 0 | 0 | 0 | 0 | 0