NICST_MEMO_11_003

NICST Internal Memo

Date: February 25, 2011

From: Jeff McIntire

To: Bruce Guenther, Jim Butler, and Jack Xiong

Subject: Plateau and Electronics Side Comparison of Thermal Band Radiometric Calibration from VIIRS F1 RC-05 Testing

---------------------------------------------------------------------------------------------------------

References:

[1] ‘Sensor Performance Verification Plan,’ PVP154640-101.

[2] ‘Performance Specification Sensor Specification,’ ps154640-101c.

[3] NICST_MEMO_11_001, ‘Analysis of the Radiometric Calibration from VIIRS F1 RC-05 Part 1 Test (Nominal Plateau, Electronic Side B),’ Jeff McIntire and Sanxiong Xiong, January 3, 2011.

[4] NICST_MEMO_11_002, ‘Analysis of the Radiometric Calibration from VIIRS F1 RC-05 Part 2 Test (Nominal Plateau),’ Jeff McIntire and Sanxiong Xiong, January 31, 2011.

[5] IT_OBCBB_EDU_Temp_081112, ‘OBC Blackbody Temperature Calculation,’ Courtney Ranshaw, November 11, 2008.

1.  Introduction

During the thermal vacuum testing of VIIRS F1 sensor, RC-05 was performed under a variety of instrument conditions. RC-05 Parts 1 and 2 were conducted at three instrument temperature plateaus (Cold, Nominal, and Hot). In addition, RC-05 Part 1 was tested using both electronics sides A and B at each plateau. The purpose of this memo is to compare and trend the thermal band performance across different instrument conditions, including compliance with a number of specifications [1,2].

The methodology, processing, and results of RC-05 at Nominal plateau, electronic side B have been reported for both Parts 1 and 2 [3,4]. Tables 1 - 7 list the data used in this work. Tables 1 – 6 contain the data from RC-05 Part 1 (all plateaus and electronics sides) used in the high gain determination on the left and the low gain determination on the right. All UAIDs, collects, gain modes, and relevant source temperatures are listed. Table 7 lists the RC-05 Part 2 data for both warm up and cool down at all three instrument temperature plateaus. Note that Part 2 data taken at Cold plateau was on electronics side A and data taken at Nominal and Hot plateaus was on electronics side B. The results shown in this work are from HAM side A (HAM sides A and B yield similar results).

2.  Analysis

2.1  RC-05 Part 1 High Gain (BCS source)

The radiometric coefficient a0 for each high gain band is plotted in Figure 1 versus detector for all plateaus and electronics sides. In general, the a0 coefficient is on the order of 10-2 or less and is roughly consistent over plateaus and electronic sides. M14 has the largest offset of around 0.01 – 0.02 for all detectors. The detector dependence is fairly stable versus instrument condition for the LWIR; some variation is observed in the MWIR. In addition, the detector 1 offset seems to be Out Of Family (OOF) in the LWIR. In Figure 2, the a0 coefficients for the all detectors in each band are graphed versus TOMM temperature (solid lines for electronics side A and dashed lines for electronics side B). There is a good deal of detector variation in the a0 trends versus TOMM. However, it should be noted that the variation of the offset with plateau is on the order of 10-2 or less. Figure 3 shows the same trend versus TOMM for a middle detector (8 for M bands and 16 for I bands) along with 2 sigma error bars (NIST standard). Note that in general the error bars overlap for the different plateaus (electronics side dependent), and as a result, a0 is consistent over instrument temperature conditions. For almost all bands and detectors, the error bars overlap across all instrument conditions.

Figure 4 contains the gains (1/a1) for all high gain bands (all plateaus and electronics sides). The detector to detector patterns are very consistent from plateau to plateau and across electronics side. Also, note that strong odd – even detector dependence is observed in I5, M12, and M13 and that M12 detector 1 is consistently OOF. The gain trends versus TOMM are shown in Figure 5. The gains trend fairly linearly over plateau. In addition, the gains tend to decrease with TOMM, with the exception of M13, which registers a slight increase. In general, the LWIR bands decrease by about 2 % or less and the MWIR bands change by less than 1 % over the ~16 K TOMM range. Figure 6 shows the gain trend for a middle detector versus TOMM along with the 2 sigma error bars. Here the temperature variation is outside the error bars in most cases, particularly for the LWIR bands. There is some overlap between plateaus for a given electronics side for I4 and M13, but the electronics sides are not consistent.

The nonlinear term a2 is plotted in Figure 7 for all high gain bands (across plateaus and electronics sides). For the LWIR bands, the detector pattern of the nonlinear term is roughly constant over plateau and electronics side; however, for the MWIR, the pattern is less well defined. In terms of magnitude, the nonlinear term is consistently below 10-8. In Figure 8, the trending versus TOMM for a2 is shown (all detectors). For most bands and electronics sides, the nonlinear term exhibits a small, roughly linear increase with TOMM (less well defined for the MWIR bands). There is some noticeable detector variation in the trending for the MWIR. The a2 trend versus TOMM is plotted in Figure 9 for a middle detector along with the 2 sigma error bars. For the majority of cases, the error bars overlap. The most glaring exceptions are in I4, M12, and M14.

The NEdT at TTYP for high gain bands, plateaus, and electronics sides is shown in Figure 10. The NEdT at TTYP is very consistent over the range of instrument conditions tested, both in terms of magnitude and detector dependence. Note that detectors 31 in I5, 1 in M12, and 8 in M16A are consistently OOF. Figure 11 shows the dependence on TOMM for the NEdT at TTYP (all detectors). The NEdT generally increases with TOMM for all bands and both electronics sides (although there is a good deal of detector dependence). This is the result of the increasing dark noise in the detectors which occurs at higher instrument temperatures. The NEdT at TTYP increases about 10 % over the TOMM range of ~16 K (except for M12 which increases roughly 30 %). The NEdT as a function of BCS temperature is graphed in Figure 12 for all high gain bands and instrument conditions (median detector). Again, the NEdT is consistent across instrument conditions, excepting the slight increase with instrument temperature due to increasing dark noise. Note that only the data within the optimized dynamic range is plotted here [3] (the optimized M14 dynamic range begins at 210 K, which was only measured once).

The nonlinearity metric (RRNL) is graphed versus detector for all instrument conditions in Figure 13. The RRNL is generally consistent over instrument conditions. The exception is M14; the Nominal plateau, electronics side B set includes a larger set of BCS levels including 210 K (at all other instrument conditions, this data point is not measured). As a result, the fitting for M14 is more tightly constrained for Nominal plateau, electronics side B.

The ARD for all bands and instrument conditions is shown in Figure 14 (median detector). The ARD are roughly consistent over all plateaus and electronics sides for the optimized dynamic range. In the MWIR, there is a constant bias of about 0.3 % which results from the application of the scan by scan OBC correction to the retrieved radiance (see [3]). In addition, the shape that is present in the MWIR is largely the result of the radiometric fitting.

The potential for detector to detector striping was measured by the RRU, graphed for all bands, plateaus, and electronics sides in Figure 15 (worst case detector). A RRU greater then one indicates the potential for striping (dashed red line in the plots). The RRU generally increases with increasing scene temperature. For bands I4, M12, M13, M14, and M15, some striping is possible at the highest scene temperatures. Note that the specification is only valid between the specified LMIN and 0.9LMAX. Most of the evidence for striping occurs above that range. At higher temperatures, the deviation of the retrieved radiance from the band average increases, but the measured NEdL levels off; the result is a steadily increasing RRU.

2.2  RC-05 Part 1 Low Gain (TMC source)

The M13 low gain coefficients for all bands, plateaus, and electronics sides are shown in Figure 16. The detector patterns are consistent over instrument conditions for both a1 and a2; in contrast, no discernable pattern exists for a0, which is generally consistent with zero. The temperature trends for each coefficient (all detectors) are graphed in Figure 17. For electronics side B, 1/a1 tends to decrease linearly with TOMM and a2 tends to increase with TOMM. However, no such dependence is evident for electronics side A. The gain varies by up to 5 % over instrument conditions while a2 is on the 10-7 level. The coefficient trends versus TOMM with 2 sigma error bars (middle detector) are shown in Figure 18. In general, the error bars overlap for a0, but not for a1 or a2. Note that this uncertainty is in part due to the difficulty in cross calibrating the TMC to the BCS and to the fact that only five scene temperatures were measured for the low gain dynamic range (optimized).

The NEdT is plotted in Figure 19 for M13 low gain. In the left hand plot, the NEdT at TTYP is shown versus detector for all instrument conditions. The NEdT at TTYP is not very consistent from detector to detector or from instrument condition to instrument condition. This results from the fact that TTYP is 380 K for M13 low gain, and that this is a portion of the dynamic range that was under-sampled during testing. In consequence, the NEdT at TTYP is difficult to constrain. Nonetheless, all detectors meet the requirement (red dashed line) for all instrument conditions. The right hand plot in Figure 19 shows the NEdT as a function of effective TMC temperature (the blackbody temperature associated with the retrieved TMC radiance). The NEdT is between 0.1 and 0.2 K above about 450 K; below this the NEdT increases to as much as 0.5 K near 350 K.

Figure 20 shows some of the performance metrics associated with M13 low gain. In the left hand plot, the nonlinearity metric is graphed versus detector for all instrument conditions. The nonlinearity is consistent over all plateaus and electronics sides, and is well within the specified limit (dashed red line). The middle graph contains the M13 low gain ARD versus effective TMC temperature (worst case detector). The ARD is consistent over instrument conditions above about 450 K; below this there is some discrepancy between the data collections. In the rightmost plot, the RRU is graphed versus effective TMC temperature (worst case detector). Some potential striping is evident at lower instrument temperatures in the region around 500 K, but is not observed at higher instrument temperatures.

2.3  RC-05 Part 2 High Gain (OBC source)

The OBC warm up and cool down cycle was performed at all three instrument temperature plateaus. However, it was only measured on one electronics side per plateau: A side for Cold plateau and B side for Nominal and Hot plateaus. It should be noted that the number of OBC levels is not the same for warm up and cool down (cool down has three less levels at lower temperature than warm up for the same plateau). In addition, the number of levels changes from plateau to plateau: warm up at Cold plateau has one more level then warm up at Nominal plateau and two more levels than warm up at Hot plateau, all at lower temperatures (similarly for cool down). As a result, any plateau comparison must take into account the fact that the temperature range used in the fitting was not consistent across RC-05 Part 2 measurements.

The coefficient a0 derived from the OBC warm up is plotted versus detector for all high gain bands and plateaus in Figure 21. The pattern over detectors is not consistent from plateau to plateau. The additional levels at colder instrument temperatures are a large constraint on the fitting, especially the offset term. Figure 22 shows the trend versus TOMM for all bands (median detector). In most cases the relationship is nonlinear, but is within the expanded (2 sigma) error bars.

The OBC warm up gain is graphed versus detector for all high gain bands and plateaus in Figure 23. The detector pattern is consistent over the three plateaus. The median detector gain trend versus TOMM is shown in Figure 24. In general, the gains have a roughly linear relationship with TOMM; one exception is M12, which appears to have a more nonlinear trend. The gain variation with TOMM is roughly less than 1 % (MWIR) or 2 % (LWIR) over about 16 K. In general, the gains are not consistent over instrument conditions (excepting I4, I5, and M15).

Figure 25 shows the detector variation in the coefficient a2 for all high gain bands and plateaus. As with a0, the detector to detector variation in a2 is generally not consistent from plateau to plateau. As noted before, the additional low temperature levels are a large constraint on the fitting, especially affecting both the offset and nonlinear terms (particularly in the MWIR). The median detector a2 trend versus TOMM is plotted in Figure 26 and shows a nonlinear relationship for all bands. However, the trend is within the 2 sigma expanded uncertainties for all bands.

In contrast to the OBC warm up gains, Figure 27 graphs the OBC cool down gains versus detector for all plateaus. The detector patterns are less well defined compared the warm up gains (see Figure 23). The number of levels available in the cool down portion of the cycle is three less than the warm up at any given plateau. In addition, the extra levels that are available in warm up are at lower temperatures. Thus, the ability to constrain the fitting in the cool down cycle is significantly reduced.

The NEdT at TTYP derived from both OBC warm up and cool down is shown for all bands versus detector in Figure 28. The NEdT at TTYP is generally consistent over plateau and electronics side. In the case of I5, Hot plateau warm up shows some deviation from the other data collections; for I5, TTYP is 210 K (well below the measured range) and the extrapolation may not be well constrained. In addition, a slight increase in NEdT over instrument temperature is observed (consistent with increasing dark noise).

Figure 29 plots the BCS ARD for both warm up and cool down for all three plateaus versus TOBC. In general, the warm up and cool down ARD tend to decrease with increasing TOBC. In addition, the warm up ARD trend slightly higher with instrument temperature. The cool down ARD are offset from the warm up ARD by about 0.1 % in the MWIR and 0.05 % in the LWIR.

The deviation from the specified OBC temperature setpoint is shown in Figure 30 as a function of TOBC. The dashed red lines indicate the allowed tolerance of ±0.2 K (SRV0654). The deviation exceeds the tolerance as TOBC decreases below 292 K for all plateaus. This is in part due to the updated OBC temperature calculation [5], which was not used to set the OBC temperature during testing. This correction increases as TOBC decreases. In Figure 31, the standard deviation of the temperatures extracted from the six thermistors imbedded in the OBC is plotted versus TOBC. The specification requires this standard deviation to be below 0.03 K (SRV0095) for temperature controlled conditions (warm-up). The standard deviation meets this requirement for all scene temperatures and plateaus.

3.  Summary

- External source (BCS) – high gain

  - The a0 coefficients are on the order of 10-2 or less and are consistent over both electronics side and instrument temperature plateau (within 2 sigma).

  - The a2 coefficients are consistently below 10-8 and are largely within the 2 sigma errors.

  - In general, the gains trend lower with increased instrument temperature (about 2 % or less). The exception is M13, which exhibits a slight increase. The trends are largely outside the 2 sigma errors.

  - Detector to detector differences in the gain are consistent over all instrument conditions. Odd – even dependence is observed in I5, M12, and M13.

  - The NEdT at TTYP is consistent across instrument conditions, excepting the slight increase with instrument temperature due to increasing dark noise (roughly a 10 % increase over the measured range).

  - The nonlinearity metric is consistent over all instrument conditions.

  - The ARD is consistent over all instrument conditions within the dynamic range. There is a small bias in the MWIR bands due to the scan by scan OBC correction.

  - The potential for detector to detector striping increases with scene temperature consistently over instrument conditions, with striping resolved from the noise observed at the highest measured scene temperatures for bands I4, M12, M13, M14, and M15.

- External source (TMC) – low gain

  - The coefficient a0 is consistent with zero for all measurements. The gain varies by up to 5 % (which is outside the 2 sigma uncertainty). The a2 coefficient increases with instrument temperature and is generally inconsistent over instrument conditions.

  - The NEdT at TTYP is not very consistent across instrument conditions, but is nonetheless below the specified limit.

- Internal source (OBC) – high gain

  - The warm-up a0 and a2 coefficients are consistent over plateaus within the expanded uncertainty.

  - The warm-up gains are also consistent over instrument conditions within the 2 sigma uncertainty (except for bands M12 and M13). The gains trend lower by about 2 % (except for bands M12 and M13, which trend higher by about 1 %). Cool-down gains are less well defined.

  - The NEdT at TTYP is consistent over instrument plateaus (excepting a slight increase due to increasing dark noise) and well within the specification.

  - The BCS ARD tends to decrease slightly with increasing TOBC. In addition, the cool-down ARD is slightly higher than the warm-up.

  - The standard deviation of the six OBC thermistors is below 0.03 K for all instrument conditions.

Acknowledgement

The sensor test data used in this document was provided by the Raytheon El Segundo testing team. Approaches for data acquisition and data reductions, as well as data extraction tools were also provided by the Raytheon El Segundo team. We would like to thank the Raytheon El Segundo team for their support. The data analysis tools were developed by the NICST team.

Table 1: F1 RC-05 Part 1 data (Cold plateau, electronics side A)

Table 2: F1 RC-05 Part 1 data (Nominal plateau, electronics side A)

Table 3: F1 RC-05 Part 1 data (Hot plateau, electronics side A)

Table 4: F1 RC-05 Part 1 data (Cold plateau, electronics side B)

Table 5: F1 RC-05 Part 1 data (Nominal plateau, electronics side B)

Table 6: F1 RC-05 Part 1 data (Hot plateau, electronics side B)

Table 7: F1 RC-05 Part 2 data

Table 8: Specifications addressed in this work [2]

Table 9: Sensitivity requirements for the emissive bands [2]

Table 10: Absolute radiometric calibration uncertainty of the spectral radiance for the emissive bands [2]

Figure 1: a0 coefficient versus detector

Figure 2: a0 coefficient versus TOMM (all detectors)

Figure 3: a0 versus TOMM (middle detector with 2 sigma error bars)

Figure 4: 1/a1 coefficient versus detector

Figure 5: 1/a1 coefficient versus TOMM (all detectors)

Figure 6: 1/a1 versus TOMM (middle detector with 2 sigma error bars)

Figure 7: a2 coefficient versus detector

Figure 8: a2 coefficient versus TOMM (all detectors)

Figure 9: a2 versus TOMM (middle detector with 2 sigma error bars)

Figure 10: NEdT at TTYP versus detector

Figure 11: NEdT at TTYP versus TOMM (all detectors)

Figure 12: NEdT versus TBCS (median detector)

Figure 13: RRNL versus detector

Figure 14: ARD versus TBCS (median detector)

Figure 15: RRU versus TBCS (median detector)

Figure 16: M13 low gain coefficients versus detector

Figure 17: M13 low gain coefficients versus TOMM (all detectors)

Figure 18: M13 low gain coefficients versus TOMM (middle detector with 2 sigma error bars)

Figure 19: NEdT at TTYP versus detector as well as NEdT versus TTMC (median detector)

Figure 20: RRNL versus detector as well as ARD and RRU versus TTMC (median detector)

Figure 21: a0 coefficient versus detector from warm up

Figure 22: a0 coefficient versus TOMM from warm up (median detector)

Figure 23: 1/a1 coefficient versus detector from warm up

Figure 24: 1/a1 coefficient versus TOMM from warm up (median detector)

Figure 25: a2 coefficient versus detector from warm up

Figure 26: a2 coefficient versus TOMM from warm up (median detector)

Figure 27: 1/a1 coefficient versus detector from cool down

Figure 28: NEdT at TTYP versus detector

Figure 29: BCS ARD versus TOBC (median detector)

Figure 30: Deviation of TOBC from OBC setpoint

Figure 31: Standard Deviation of the six OBC thermistors versus TOBC