VCST Internal Memo

Title: Updated Estimate of Radiometric Uncertainty for the S-NPP Thermal Bands On-orbit

Memo Number: 2014_002

Revision: 01

Date: February 25, 2014

Author: Jeff McIntire and Boryana Efremova

To: Xiaoxiong Xiong and James Butler

Cc: Hassan Oudrari, Kwo-Fu (Vincent) Chiang, Jon Fulbright and Aisheng Wu

# References

- [1] An Introduction to Error Analysis, John R. Taylor, University Science Books, 1997.

- [2] NICST_MEMO_11_014, ‘Uncertainty Estimate for the Thermal Band Radiometric Retrievals,’ Jeff McIntire, October 5, 2011.

- [3] VCST_MEMO_2013_002, ‘Estimate of Radiometric Uncertainty for the S-NPP Thermal Bands On-orbit,’ Jeff McIntire and Boryana Efremova, September 5, 2013.

- [4] NICST_MEMO_11_001, ‘Analysis of the Radiometric Calibration from VIIRS F1 RC-05 Part 1 Test (Nominal Plateau, Electronic Side B),’ Jeff McIntire and Sanxiong Xiong, January 3, 2011.

- [5] ‘Performance Verification Report – VIIRS FU1 Emissive Band Calibration (PVP Section 4.2.4),’ Karen Galang, December 4, 2009, revision E.

- [6] VCST_TECH_REPORT_12_028, ‘TEB On-orbit Calibration During 09/2012 BB WUCD for VIIRS F1,’ Boryana Efremova, Jeff McIntire, and Aisheng Wu, October 2, 2012.

- [7] PL3095-N06593, ‘Absolute Radiometric Calibration Thermal Region: Uncertainty Estimates,’ Jim Young, 1997.

- [8] NICST_MEMO_08_008, ‘Response vs. Scan Angle for VIIRS FU1 TEB (FP-10, part 2),’ Sanxiong Xiong, Chunhui Pan, and Nianzeng Che, March 11, 2008.

- [9] VCST_TECH_MEMO_2012_029, ‘Calibrating NPP VIIRS M13LG fire detection band by comparing M13LG and M13HG images of the Moon in the Space View,’ Boryana Efremova and Jeff McIntire, December 3, 2012.

- [10] ‘NPP Visible/infrared Imager Radiometer Suite (VIIRS) Radiance Uncertainty, Emissive Bands – Tested Performance’, E. Johnson, K. Galang, C. Ranshaw, and B. Robinson, Proc of SPIE 7808, 78081F (2010).

1. Introduction

VIIRS thermal band SDR products (TOA radiance and brightness temperature) are used in a number of different science products. It is important to estimate the uncertainty of these measurements for science team evaluations of the down stream products. This work will focus on propagating the uncertainty inherent in the S-NPP VIIRS TEB measurements to the SDR product level. The uncertainty is propagated using the standard formulation [1], as is described in the following section. Results of the error propagation are investigated for different Earth scene temperatures and conditions (scan angle and aggregation mode).

The uncertainty estimation contained in this work follows the general methodology established at the time of sensor level TV testing [2]. Here the uncertainties are first propagated into the radiometric calibration coefficients determined pre-launch; then these coefficients are used to investigate the uncertainty of the SDR products on-orbit. This work is an update to [3], with improvements in traceability and consistency in the model.

2. Error Propagation

For the purposes of this work, we follow the standard propagation of error for a function y of variables xi is described by [1]

u2y=i=1N∂y∂xi2u2xi+2i=1N−1j=i+1N∂y∂xi∂y∂xjuxi,xj.				(1)

Here u(xi) is the uncertainty of the variable xi that goes into the calculation of the y and u(xi,xj) is the covariance between xi and xj. The Earth View (EV) retrieved radiance was defined in [4] and is a function of

LretEV=f(c0,c1,c2,LBB,LHAM,LRTA,LCAV,LSH,FSH,FCAV,FRTA,RVSBB,RVSEV,RVSSV,εBB,ρRTA,dnEV,dnBB).												(2)

The partial derivatives for each variable are listed in Appendix A for the path difference radiance viewing an external blackbody pre-launch (BCS) and in Appendix B for the EV retrieved radiance on-orbit. Note that the total uncertainty here is for a single EV pixel; the effects of aggregation and scan angle will also be investigated in Section 4.

In general, the covariance terms were not directly calculated (the exceptions are the covariance terms between the radiometric coefficients); a direct calculation of these terms is beyond the scope of this work. However, an upper bound on the covariance terms is determined through use of the Schwarz inequality [1], or

uxi,xj≤uxiuxj.									(3)

Note that when the Schwarz inequality is used in conjunction with Eq. (1), the absolute value of the partial derivatives in the covariance terms are used. Results will be presented without covariance terms (excepting those terms between the radiometric coefficients) as a baseline and with covariance terms determined using the Schwarz inequality as a worst case estimate. Note that covariance terms are only included in the worst case estimate if the sources of uncertainty are considered interdependent.

3. Individual Error Sources

In this section we describe the individual uncertainty contributors. The descriptions follow from [2] with some additions unique to on-orbit operations.

3.1 Radiances

The radiance uncertainty for each of the radiances that factor into the present calculation (LBB, LHAM, LRTA, LSH, and LCAV) is the RSS of the uncertainty contributors. Each of these radiances was converted from a temperature reading provided by one or more thermistors once per scan using the Planck equation, integrated over the spectral response of the instrument. In the cases where more than one thermistor was used, the average was employed in the Planck equation. The error for each is composed of four components: temperature, spectral, interpolation, and statistical uncertainties. The spectral and interpolation uncertainties are considered biases, the statistical uncertainty is considered random, and the temperature uncertainty is a combination of both for the BB and purely bias for all other temperatures.

The radiance uncertainty associated with temperature error was determined by taking the absolute value of the maximum difference between the Planck radiance at a given temperature and at the ends of the temperature uncertainty interval, or

uL=maxLT,λ−LT±ΔT,λ.							(4)

The temperature uncertainties used in the pre-launch analysis are listed in Table 1 [5]. An additional temperature uncertainty of the maximum standard deviation over the six Blackbody (BB) thermistors over an orbit is included on-orbit; this accounts for observed variability in the BB temperature uniformity on-orbit [6] (an effect which was not observed pre-launch).

The spectral errors used here are listed in Table 2 [7]. The radiance uncertainties due to spectral errors were determined in the same manner as the temperature uncertainties,

uL=maxLT,λ−LT,λ±Δλ.							(5)

On-orbit, the calibration relies on LUTs to convert source temperature to radiance. A linear interpolation is used to determine the radiance at the actual source temperature every scan. The interpolation uncertainty is determined by the following,

uL=12T−T0T−T1f2T.							(6)

where T0 and T1 are the LUT values bracketing the measured T and the second derivative is of Planck’s radiation law with respect to the measured T. Since the LUTs are only used on-orbit, these uncertainty contributors are included in the estimation of the retrieved radiance uncertainty but not in the radiometric calibration coefficients.

The statistical uncertainties were the standard deviation of the radiances determined within one granule (for the retrieved radiance) or within one collect (for the calibration coefficients).

3.2 Response Versus Scan (RVS)

The uncertainties for the RVS factors were determined in [8]. The on-orbit RVS was normalized relative to the solar diffuser HAM angle. The RVS uncertainty derived from FP-10 was a combination of fitting error and measurement error [8]; the band averaged values used in this work are given in Table 3 (where the average of the M16A and M16B uncertainties was used for M16). The application of these uncertainties is considered bias. Any scan angle dependence to the uncertainty was not considered and no uncertainty due to emission versus scan was included in the present work.

3.3 BB Reflectance Shape Factors, Reflectance of the RTA, and Emissivity of the BB

The uncertainties for the BB reflectance shape factors are all taken to be 100% (these terms were considered bias uncertainties here). The uncertainty for the reflectance of the RTA is 0.5% at 270 K [6]. This uncertainty was used for all source temperature levels. The BB emissivity was measured at 3.39 μm; the error in the emissivity measurement was determined to be 0.07% [5]. This uncertainty was used for all thermal bands. The RTA reflectance and BB emissivity uncertainties used in this work were considered biases.

3.4 EV and BB Response

The uncertainty in the response was the RSS of the random and bias errors for the background subtracted digital response. The precision error was the standard deviation of the mean over all analyzed samples and scans; for the purposes of this work:

- The random error for a single EV pixel cannot be calculated directly on-orbit. A fit of standard deviations over the BCS temperature range derived from pre-launch data collected in instrument TV at nominal plateau was used to estimate the random error in a given EV pixel as a function of detector signal. The effect of aggregation zone upon the on-orbit response uncertainty was accounted for by an additional factor of 1/√2 or 1/√3 in the 2:1 and 3:1 aggregation zones, respectively. The random error for the single EV pixel used in the pre-launch coefficient determination was standard deviation of the mean of all samples used.

- The random error for the BB view or Space View (SV) was the standard deviation of the mean over the 48 samples per scan and all scans (in a collect for the pre-launch data and in a granule for the on-orbit data).

The total bias error was the RSS of the known biases. Any bias common to all sectors would be removed in the background subtraction. The known biases between sectors are listed below [5]:

- The EV response was reported in a truncated 12 bit format, whereas the calibration sectors were reported in 14 bits; this creates a bias of about 0.375 dn in the SV relative to the EV. To compensate, all calibration view data can be truncated to 12 bits at the beginning of the processing. The bias is then common to both sectors and is eliminated in the background subtraction.

- There is an additional bias in the EV data due to aggregation. In a three pixel aggregation zone, each pixel is truncated to 12 bits. Then the three pixels are averaged, and the average is truncated to 12 bits. A similar process was used for the two pixel aggregation zone. To compensate in the analysis, all calibration view data can be processed in a manner similar to the EV data. The bias is then common to both sectors and is eliminated in the background subtraction.

- For M13, there is a small bias between fixed and auto gain configurations (a few tenths of a DN). Currently, the cause has not been determined [5]. This bias was not included in the present analysis.

- There is also a bias in the EV data for M16 TDI. Each M16A and M16B detector response is first truncated to 12 bits. Then the average of the two detector responses is taken, and that average is again truncated to 12 bits. To compensate in the analysis, all calibration view data can be processed in a manner similar to the EV data. The bias is then common to both sectors and is eliminated in the background subtraction.

As this analysis was conducted at model dn values (as described below), all the bias errors that can be removed via processing were not included in this analysis. The exception is the M13 bias between auto and fixed gain modes; as this bias is not currently understood, it was not included in this analysis.

3.5 Radiometric Coefficients

The vertical least-squares fitting algorithm used in this memo determined the vertical deviations of the set of data points from the fit, or

R2=i1u2ΔLiΔLi−c0−c1dni−c2dni22.						(7)

The minimum of the vertical deviations was computed by setting the partial derivatives with respect to coefficients equal to zero, or

∂R2∂ci=0.										(8)

This led to the following matrix equation:

c0c1c2=1u2dnu2dn2u2dnu2dn2u2dn3u2dn2u2dn3u2dn4u2−1ΔLu2dnΔLu2dn2ΔLu2,					(9)

and we define

A=1u2dnu2dn2u2dnu2dn2u2dn3u2dn2u2dn3u2dn4u2−1.							(10)

The solution to this matrix equation determined the radiometric coefficients using pre-launch data collected during instrument level TV testing. This algorithm also produced 1-sigma uncertainties and covariance terms, which are defined by the following:

uci=Ai,i,										(11)

uci,cj=Ai,j.									(12)

The uncertainties in ΔL were estimated using Eq. (1) and the partial derivatives in Appendix A. The individual contributors were listed in the preceding subsections.

Furthermore, it should be noted that the uncertainties are valid only insomuch as the model itself is valid. The quality of the radiometric fitting for the thermal bands has been discussed in previous work [2]. For the purposes of this memo, the radiometric model is considered sufficiently valid to proceed with the uncertainty analysis.

4. Analysis

The uncertainty propagation was performed for each band at a series of model dn levels (5, 20, 100, 250, 500, 750, 1000, 1500, 2000, 2500, 3000, and 3500). The radiometric coefficients used in this work were derived from the pre-launch data collected during instrument level TV testing (nominal plateau, electronics side B). The uncertainties for these coefficients were also determined from the pre-launch data as described above. The on-orbit retrieved radiance uncertainties were then determined using Eq. (1), the partial derivatives in Appendix B, and the individual contributors listed above.

The exception is M13 low gain, where uncertainties roughly the difference between the derived on-orbit coefficients from lunar analysis [9] and the operational coefficients are used in conjunction with the operational coefficients. The lunar analysis derived coefficients are in general agreement with the pre-launch coefficients derived in [4]. As the offset is consistent with zero, an uncertainty of 100% is used for the operational offset uncertainty. An uncertainty of 7.4% is used for the operational linear coefficient. The nonlinear term is neglected (its contribution is negligible and cannot be derived on-orbit).

Aggregation zone and scan angle effects were also investigated. The aggregation zone influence was estimated by including an additional factor of 1/√2 or 1/√3 in the EV dn uncertainty. The scan angle effects were investigated by determining the uncertainty at the following scan angles: -54.2, -35.4, 0.0, 35.4, and 54.2 degrees off nadir.

In [2], it was postulated that propagating the uncertainty at the retrieved radiance level includes some double counting, as the errors in the source path difference radiance are propagated into the fitting coefficient uncertainties as well as contribute at the retrieved radiance level. In that work, the uncertainties in the BCS path difference radiance were propagated into the coefficients errors, and not included at the product level error propagation. In the current work, the uncertainties present in the radiometric coefficients were included at the product level error propagation as well with the weights defined by the partial derivatives in Appendix A.

5. Results

Figures 1 – 8 plot the individual uncertainty contributors in % of the retrieved radiance for each band (detector 9, HAM side A) as a function of scene temperature. For the MWIR, the c2 and BB radiance terms are the most important at higher temperatures and the EV dn term dominates at lower temperatures, increasing rapidly as the temperature decreases (note the BB is nominally controlled at ~ 292 K). In the LWIR, no terms are particularly dominant at higher temperatures, while the EV dn, c0, and RVS terms predominate at the lowest scene temperatures (again increasing rapidly). The case of M13 LG shows that the largest terms are the radiometric coefficients (c0 at lower temperature and c1 at higher temperatures).

The total uncertainties per band for detector 9, HAM side A are shown in Figure 9 (in % of radiance) and Figure 10 (in K) versus scene temperature. All bands exhibit roughly constant uncertainty at higher scene temperatures. In Figure 8, the bands show less than ~1 % radiance uncertainty above scene temperatures ~280 K (I4), ~260 K (M12 and M13), ~230 K (I5 and M14), and ~210 K (M15 and M16). In terms of temperature uncertainty, all bands show less that ~0.5 K above scene temperatures ~260 K (I4), ~ 240 K (M12 and M13), ~230 K (I5 and M14), and ~210 K (M15 and M16). The total radiance uncertainties for M13 LG in %, and temperature uncertainties in K are shown in Figures 11 and 12, respectively. The radiance uncertainty is roughly constant at ~7.5 % at higher scene temperatures (due mainly to c1) and rises below 400 K to roughly 25 % (largely as the result of c0); in terms of temperature uncertainty, M13 LG exhibits uncertainties of about 8 K near the 350 K, decreasing to about 4 K between 400 and 450 K, and then increasing slowly to about 8 K at scene temperatures of 650 K.

The total uncertainties are compared to the specifications in Tables 4 and 5 (in terms of % of radiance and temperature [K], respectively) for all bands, detector 9, HAM side A. There are some non-compliances at lower scene temperatures for bands M12 and M13; however, in all cases, the non-compliances are small (within 0.19 K). It should be noted that pre-launch analysis showed these bands to be compliant (but in general only marginally for the non-compliances reported here) [2,10]. The differences from pre-launch uncertainty estimates are largely attributed to the estimation of the EV dn uncertainty. Pre-launch this term was estimated from an extended, external source over many scans; post-launch, it was estimated from fitting the pre-launch data.

The variation with scan angle (and aggregation mode) was also calculated as seen in Figures 13 – 19 for all bands (detector 9, HAM side A). Note that the specification is shown with a dashed red line. The MWIR bands show dependence on aggregation zone, where the interior aggregation zones have lower uncertainty. This is the result of the lower uncertainty in the EV dn contribution (included here as the decrease in uncertainty associated with averaging 2 or 3 pixels: 1/√2 or 1/√3). In M14, the uncertainty exhibits a dependence on scan angle, where lower uncertainty is observed at end of scan. This difference decreases from beginning of scan to end of scan (the RVS varies by ~10 % for M14). The other LWIR bands do not show much variation (some small scan angle dependence is observed).

Figures 20 and 21 show the total uncertainties for both the baseline (solid lines) and worst-case (dashed lines) approaches. The worst-case uncertainties are always larger than the baseline, with the smallest difference at ~292 K. Below this temperature, the worst-case uncertainty is about 0.2 K larger. Above 292 K, the differences increase to about 0.3 K at scene temperatures of 350 K. The worst-case uncertainty likely represents an overestimate as covariance terms must be smaller than the Schwarz inequality and may even be negative.

6. Summary

Uncertainty estimates for the thermal band radiometry retrievals were determined by propagating the error estimates from the individual uncertainty sources. This uncertainty model was applied to the case of a single EV pixel. The following is a list of findings:

- Total uncertainties for each band were compared to specified uncertainty limits at a series of scene temperatures. For most cases, the estimated uncertainties were below the limits. However, some lower scene temperatures showed uncertainties above the specification for M12 and M13.

- In general, the modeled uncertainty estimates were in good agreement with pre-launch estimates. Some lower scene temperatures had slightly higher uncertainties, resulting in non-compliances with the specification that were absent in the pre-launch calculation.

- MWIR EV retrieved radiance uncertainties are between 0.1 – 0.3 K at high scene temperatures and increase rapidly below 250 K for M12 and M13 and 270 K for I4. LWIR EV retrieved radiance uncertainties are between 0.1 – 0.4 K above scene temperatures 210 – 230 K.

- Dominant terms in MWIR EV uncertainties are c2 (important at high temperatures), BB radiance error (roughly constant with scene temperature), and EV dn uncertainties (increasing rapidly with decreasing scene temperature). No term was dominant in the LWIR EV uncertainties at higher scene temperatures whiles c0, RVS, and EV dn uncertainties all increased at lower scene temperatures.

- MWIR band uncertainties showed decreased uncertainty at lower scene temperatures for the interior aggregation zones. M14 uncertainties showed some dependence with scan angle at lower scene temperatures.

- M13 LG uncertainties were estimated to be between 4 and 8 K; the uncertainties were dominated by the c0 and c1 terms (lower and higher scene temperatures respectively).

Appendix A

The following are the partial derivatives of the BCS view path difference radiance with respect to the various contributors:

∂ΔLBCS∂RVSBCS=LBCS−1ρRTALHAM−1−ρRTALRTA

∂ΔLBCS∂RVSSV=1ρRTALHAM−1−ρRTALRTA

∂ΔLBCS∂LBCS=RVSBCS

∂ΔLBCS∂LHAM=RVSSV−RVSBCSρRTA

∂ΔLBCS∂LRTA=−RVSSV−RVSBCSρRTA1−ρRTA

∂ΔLBCS∂ρRTA=−RVSSV−RVSBCSρRTA2LHAM−LRTA

Appendix B

The following are the partial derivatives of the retrieved EV radiance with respect to the various contributors:

∂LretEV∂c0=ΔLBBRVSEVc1dnBB−dnEV+c2dnBB2−dnEV2c0+c1dnBB+c2dnBB22

∂LretEV∂c1=ΔLBBRVSEVc0dnEV−dnBB+c2dnbB2dnEV−dnEV2dnBBc0+c1dnBB+c2dnBB22

∂LretEV∂c2=ΔLBBRVSEVc0dnEV2−dnBB2+c1dnEV2dnBB−dnBB2dnEVc0+c1dnBB+c2dnBB22

∂LretEV∂RVSBB=εBBLBB+1−εBBFSHLSH+FCAVLCAV+FRTALRTA−1ρRTALHAM−1−ρRTALRTAc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2∂LretEV∂RVSEV=−ΔLBBc0+c1dnEV+c2dnEV2RVSEV2c0+c1dnBB+c2dnBB2−RVSSVρRTARVSEV2LHAM−1−ρRTALRTA

∂LretEV∂RVSSV=1ρRTALHAM−1−ρRTALRTA1+c0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂LRTA=1−εBBRVSBBFRTA−RVSSV−RVSBBρRTA1−ρRTAc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2−RVSSV−RVSEVRVSEVρRTA1−ρRTA

∂LretEV∂LBB=εBBRVSBBc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂LSH=1−εBBRVSBBFSHc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂LCAV=1−εBBRVSBBFCAVc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂LHAM=RVSSV−RVSBBρRTAc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2+RVSSV−RVSEVRVSEVρRTA

∂LretEV∂FSH=1−εBBRVSBBLSHc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂FCAV=1−εBBRVSBBLCAVc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂FRTA=1−εBBRVSBBLRTAc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂εBB=RVSBBLBB+FSHLSH+FCAVLCAV+FRTALRTAc0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2

∂LretEV∂ρRTA=−LHAM−LRTAρRTA2c0+c1dnEV+c2dnEV2RVSEVc0+c1dnBB+c2dnBB2RVSSV−RVSBB+RVSSV−RVSEVRVSEV∂LretEV∂dnEV=ΔLOBCRVSEVc1+2c2dnEVc0+c1dnBB+c2dnBB2

∂LretEV∂dnBB=ΔLOBCRVSEVc1+2c2dnBBc0+c1dnEV+c2dnEV2c0+c1dnBB+c2dnBB22

Table 1: Temperature biases [5]

| Source | Temperature Bias (K) |
| --- | --- |
| BCS | 0.05 |
| BB | 0.04 |
| HAM | 1.0 |
| RTA | 9.0 |
| SH | 3.0 |
| CAV | 6.0 |

Table 2: Spectral biases [7]

| Band | Spectral Bias (nm) |
| --- | --- |
| I4 | 1.2 |
| I5 | 4.0 |
| M12 | 1.2 |
| M13 | 1.2 |
| M14 | 4.0 |
| M15 | 4.0 |
| M16 | 4.0 |

Table 3: RVS uncertainties [8]

| Band | RVS uncertainty |
| --- | --- |
| I4 | 0.000811 |
| I5 | 0.000986 |
| M12 | 0.000818 |
| M13 | 0.000798 |
| M14 | 0.001003 |
| M15 | 0.000875 |
| M16 | 0.000782 |

Table 4: Total uncertainties compared to specifications in % of radiance

| Band | Scene Temperature |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 190 | 230 | 267 | 270 | 310 | 340 |
| I4 spec | -- | -- | 5.0 | -- | -- | -- |
| I4 | -- | -- | 2.57 | -- | -- | -- |
| I5 spec | -- | -- | 2.5 | -- | -- | -- |
| I5 | -- | -- | 0.42 | -- | -- | -- |
| M12 spec | -- | 7.0 | -- | 0.7 | 0.7 | 0.7 |
| M12 | -- | 8.95 | -- | 0.78 | 0.42 | 0.45 |
| M13 spec | -- | 5.7 | -- | 0.7 | 0.7 | 0.7 |
| M13 | -- | 7.47 | -- | 0.75 | 0.37 | 0.40 |
| M14 spec | 12.3 | 2.4 | -- | 0.6 | 0.4 | 0.5 |
| M14 | 4.56 | 0.85 | -- | 0.35 | 0.29 | 0.32 |
| M15 spec | 2.1 | 0.6 | -- | 0.4 | 0.4 | 0.4 |
| M15 | 1.44 | 0.45 | -- | 0.25 | 0.22 | 0.23 |
| M16 spec | 1.6 | 0.6 | -- | 0.4 | 0.4 | 0.4 |
| M16 | 1.06 | 0.34 | -- | 0.22 | 0.20 | 0.20 |

Table 5: Total uncertainties compared to specifications in temperature [K]

| Band | Scene Temperature |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 190 | 230 | 267 | 270 | 310 | 340 |
| I4 spec | -- | -- | 0.91 | -- | -- | -- |
| I4 | -- | -- | 0.47 | -- | -- | -- |
| I5 spec | -- | -- | 1.4 | -- | -- | -- |
| I5 | -- | -- | 0.23 | -- | -- | -- |
| M12 spec | -- | 0.92 | -- | 0.13 | 0.17 | 0.21 |
| M12 | -- | 1.11 | -- | 0.14 | 0.10 | 0.13 |
| M13 spec | -- | 0.85 | -- | 0.14 | 0.19 | 0.23 |
| M13 | -- | 1.01 | -- | 0.15 | 0.10 | 0.13 |
| M14 spec | 2.6 | 0.75 | -- | 0.26 | 0.23 | 0.34 |
| M14 | 0.90 | 0.26 | -- | 0.15 | 0.16 | 0.22 |
| M15 spec | 0.56 | 0.24 | -- | 0.22 | 0.28 | 0.34 |
| M15 | 0.38 | 0.17 | -- | 0.13 | 0.16 | 0.19 |
| M16 spec | 0.48 | 0.26 | -- | 0.24 | 0.31 | 0.37 |
| M16 | 0.30 | 0.15 | -- | 0.13 | 0.16 | 0.19 |

Figure 1: I4 individual uncertainty terms in EV Lret in % of radiance

Figure 2: I5 individual uncertainty terms in EV Lret in % of radiance

Figure 3: M12 individual uncertainty terms in EV Lret in % of radiance

Figure 4: M13 HG individual uncertainty terms in EV Lret in % of radiance

Figure 5: M13 LG individual uncertainty terms in EV Lret in % of radiance

Figure 6: M14 individual uncertainty terms in EV Lret in % of radiance

Figure 7: M15 individual uncertainty terms in EV Lret in % of radiance

Figure 8: M16 individual uncertainty terms in EV Lret in % of radiance

Figure 9: Total uncertainties in EV Lret in % of radiance

Figure 10: Total uncertainties in EV Tret in temperature [K]

Figure 11: Total uncertainties in M13 LG EV Lret in % of radiance

Figure 12: Total uncertainties in M13 LG EV Tret in temperature [K]

Figure 13: Variation of total uncertainties in I4 EV Lret with scan angle in temperature [K]

Figure 14: Variation of total uncertainties in I5 EV Lret with scan angle in temperature [K]

Figure 15: Variation of total uncertainties in M12 EV Lret with scan angle in temperature [K]

Figure 16: Variation of total uncertainties in M13 HG EV Lret with scan angle in temperature [K]

Figure 17: Variation of total uncertainties in M14 EV Lret with scan angle in temperature [K]

Figure 18: Variation of total uncertainties in M15 EV Lret with scan angle in temperature [K]

Figure 19: Variation of total uncertainties in M16 EV Lret with scan angle in temperature [K]

Figure 20: Total uncertainties in EV Lret (baseline and worst-case) in % of radiance

Figure 21: Total uncertainties in EV Lret (baseline and worst-case) in temperature [K]