VCST Internal Memo

Title: Assessment of ETP-511 AOA Level Polarization Sensitivity Measurements for the JPSS-2 VIIRS VisNIR Bands

Memo Number: 2016_004

Revision: 01

Date: January 20, 2016

Author: Jeff McIntire

To: Xiaoxiong Xiong and James Butler

Cc: Hassan Oudrari, Kwo-Fu (Vincent) Chiang, Jon Fulbright and Aisheng Wu

# References

- [1] BDAY.ETP-511AOA Polarization Characterization Test Rev 5, ‘AOA Polarization Characterization test,’ Karthik Naidu.

- [2] VIIRS.03.02.072_J2 AOA PTSA Test Plan and STE Character Analysis v1, ‘J2 AOA Polarization test Source Assembly (PTSA) test Plan, Test Configuration, test analysis in prep for AOA polarization test,’ Tung R Wang and Tim P Johnson, October 21, 2015.

- [3] VCST_TECH_REPORT_15_021, ‘Preliminary JPSS-2 VIIRS AOA Polarization Analysis,’ Jeff McIntire, November 9, 2015.

- [4] J2 AOA pol test analysis v5, ‘J2 AOA polarization data analysis,’ Tim Johnson and Tung R Wang, November 12, 2015.

- [5] VIIRS_J2_AOA_polarizaion_Aero_151029, ‘J2 VIIRS Polarization Testing at AOA Level,’ David Moyer, October 29, 2015.

[6] VIIRS PRD Rev C, ‘Joint Polar-orbiting Satellite System (JPSS) Visible Infrared Imaging Radiometer Suite (VIIRS) Performance Requirements Document (PRD).’

[7] VCST_TECH_MEMO_14_002, ‘Assessment of FP-11 Polarization Sensitivity for the VIIRS F2 VisNIR Bands,’ Jeff McIntire, March 13, 2014.

[8] Mathematical Methods in the Physical Sciences, Mary L. Boas, John Wiley & Sons, 1985.

[9] VCST_TECH_MEMO_14_006, ‘Assessment of ETP-078 AOA Level Polarization Sensitivity for the VIIRS F2 VisNIR Bands,’ Jeff McIntire, April 9, 2014.

1. Introduction

Measurements of the JPSS-2 VIIRS sensor polarization sensitivity for the VisNIR bands were made on the Aft Optics Assembly (AOA) during ETP-511 [1-2]. Preliminary Analysis was reported in [3-5]. This work provides an overview of the test setup and objectives, analysis methodology, and results.

2. Objective

The objective of the ETP-511 test was to determine the sensitivity of the VIIRS instrument to input linearly polarized light. There are two specifications which relate to the polarization sensitivity of VIIRS [6]:

V_PRD-12624 – The VIIRS Sensor linear polarization sensitivity of the VIS and NIR bands shall be less than or equal to the values indicated in Table 1 for scan angles less than 45 degrees off Nadir.

V_PRD-12667 – The VIIRS Sensor linear polarization sensitivity shall be measured within a characterization uncertainty of 0.5% (one sigma) for scan angles less than 55.84 degrees off Nadir.

Both specifications are with respect to the sensor level polarization sensitivity, which will be measured in upcoming ambient phase testing. However, the results of the AOA testing will be compared to the first specification in this memo; it should be noted that the specification is not binding here.

2. Test Equipment

The test setup for ETP-511 consisted of an integrating sphere, polarizers, and various baffling [2]. A general schematic of the test setup is shown in Figure 1 indicating the placement of the polarizers relative to the optical path.

The SIS-100-2 was used as a source for ETP-511 testing [2]. The lamp configuration used was BDF.

Two different polarizers were used during testing: BVONIR and MOXTEK. The BVONIR sheet polarizer polarizes through scattering while the MOXTEK polarizer is a wire – grid polarizer. These polarizers were mounted on a rotary stage which could be cycled between 0 and 360 degrees. A second fixed polarizer of each type was used in conjunction with the first to measure the polarizer efficiency.

3. Test Configuration

The test was conducted at two M band integration times, 48.56 µs and 8.83 µs, and one I band integration time 26.49 µs [2]. The first M band setting was used in the measurements of bands M1 – M3 and the second was used for bands M4 – M7. Three test configurations were used during ETP-511 testing: BVONIR with the first M band integration time, BVONIR with the M band second integration time, and MOXTEK with the M band first integration time. Dark, efficiency, and polarization sensitivity tests were performed for each configuration.

First, a dark test was performed to determine a stray light pedestal to remove from the response. The rotating polarizer was fixed at 0 degrees. Dark data was taken before and after each efficiency or sensitivity test.

Next, a second fixed polarizer (the same type as the rotating polarizer) was inserted into the path between the rotating polarizer and VIIRS. The rotating polarizer was then cycled through 25 different angles from 0 to 360 degrees in 15 degree increments.

Lastly, the polarization sensitivity of VIIRS instrument was measured for all three configurations (the BVONIR configurations were measured twice and the MOXTEK configuration was measured three times). The rotating polarizer was then cycled through 25 different angles from 0 to 360 degrees in 15 degree increments.

Additional data was collected with half of the AOA aperture blocked (either the upper or lower half). These tests were performed using the MOXTEK polarizer.

4. Methodology

A standard Fourier analysis was used to determine the polarization sensitivity of the VIIRS instrument [7,8]. The Fourier expansion is written as

dn=12c0+ncncosnα+ndnsinnα,							(1)

where the Fourier coefficients are defined by the following:

cn=1πcosnθdnθ∂θ,									(2)

and

dn=1πsinnθdnθ∂θ.									(3)

In this work, only the zeroth through fourth order Fourier coefficients were calculated (for the polarizer efficiency tests, only the zeroth and second order terms were calculated). Eq. (1) can be rewritten as

dn=12c01+n=14ancosnα−δn,								(4)

where the linear polarization sensitivity of the instrument (n=2) is defined as

a2=c22+d2212c0a2eff										(5)

and the phase angle is defined as

δ2=tan−1d2c2.										(6)

The linear polarization sensitivity is also referred to as the degree of linear polarization (DoLP). The efficiency of each polarizer type is defined as

a2eff=c22+d2212c0,										(7)

determined from the cross polarizer testing.

5. Analysis

The data received was already pre-processed into background subtracted dn. The ETP-511 data is listed in Table 2 (including the type of test and configuration). The averaged dark offset was removed from each measurement using the dark offset data listed in Table 2; each of the packets included a dark measurement before and after the measurement, the average of which was used in this work Inside the data packets were measurements corresponding to discrete polarizer sheet angles, cycled from 0 to 360 degrees in 15 degree increments. This set of 25 measurements was then used to determine the Fourier coefficients in Eqs. (2) – (3), and finally the linear polarization sensitivity and phase from Eqs. (5) and (6). First, the cross polarizer efficiency test was analyzed; then the efficiency correction defined by Eq. (7) was used in the final polarization sensitivity calculation.

6. Results

Cross polarizer test data was analyzed for the BVONIR and MOXTEK configurations and the DoLP for each configuration was determined. The efficiency was calculated for each band and detector; the results for the BVONIR polarizer using the higher and lower M band integration time (black and red respectively) along with the MOXTEK results (blue) are shown in Figure 2. For each band, the detectors are plotted on the horizontal axis. Note that there is reasonable agreement between the BVONIR measurements for most detectors in M3; the signal is much lower for this band with the reduced integration time, and so the efficiency is less well determined. Also, the I1 and I2 efficiencies agree well as both measurements use the same integration time. The BVONIR efficiencies are in all cases above 93.5 % and below 96 %. The MOXTEK polarizer, for which data is available for only bands I1 – I2 and M1 – M3, shows efficiencies above 99 %. For bands I1 and I2, the MOXTEK efficiencies are very close to, and in some cases above, 100 %. An efficiency above 100 % is unphysical; in these cases, an efficiency of 100 % is used.

The polarization sensitivity was derived for all VisNIR bands and detectors. Figure 3 plots the measured dn for band M1 data from TC12 using the BVONIR polarizer and the nominal M band integration time. The lines indicate the calculated Fourier series using the zeroth through fourth order terms; the Fourier series reproduces the observed behavior very well. Note that the amplitude and phase of the Fourier series varies with detector. Figure 4 shows the zeroth through fourth order Fourier coefficients for M1 versus detector. The zeroth and second order terms are dominant, but the first order term is non-negligible; the third and fourth order terms are very small. The results shown in Figures 3 and 4 are indicative of all VisNIR bands results.

Figure 5 compares the second order term a2 across AOA test configurations using the full AOA aperture. The different test configurations as well as repeated measurements generally agree on the overall magnitude and detector dependence (the MOXTEK repeatability was better than the BVONIR). The results for band M1 range from 3.5 % to 4.5 % based on the MOXTEK measurements; note that this is above the specified limit of 3 %. All other bands shown polarization sensitivities of less than 2 %. For JPSS-1 VIIRS, the high polarization sensitivities in bands M1 – M4 [6] led to the filters being redesigned; bands M2 – M4 show significant improvement, while band M1 is similar in magnitude to JPSS-1 AOA measurements. The other bands show similar results to JPSS-1 AOA measurements as expected since their filters used the same prescriptions. Figure 6 shows the JPSS-1 AOA results [9]. The configurations were slightly different, but the band M1 results are consistent with JPSS-2 (only slightly higher). As a result, it is expected that band M1 will not meet the specified limit of 3 % at the system level as was the case for JPSS-1.

The corresponding AOA polarization phases (δ2) are compared for all bands and detectors in Figure 7. In general, there is good agreement between the AOA configurations. Note that the AOA measurements were measured with a different focal plane orientation relative to the upcoming system level test, so the AOA phases will need to be modified by π.

Additional testing was performed where half of the AOA aperture was blocked (upper half or lower half) using the MOXTEK polarizer. The polarization factors and phases derived from this data are shown in Figures 8 and 9 in comparison to the three MOXTEK measurements where no aperture blocking was performed. These tests were performed in order to further refine the polarization model ahead of system level testing.

7. Summary

ETP-511 polarization sensitivity testing was performed under at the AOA level for the JPSS-2 VIIRS sensor. Analysis showed the following:

- Linear polarization sensitivity for band M1 was observed to be higher than the specified limit (as high as ~4.5 %) based on MOXTEK results. Note that the specification applies to the system level measurements.

- All other bands show polarization sensitivities below ~2 %, with detector to detector variations of ~1 % or less. These detector to detector differences are likely the result of angle of incidence changes on the filter assembly.

- Comparisons between test configurations in general agreed (using the BVONIR polarizer sheet and the wire – grid MOXTEK polarizer).

Table 1: specified maximum polarization sensitivity [6]

| Band | Sensitivity [%] |
| --- | --- |
| I2, M1, M7 | 3 |
| I1, M2, M3, M4, M5, M6 | 2.5 |

Table 2: data used in the ETP-511 test analysis. The dark offset data used for each test is listed in parentheses. Nominal and reduced refer to the M band integration time.

| Test | AOA Aperture | Configuration |  |  |
| --- | --- | --- | --- | --- |
|  |  | BVONIR (Nominal) | BVONIR (Reduced) | MOXTEK |
| Efficiency | Full | TC7 (TC6) | TC8 (TC6) | TC10 (TC9) |
| Sensitivity | Full | TC12 (TC11) TC17 (TC16) | TC13 (TC11) TC18 (TC16) | TC15 (TC14) TC20 (TC19) TC22 (TC21) |
|  | Upper Half |  |  | TC25 (TC24) |
|  | Lower Half |  |  | TC27 (TC26) |

Figure 1: Schematic of the polarization sensitivity test setup

Figure 2: Polarizer efficiencies determined from ETP-511 testing

Figure 3: dn as a function of polarizer angle for M1 using BVONIR using the nominal M band integration time

Figure 4: Fourier coefficients for M1 using BVONIR using the nominal M band integration time

Figure 5: Polarization factor a2 in [%] across AOA test configurations

Figure 6: Polarization factor a2 in [%] across AOA test configurations from JPSS-1 [9]

Figure 7: Polarization phase δ2 in [radians] across AOA test configurations

Figure 8: Polarization factor a2 in [%] using the MOXTEK polarizer with and without half aperture blocking

Figure 9: Polarization phase δ2 in [radians] using the MOXTEK polarizer with and without half aperture blocking