## Slide 1
S-NPP VIIRS Thermal Emissive Bands On-Orbit Performance and Calibration
Xiaoxiong (Jack) Xiong……
VCST, NASA/GSFC
Acknowledgements: VIIRS SDR Team Members
VCST Members
May 18 2015
Notes: 1
1

## Slide 2
Outline
  TEB Calibration
  On-orbit Performance
    BB performance
    Detector short-term stability and long-term response (F-factors)
    Detector noise characterization (NEdT)
    Trending during WUCD
    Potential Improvements and Uncertainty Estimates
      Uncertainty assessment
      Improving M13LG calibration
      Average vs per-scan F-factor
      Moon in SV processing
    Conclusions
Notes: 2
2

## Slide 3
Calibrated using an on-board blackbody (BB):
  Scaling factor “F-factor” is derived and applied each scan.
  Warm-up and cool-down (WUCD) cycles are performed quarterly to fully characterize TEB detector response, including offset and nonlinear terms.
Band | I4 | I5 | M12 | M13 | M14 | M15 | M16
Wavelength [µm] | 3.74 | 11.45 | 3.70 | 4.05 | 8.55 | 10.76 | 12.01
5 M-bands and 2 I-bands, covering wavelengths from 3.7 to 12mm
Thermal Emissive Bands (TEB)
Notes: 3
3

## Slide 4
TEB Calibration Methodology
where the Lbg(B,) is the background difference between the EV and SV path:
VIIRS Earth View radiance is retrieved following ATBD Eq.(116)
the F-factor is derived each scan for each band, detector, and HAM-side:
Estimated BB radiance
Retrieved BB radiance
and the aperture radiance from the BB is:
dn: detector response
ci: calibration coefficients
RVS: response versus scan-angle
Notes: 4
4

## Slide 5
Long-term trend of daily-averaged TBB
  Stable to within a few mK.
  ~15mK offsets were due to the use of two different TBB settings.
BB Performance
Short-term stability (scan-by-scan TBB):
  Orbital variations of individual thermistors up to 40mK
  Variations in average temperature ~ 20mK
  Temperature difference between individual thermistors up to 60mK
  BB uniformity meets the requirement with standard deviation less than 30mK
30mK uniformity requirement
Day / Night
Notes: 5
5

## Slide 6
Orbits: 17553, 17554, 17555
* For clarity the F-factors are shifted.
Detector Short-term Stability
Scan-by-scan (HAM-A) Granule average (HAM-A)
Detector responses (F-factors) show small orbital variations:
±0.2% or less for scan-by-scan
±0.1% or less for granule average
Would using averaged F-factors improve SDR product?
F-factor orbital variations correlate with TBB variations and instrument temperatures variations.
Notes: 6
6

## Slide 7
* For clarity the F-factors are shifted.
Short-term Stability- Individual Detectors
M16 (not shown) similar to M15; same D16 our of family behavior
Orbits: 17553, 17554, 17555
Granule average (HAM-A)
Notes: 7
7

## Slide 8
Detector Long-term Response
Band | I4 | I5 | M12 | M13 | M14 | M15 | M16
Average F-factor: 03 26 2012 | 1.0105 | 1.0040 | 1.0035 | 1.0070 | 0.9946 | 1.0056 | 1.0101
Average F-factor: 03 25 2015 | 1.0142 | 1.0146 | 1.0072 | 1.0094 | 0.9964 | 1.0069 | 1.0124
Trend [%] | 0.36 | 1.05 | 0.37 | 0.24 | 0.19 | 0.13 | 0.22
Daily average F-factor trend:
  From Jan 20, 2012 (orbit 1200) to Mar 24, 2015 (orbit 17637)
  I5 shows the most noticeable trend of ~1%, while the rest of the bands are within 0.37% (M12).
  Discontinuities in the trend are coincident anomalies during which the cold FPA and/or instrument temperatures changed.
  Features in LWIR bands F-trend appear to coincide with the passage of the Earth through perihelion.
Petulant anomaly
SC anomaly
20120324
Earth in perihelion
Notes: 8
8

## Slide 9
Instrument Temperatures Trend
  Discontinuities in the instrument temperatures trends coincident with discontinuities in the F-factor trends shown on previous slide.
  Features in instrument temperature trends appears to coincide with the passage of the Earth through perihelion. The F-factor for LWIR bands shows features at the same time.
  There is a small increasing trend of SMWIR focal plain temperature, which can explain the trend observed in MWIR bands (I4, M12, M13).
Notes: 9
9

## Slide 10
Detector Noise Characterization (NEdT)
NEdT routinely trended at 292.5K: stable since the CFPA temperatures reached ~80K (orbit 1200). Band averaged values are within 0.2 K for I bands and 0.07 K for M bands
NEdT at TTYP derived periodically from BB WUCD data: stable and meeting the sensor design requirements by a wide margin:
Notes: 10
10

## Slide 11
Detector Noise Characterization (NEdT)
NEdT at TTYP (derived from BB cool-down data)
Continue to meet the sensor design requirements
Band | TTYP [K] | NEdT at Ttyp [K] |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  | Requirement | 02/12 | 05/12 | 09/12 | 12/12 | 03/13 | 06/13 | 09/13 | 12/13 | 03/14 | 06/14 | 09/14 | 12/14 | 03/15
I4 | 270 | 2.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
I5 | 210 | 1.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4
M12 | 270 | 0.396 | 0.13 | 0.13 | 0.13 | 0.11 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12
M13 | 300 | 0.107 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04
M14 | 270 | 0.091 | 0.05 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06
M15 | 300 | 0.070 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
M16 | 300 | 0.072 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03
Notes: 11
11

## Slide 12
Band-average c1 coefficients, as derived from the thirteen WUCD till Mar 2015, are shown in red (WU data), and blue (CD data) and compared to pre-launch (green). I5 c1 starts to show noticeable trend, consistent with results from F-factor trending.
Band-average c1 coefficients derived during WUCD cycles are within 1.9% on average (at M16 CD) from pre-launch values.
An offset between WU and CD results is present through the thirteen WUCDs, most prominent for LWIR bands.
12
Calibration Coefficients – Band-average c1
Y-range spans c1LUT 4% c1LUT
Notes: 12
12

## Slide 13
13
C0 Coefficients
Y-range spans: c0LUT 0.002 (MWIR),
c0LUT 0.1 (LWIR)
Band average c0
Detector specific c0/c0LUT
Notes: 13
13

## Slide 14
EV retrieved radiance uncertainty propagated using standard NIST formulation (k=1)
Some uncertainty contributors determined pre-launch by the instrument vendor: RTA reflectance BB emissivity
Radiometric coefficient and RVS uncertainties determined from NASA pre-launch analysis
Uncertainties investigated for a range of input signal levels and scan angles
Uncertainty Estimates
Total Uncertainty
Radiance (BB, SH, CAV, HAM, RTA)
Reflectance factors off BB
Radiometric coefficients
Response
RVS
RTA reflectance
BB emissivity

## Slide 15
Band | 267 K
I4 spec | 0.91
I4 estimate | 0.468
I5 spec | 1.4
I5 estimate | 0.226
Comparison to Requirement [K]
Uncertainty specifications
Defined in terms of %, at particular uniform scene temperatures, converted to K
Estimates exceed the specification at lower scene temperatures for bands M12 and M13
Band | 190 K | 230 K | 270 K | 310 K | 340 K
M12 spec | --- | 0.92 | 0.13 | 0.17 | 0.21
M12 estimate | --- | 1.11 | 0.13 | 0.07 | 0.09
M13 spec | --- | 0.85 | 0.14 | 0.19 | 0.23
M13 estimate | --- | 1.01 | 0.14 | 0.07 | 0.10
M14 spec | 2.60 | 0.75 | 0.26 | 0.23 | 0.34
M14 estimate | 0.95 | 0.26 | 0.12 | 0.12 | 0.20
M15 spec | 0.56 | 0.24 | 0.22 | 0.28 | 0.34
M15 estimate | 0.42 | 0.18 | 0.12 | 0.13 | 0.19
M16 spec | 0.48 | 0.26 | 0.24 | 0.31 | 0.37
M16 estimate | 0.35 | 0.16 | 0.12 | 0.14 | 0.19

## Slide 16
Uncertainty [K]
Uncertainty contributors:
Dominant for MWIR bands are the relative BB radiance uncertainty and the relative EV dn uncertainty (increasing rapidly with decreasing scene temperature).
The LWIR bands uncertainties are dominated by the c0, RVS, and EV dn relative uncertainties, which increase with decreasing scene temperatures.

## Slide 17
M13 LG Calibration
M13 low gain: No scan by scan F factor correction
Prelaunch analysis differs between Government team (Aerospace and VCST ) and sensor subcontractor – current LUT. Government team results are:
  c1 = 0.142 - 7% higher than LUT value c1LUT = 0.132;
  c0 = 0 - inconsistent with c0LUT = 1.15
Proposal:
Update M13 low gain coefficients based on Government team pre-launch analysis, which is consistent with results from on-orbit calibration
On-orbit comparison of lunar images in M13 LG and M13 HG - supports Government team pre-launch results:
  c1 = 0.142; 7% higher than c1LUT -consistent with Gov. team pre-launch
  c0 = 0 consistent with Gov. team pre-launch
M13 LG c1LUT, c0LUT
M13 LG c1=0.142, c0=0

## Slide 18
Effect of F-factor Noise
Evaluating the effect of using average F-factors
The VCST VIIRS SDR code was modified to apply average F-factors instead of per-scan F-factors for TEB calibration.
The F-factors for each band, detector, HAM are averaged over 24 scans.
Using average F-factors does not significantly impact the SDR product.
Striping on the noise level affects SST products based on M15 and M16 brightness temperatures.
T(M15)-T(M16)
SDR: d20130121_t0736504_e0742307
Av.F
Orig

## Slide 19
TEB Calibration when Moon in SV
Currently for TEB, Fill values are assigned in EV SDR when the Moon is in the SV.
Improved algorithm computes the mean and standard deviation of a 48-frame sample each scan. Then the outlier samples (Moon intrusion) with selected rejection scheme are identified and excluded from the SV average for background subtraction.
Images of calibrated radiance from 4 consecutive Band M12 SDRs, generated with current SDR code (left) and modified (right) calibration algorithms (Data: Jan 22, 2013; Time 22:24:02). [Reference SPIE 2013, 8866-72]
Before
After
Notes: 19
19

## Slide 20
F-factors Orbital Variation Reduction
F-factor orbital variations are present, on the order of 0.05-0.1 %.
Changing the BB thermistor weighting can reduce the F-factor orbital variations. Using T3 and T6 yield less variation for most bands (except M13).
Improving the background model which would also reduce the F-factor orbital variations.
T3 & T6
T2 & T5
T1 & T4
average T1-T6

## Slide 21
Conclusions
S-NPP VIIRS on-orbit BB long-term (3+ yr) performance is very stable. Short-term (orbital) temperature variations are present but within the uniformity requirement of 30mK.
Detector response (F-factor) trending is stable, with I5 showing maximum band-average trend of 1% followed by M12 and I4. Small orbital variations are present (0.05-0.1%).
No change is observed for TEB detector noise characteristics. NEdT at Ttyp is in compliance with the requirements.
Uncertainty estimates: TEB meet calibration requirements for most scene temperatures; M12 and M13 have slightly larger than specified UC at low scene temperatures; Larger uncertainties in M13 low gain (above 350 K).
Improvements: Updates to M13 LG offset and linear coefficients to improve calibration; Modifications to the SDR code/algorithm to allow TEB calibration to be performed when the Moon is in SV; Modifications to SDR code to apply average F-factor do not have significant impact.
Notes: 21
21

## Slide 22
Back Up

## Slide 23
Dynamic Range Verification
Dynamic range verified using scheduled Lunar observations
  All detectors of all TEB bands meet the Tmin (marginal non-compliance at I4) and Tmax requirements
  For some detectors of some bands the radiance limits in the Radiance-to-Temperature LUT do not extend to the largest possible unsaturated radiance
Requirement
LUT limit
I4
I5
M12
M13HG
M14
M15
M16

## Slide 24
Detector Specific NEdT
Detector specific NEdT is stable through the mission.

## Slide 25
May 22 - 25 2012 ;
orb.: 2939 – 2984;
~46 orbits.
25
Warm-up Cool-down (WUCD) Cycles
WUCD cycles performed: Feb, May, Sep, Dec 2012; Mar, Jun, Sep, Dec 2013, Mar 2014
Feb. 6 -10 2012;
orb.: 1436 – 1494;
~59 orbits.
Sept. 10-12 2012 ;
orb.: 4509 – 4536;
~28 orbits.
Dec. 17-19 2012 ;
orb.: 5900 – 5928;
~29 orbits.
Mar. 18-20 2013 ;
orb.: 7191 – 7219;
~29 orbits.
June 17-19 2013 ;
orb.: 8482 – 8510;
~28 orbits.
Sept. 16-18 2012 ;
orb.: 9773 – 9801;
~28 orbits.
Dec. 16-18 2013 ;
orb.: 11064 – 11092;
~28 orbits.
Mar. 16-18 2014 ;
orb.: 11064 – 11092;
~28 orbits.
Notes: 25
25

## Slide 26
26
Warm-up Cool-down (WUCD) Cycles
WUCD cycles performed previously: Jun. 2014, Sept. 2014, Dec. 2014
June 23-25 2014 ;
orb.: 13746 – 13774;
~28 orbits.
Sept 17-19 2014 ;
orb.: 14937 – 14966;
~29 orbits.
Dec 15-17 2014 ;
orb.: 16228 – 16257;
~30 orbits.
Notes: 26
26

## Slide 27
27
WUCD 16-18 Mar 2015 Data Selection
Cool-down:
Orbits: 17528 – 17543.
TBB range: 266.9 K to 315K;
The scans used (~48624) are shown in blue.
Warm-up:
Orbits: 17519 – 17528; 17543 – 17548
TBB set to: 297.5K, 302.5K, 307.5K, 312.5K, 315.0K and 272.5K, 282.5K, 292.5K,
The scans used (~41616) are
highlighted in red.
Notes: 27
27

## Slide 28
28
Calibration Coefficients – c1/LUT
Band average:100* (c1on-orbit- c1LUT )/c1LUT:
 | I4 | I5 | M12 | M13 | M14 | M15 | M16
WU 02/12 [%] | 1.2 | -0.8 | 0.4 | 1 | -1.1 | -0.2 | -0.3
CD 02/12 [%] | 1.5 | 0.6 | 0.6 | 1.2 | 0.2 | 0.4 | 1.6
WU 05/12 [%] | 1.2 | -0.6 | 0.4 | 0.9 | -1.7 | -0.6 | -0.8
CD 05/12 [%] | 1.6 | 0.5 | 0.7 | 1.3 | -0.6 | 0.3 | 1.1
WU 09/12 [%] | 1.3 | 0.2 | 0.6 | 1.2 | -0.8 | 0.2 | 0.5
CD 09/12 [%] | 1.6 | 1 | 0.8 | 1.7 | 0.3 | 0.9 | 2.2
WU 12/12 [%] | 1.3 | -0.2 | 0.6 | 1.2 | -1.2 | 0.1 | 0.03
CD 12/12 [%] | 1.6 | 0.7 | 0.8 | 1.2 | -0.2 | 0.3 | 1.6
WU 03/13 [%] | 1.4 | 0.4 | 0.6 | 1.2 | -1.1 | 0.1 | 0.4
CD 03/13 [%] | 1.7 | 0.8 | 0.9 | 1.3 | -0.1 | 0.6 | 1.8
WU 06/13 [%] | 1.4 | 0.6 | 0.7 | 1.2 | -0.7 | 0.4 | 0.9
CD 06/13 [%] | 1.7 | 1.1 | 0.9 | 1.3 | -0.01 | 0.5 | 2
WU 09/13 [%] | 1.5 | 0.3 | 0.7 | 1.2 | -1.1 | 0.1 | 0.3
CD 09/13 [%] | 1.7 | 1.1 | 0.9 | 1.3 | 0.05 | 0.6 | 2
WU 12/13 [%] | 1.4 | -0.18 | 0.7 | 1.2 | -1.2 | 0.1 | 0.05
CD 12/13 [%] | 1.7 | 1.7 | 0.9 | 1.4 | 0.4 | 0.6 | 2.2
WU 03/14 [%] | 1.4 | 0.5 | 0.7 | 1.2 | -1.0 | 0.2 | 0.7
CD 03/14 [%] | 1.7 | 1.6 | 0.9 | 1.3 | 0.3 | 0.7 | 2.4
WU 06/14 [%] | 1.4 | 1.0 | 0.7 | 1.3 | -0.5 | 0.4 | 1.2
CD 06/14 [%] | 1.8 | 2.3 | 1.0 | 1.3 | 0.6 | 0.9 | 2.8
WU 09/14 [%] | 1.4 | 0.2 | 0.7 | 1.3 | -1.1 | 0.1 | 0.3
CD 09/14 [%] | 1.8 | 1.6 | 1.0 | 1.4 | 0.2 | 0.7 | 2.2
WU 12/14 [%] | 1.4 | 0.2 | 0.7 | 1.3 | -0.8 | 0.2 | 0.4
CD 12/14 [%] | 1.8 | 2.3 | 1.0 | 1.4 | 0.3 | 0.7 | 2.4
WU 03/15 [%] | 1.5 | 0.8 | 0.7 | 1.3 | -1.0 | 0.1 | 0.6
CD 03/15 [%] | 1.9 | 2.1 | 1.0 | 1.4 | 0.2 | 0.8 | 2.4
Notes: 28
28

## Slide 29
29
C2 Coefficients
Y-range spans
c2LUT 3 x c2LUT
Band average c2
Detector specific c2/c2LUT
Notes: 29
29

## Slide 30
Band | 190 K | 230 K | 270 K | 310 K | 340 K
M12 spec | --- | 7.00 | 0.70 | 0.70 | 0.70
M12 estimate | --- | 8.98 | 0.71 | 0.27 | 0.32
M13 spec | --- | 5.70 | 0.70 | 0.70 | 0.70
M13 estimate | --- | 7.50 | 0.69 | 0.26 | 0.31
M14 spec | 12.30 | 2.40 | 0.60 | 0.40 | 0.50
M14 estimate | 4.82 | 0.84 | 0.28 | 0.21 | 0.29
M15 spec | 2.10 | 0.60 | 0.40 | 0.40 | 0.40
M15 estimate | 1.59 | 0.47 | 0.22 | 0.19 | 0.22
M16 spec | 1.60 | 0.60 | 0.40 | 0.40 | 0.40
M16 estimate | 1.24 | 0.37 | 0.21 | 0.18 | 0.20
Band | 267 K
I4 spec | 5.00
I4 estimate | 2.55
I5 spec | 2.50
I5 estimate | 0.41
Uncertainty specifications
Defined in terms of %, at particular uniform scene temperatures
Estimates exceed the specification at lower scene temperatures for bands M12 and M13
Comparison to Requirement [%]

## Slide 31
Uncertainty [K]
Uncertainty contributors:
Dominant for MWIR bands are the relative BB radiance uncertainty and the relative EV dn uncertainty (increasing rapidly with decreasing scene temperature).
The LWIR bands uncertainties are dominated by the c0, RVS, and EV dn relative uncertainties, which increase with decreasing scene temperature.